Supplementary Material

Stabilization and quantitative measurement of nicotinamide adenine dinucleotide in human whole blood using dried blood spot sampling

Ryo Matsuyama1, 2, Tomoyo Omata2, Michiharu Kageyama2, Ryota Nakajima1, 3, Masanobu Kanou1, 3, Kei Yamana1, 3

1 Nutraceutical Group, Healthcare New Business Division, Teijin Limited, Hino, Tokyo, Japan

2 Discovery DMPK Research Group, Toxicology & DMPK Research Department, Teijin Institute for Bio-medical Research, Teijin Pharma Limited, Hino, Tokyo, Japan

3 NOMON Co. Ltd., Kasumigaseki, Chiyoda-ku, Tokyo, Japan

Corresponding author: Ryo Matsuyama

Nutraceutical Group, Healthcare New Business Division, Teijin Limited, Hino, Tokyo, Japan

r.matsuyama@teijin.co.jp

Table S1 Storage stability of nicotinamide adenine dinucleotide (NAD+) and nicotinamide mononucleotide (NMN) on dried blood spot (DBS) cards

- Storage at 4 ℃

|  |  | Storage stability (%) |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | NAD+ |  |  |  |  | NMN |  |  |  |
| DBS card | Storage period | Mean | SD | %CV | n |  | Mean | SD | %CV | n1) |
| DMPK-A | 2 days | 81.5 | 10.9 | 13.3 | 4 |  | 73.1 | 20.2 | 27.7 | 4 |
|  | 2 weeks | 79.1 | 12.5 | 15.8 | 18 |  | 80.6 | 18.4 | 22.9 | 18 |
|  | 1 month | 71.8 | 20.9 | 29.1 | 4 |  | 99.9 | 41.9 | 42.0 | 4(1) |
| DMPK-B | 2 days | 80.7 | 11.0 | 13.6 | 4 |  | 95.4 | 33.1 | 34.7 | 4 |
|  | 2 weeks | 87.8 | 11.2 | 12.8 | 23 |  | 93.6 | 38.1 | 40.7 | 23(1) |
|  | 1 month | 79.8 | 13.9 | 17.5 | 9 |  | 78.6 | 40.2 | 51.1 | 9(3) |
| DMPK-C | 2 days | 70.1 | 12.2 | 17.4 | 5 |  | 68.4 | 19.3 | 28.2 | 5 |
|  | 2 weeks | 66.6 | 16.3 | 24.4 | 19 |  | 52.9 | 15.3 | 29.0 | 19(3) |
|  | 1 month | 54.3 | 8.4 | 15.4 | 5 |  | NC | NC | NC | 5(3) |
| 903 protein saver | 2 days | 74.1 | 10.5 | 14.2 | 5 |  | 60.1 | 9.3 | 15.5 | 5 |
|  | 2 weeks | 77.2 | 12.6 | 16.3 | 14 |  | 64.5 | 29.6 | 46.0 | 14(1) |
|  | 1 month | 82.6 | 11.0 | 13.4 | 5 |  | 56.3 | 16.8 | 29.8 | 5 |

1) Numbers in parentheses represent the number of exclusions based on the limit of quantification (BLQ)

NC: Not calculated, as the concentration in most samples were below BLQ

Lower limit of quantification: 0.5 to 2 μM (NAD+), 0.25 to 1 μM (NMN)

Storage stability (%) = concentration after storage on DBS cards / concentration without cards before storage (day 0) × 100

- Storage at room temperature

|  |  | Storage stability (%) |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | NAD+ |  |  |  |  | NMN |  |  |  |
| DBS card | Storage period | Mean | SD | %CV | n |  | Mean | SD | %CV | n1) |
| DMPK-A | 2 days | 71.6 | 7.7 | 10.7 | 8 |  | 122.7 | 46.0 | 37.5 | 8 |
|  | 1 week | NT | NT | NT | NT |  | NT | NT | NT | NT |
|  | 2 weeks | 51.6 | 5.6 | 10.9 | 5 |  | 124.2 | 49.1 | 39.5 | 5 |
| DMPK-B | 2 days | 82.0 | 9.9 | 12.0 | 13 |  | 96.5 | 40.4 | 41.8 | 13 |
|  | 1 week | 87.2 | 21.4 | 24.5 | 7 |  | 69.1 | 28.3 | 40.9 | 7 |
|  | 2 weeks | 67.6 | 4.6 | 6.7 | 9 |  | 66.6 | 39.6 | 59.6 | 9(2) |
| DMPK-C | 2 days | 55.9 | 7.1 | 12.6 | 7 |  | NC | NC | NC | 7(4) |
|  | 1 week | NT | NT | NT | NT |  | NT | NT | NT | NT |
|  | 2 weeks | 50.2 | 7.9 | 15.7 | 5 |  | NC | NC | NC | 5(5) |
| 903 protein saver | 2 days | 66.2 | 7.6 | 11.4 | 7 |  | 79.8 | 15.6 | 19.6 | 7(1) |
|  | 1 week | NT | NT | NT | NT |  | NT | NT | NT | NT |
|  | 2 weeks | 50.3 | 6.4 | 12.7 | 5 |  | NC | NC | NC | 5(5) |

1) Numbers in parentheses represent the number of exclusions based on the limit of quantification (BLQ)

NC: Not calculated, as the concentrations in most samples were below BLQ

NT: not tested

Lower limit of quantification: 0.5 to 2 μM (NAD+), 0.25 to 1 μM (NMN)

Storage stability (%) = concentration after storage on DBS cards / concentration without cards before storage (day 0) × 100

Table S2. Recovery of nicotinamide adenine dinucleotide (NAD+) in human whole blood using the standard addition method

|  | Mean | SD | %CV |
| --- | --- | --- | --- |
| Initial (without NAD+ standard addition) n=9 |  |  |  |
| Concentration (μM) | 24.94 | 4.64 | 18.6 |
| NAD+ Standard Addition: Low (Amount added 0.45 nmol) n=8 |  |  |  |
| Calculated Concentration (μM) | 50.86 | 4.12 | 8.1 |
| Measured Concentration (μM) | 43.63 | 6.82 | 15.6 |
| Recovery (%) | 85.8 | 11.8 | 13.8 |
| NAD+ Standard Addition: High (Amount added 1.35 nmol) n=8 |  |  |  |
| Calculated Concentration (μM) | 110.75 | 4.13 | 3.7 |
| Measured Concentration (μM) | 100.68 | 7.61 | 7.6 |
| Recovery (%) | 91.1 | 8.2 | 9.1 |

Human whole blood (12.5 μL) was added in a tube with 2.5 μL of saline containing 540 μmol/L NAD+ or 180 μmol/L NAD+ physiological saline. Samples were mixed by gentle pipetting and 5 μL was quickly spotted on a funnel filter paper No.5B (Kiriyama Glass Works Co., Tokyo, Japan). After drying for 10–15 min, the filter paper with blood spots was cut with scissors and placed into a tube with 300 μL of extraction buffer containing internal standard (IS; water in 667 nmol/L d4-NAD+ and 333 nmol/L d5-NMN). Samples were mixed using a micromixer (TOMY Microtube Mixer MT-400; TOMY Digital Biology Co. Ltd., Tokyo, Japan) for 30 min at room temperature to prepare a blood extraction solution. A portion of the blood extraction solution was placed into a tube with a 2-fold volume of acetonitrile, followed by mixing. Samples were centrifuged (24,250 × g, 10 min, 4 °C) and the supernatant was analyzed using LC-MS/MS. The recovery was calculated as follows:

Recovery (%) = Measured Concentration / Calculated Concentration × 100