## RND
Index
Line | Description
30 | Weighing a cold (°C) bottle and effect on the uncertainty on mass
59 | Preparation of the "MOOS-4 spike" nitrite solution
77 | Planning the re-bottling of MOOS-4
129 | Mass loss of old MOOS-4 and SEAN-1 material: traspiration from the irradiated bottles over 6 years
197 | Cleaning bottle test
221 | Screening nitrate on the starting material left at RT for 9 months aftern irradiation: almost full recovery and homogenous
248 | Screening phosphate on the starting material left at RT for 9 months after irradiation: full recovery and homogenous
294 | Screening nitrite on the starting material left at RT for 9 months after irradiation: nitrite formation
324 | Irradiation of the nitrite spike solution eliminates nitrite
341 | Brief short-term stability of the MOOS-4 starting material
504 | Screening of nutrient on the 7 boxes before reblending
604 | EtOH interference on nitrite
627 | Preparation of the 15NO3- internal standard with exact-matching
666 | Preparation of the 15NO2- and 15NO3- internal standard with exact-matching for simultaneous determination
692 | Standardization NO2- Sigma Aldrich (verified 2024-04-22)
901 | Development of novel method for the determination of nitrite by GC-MS following derivatization with triethyloxonium
1220 | Validation figures nitrite and nitrate method and preliminary data on MOOS-4 and SEAN-1
1434 | Purification of oxonium reagent for nitrate measurement
1465 | Study matrix effects MOOS-4 by photometry (nitrite, phosphate, silicate; verified 2024-05-03)
1566 | Temperature effect on the determination of nitrite, phosphate and silicate by UV-vis
1654 | Silicate interference on phosphate
1679 | Validation of silicate by LC-ICPMS: calibration plot, analytical range, matrix effect, LOD, LOQ, signal repeatability
Weighing a cold (°C) bottle and effect on the uncertainty on mass
Experiments performed during the CCQM exercise K161
Cold (4 °C) bottle off the fridge (Nalgene 65 mL narrow neck)
Try to understand if weighing a cold liquid in a bottle is problematic for moisture
Weigh the bottle (1) then weigh the bottle (2-8) after wipeing the outside with a kim wipe
Data | Mass (g)
1 | 71.7093
2 | 71.7095
3 | 71.7089
4 | 71.7087
5 | 71.7084
6 | 71.708
7 | 71.7083
8 | 71.7076
Max | 71.7095
Min | 71.7076
Difference | 0.0019
Uncertainty | 0.002
This uncertainty on the 30 g standard solution is 0.002/30 *100% = 0.007%
From the experiments performed on the preliminary data, the impact of weighing a cold solution seems to be much lowerd than 0.002 g
When the cold LNSW is used to dilute the standard, there is no differences between direct weighing and after wiping the bottle
The 0.002 is a conservative choise which adds a modest contribution to the overall uncertainty
Preparation of the "MOOS-4 spike" nitrite solution
Date | 20200228
Bottle (Nalgene, 2 L) = 179 g
Added 70 g of NaCl from Sigma Aldrich S7653 lot SLB0434V BioExtra 99.5%
Added ~ 2L of water GenPure (B3)
Added 5 mL of 0.1 M NaHCO3 Sigma Aldrich 36486 lot BCBW9092
Final pH of 8.2
The amount of [CO3] was much less with respect of the total dissolved inorganic carbon typical of seawater
A NaNO2 solution was then prepared from Sigma Aldrich 563218-25G
NaNO2 = 0.1548 g
adding water to final mass of 48.0817
2 x 4.65 mL of the NaNO2 solution was added to the 2 L saline medium
For irradiation, 33 x 60 mL bottles where prepared and sent for irradiation along with the MOOS and SEAN samples
Planning the re-bottling of MOOS-4
1. Use of old bottle is not suitable
Date | 2023-03-03
When the cappers is applied to close the irradiated caps, we can observe some cracking of the PP material
The old bottles cannot be used again, we will need 60 mL new bottle from Fisher
The plan is to used EtOH to perform the sterilization of drum and bottles, rinse with water two times and leave the bottle with water for sometime
2. The nitrite spike
Total nitrite in the spiking solution in g:
0.019965
Total number of 60 mL bottles containing the spike:
33
The nitrite did not survive the irradiation, no spiking is necessary anymore, see line:
294
Mass loss of old MOOS-4 and SEAN-1 material: traspiration from the irradiated bottles over 6 years
The first proposed batch was bottled in 2017-03 and the mass of each unti was recorded
On 2022-06-08, I Inspected two boxes of material one box of MOOS candidate (box39 of 54) and one box of SEAN candidate (box 33 of 54)
These units were left at RT till today where the mass of the units was meaured again
The traspiration is minimal, less than:
-0.1815%
Material | Unit mass | Unit mass | Delta mass |  | Material | Unit mass | Unit mass | Delta mass
ID | at Bottling | 2023-03-01 |  |  | ID | at Bottling | 2023-03-01
| g | g | g |  |  | g | g | g
MOOS2123 | 75.7642 | 75.6766 | -0.0876 |  | SEAN1840 | 71.1949 | 71.1053 | -0.0896
MOOS2148 | 75.1234 | 75.0369 | -0.0865 |  | SEAN1828 | 74.6452 | 74.558 | -0.0872
MOOS2161 | 74.5862 | 74.4918 | -0.0944 |  | SEAN1830 | 72.1164 | 72.0281 | -0.0883
MOOS2120 | 71.6918 | 71.6022 | -0.0896 |  | SEAN1811 | 71.4389 | 71.3521 | -0.0868
MOOS2168 | 74.8295 | 74.7429 | -0.0866 |  | SEAN1829 | 73.0342 | 72.9453 | -0.0889
MOOS2124 | 72.402 | 72.3161 | -0.0859 |  | SEAN1820 | 71.2827 | 71.1953 | -0.0874
MOOS2155 | 74.6284 | 74.5382 | -0.0902 |  | SEAN1809 | 71.0982 | 71.0068 | -0.0914
MOOS2118 | 71.8595 | 71.7712 | -0.0883 |  | SEAN1810 | 73.3068 | 73.2139 | -0.0929
MOOS2143 | 71.1566 | 71.0679 | -0.0887 |  | SEAN1807 | 71.7545 | 71.6632 | -0.0913
MOOS2140 | 75.7435 | 75.6572 | -0.0863 |  | SEAN1796 | 73.202 | 73.1116 | -0.0904
MOOS2160 | 73.7569 | 73.6685 | -0.0884 |  | SEAN1839 | 73.4675 | 73.3772 | -0.0903
MOOS2117 | 72.0358 | 71.9461 | -0.0897 |  | SEAN1842 | 72.6799 | 72.588 | -0.0919
MOOS2131 | 75.268 | 75.1795 | -0.0885 |  | SEAN1819 | 72.5874 | 72.4953 | -0.0921
MOOS2127 | 74.9492 | 74.8598 | -0.0894 |  | SEAN1832 | 73.1485 | 73.0512 | -0.0973
MOOS2125 | 73.4957 | 73.4104 | -0.0853 |  | SEAN1795 | 72.8915 | 72.8024 | -0.0891
MOOS2156 | 72.8186 | 72.7293 | -0.0893 |  | SEAN1837 | 70.186 | 70.097 | -0.089
MOOS2169 | 74.2222 | 74.131 | -0.0912 |  | SEAN1806 | 72.0623 | 71.971 | -0.0913
MOOS2134 | 74.545 | 74.4567 | -0.0883 |  | SEAN1818 | 72.2 | 72.1074 | -0.0926
MOOS2149 | 73.8618 | 73.7736 | -0.0882 |  | SEAN1826 | 71.2171 | 71.1256 | -0.0915
MOOS2165 | 73.3432 | 73.2549 | -0.0883 |  | SEAN1797 | 72.5738 | 72.4838 | -0.09
MOOS2142 | 75.8582 | 75.7701 | -0.0881 |  | SEAN1805 | 71.6832 | 71.5939 | -0.0893
MOOS2141 | 72.7269 | 72.6375 | -0.0894 |  | SEAN1838 | 73.2085 | 73.1187 | -0.0898
MOOS2138 | 75.4357 | 75.3464 | -0.0893 |  | SEAN1827 | 71.0146 | 70.929 | -0.0856
MOOS2153 | 72.5916 | 72.5006 | -0.091 |  | SEAN1798 | 71.7688 | 71.6785 | -0.0903
MOOS2135 | 75.3238 | 75.238 | -0.0858 |  | SEAN1835 | 71.7783 | 71.6927 | -0.0856
MOOS2136 | 74.0812 | 73.9916 | -0.0896 |  | SEAN1815 | 71.5198 | 71.4312 | -0.0886
MOOS2129 | 71.2091 | 71.1197 | -0.0894 |  | SEAN1823 | 70.2274 | 70.1381 | -0.0893
MOOS2121 | 74.6875 | 74.5986 | -0.0889 |  | SEAN1816 | 70.8234 | 70.7311 | -0.0923
MOOS2167 | 75.2741 | 75.1879 | -0.0862 |  | SEAN1824 | 72.0332 | 71.941 | -0.0922
MOOS2122 | 75.63 | 75.5445 | -0.0855 |  | SEAN1825 | 71.4887 | 71.3798 | -0.1089
MOOS2137 | 73.0642 | 72.9747 | -0.0895 |  | SEAN1802 | 72.0258 | 71.9382 | -0.0876
MOOS2163 | 72.1762 | 72.0859 | -0.0903 |  | SEAN1794 | 73.3397 | 73.252 | -0.0877
MOOS2164 | 74.5205 | 74.4308 | -0.0897 |  | SEAN1834 | 69.4782 | 69.3759 | -0.1023
MOOS2151 | 71.7785 | 71.6913 | -0.0872 |  | SEAN1803 | 74.045 | 73.9565 | -0.0885
MOOS2119 | 73.6528 | 73.5648 | -0.088 |  | SEAN1792 | 70.4172 | 70.324 | -0.0932
MOOS2116 | 75.8791 | 75.7886 | -0.0905 |  | SEAN1822 | 72.9095 | 72.8209 | -0.0886
MOOS2145 | 74.0206 | 73.9317 | -0.0889 |  | SEAN1833 | 71.1062 | 71.0152 | -0.091
MOOS2139 | 71.3055 | 71.2171 | -0.0884 |  | SEAN1814 | 73.1096 | 73.0199 | -0.0897
MOOS2152 | 72.9703 | 72.8825 | -0.0878 |  | SEAN1791 | 72.2793 | 72.1867 | -0.0926
MOOS2128 | 75.3727 | 75.2831 | -0.0896 |  | SEAN1801 | 71.7895 | 71.6966 | -0.0929
MOOS2146 | 73.4016 | 73.3142 | -0.0874 |  | SEAN1789 | 71.6729 | 71.5835 | -0.0894
MOOS2162 | 74.7841 | 74.6929 | -0.0912 |  | SEAN1800 | 71.8634 | 71.7714 | -0.092
MOOS2126 | 75.7192 | 75.6331 | -0.0861 |  | SEAN1812 | 70.8811 | 70.7834 | -0.0977
MOOS2157 | 73.0685 | 72.9815 | -0.087 |  | SEAN1836 | 75.1508 | 75.0625 | -0.0883
MOOS2158 | 73.5029 | 73.4154 | -0.0875 |  | SEAN1813 | 74.5955 | 74.5097 | -0.0858
MOOS2166 | 74.625 | 74.5361 | -0.0889 |  | SEAN1821 | 74.8496 | 74.7583 | -0.0913
MOOS2130 | 72.4081 | 72.3199 | -0.0882 |  | SEAN1790 | 72.1467 | 72.056 | -0.0907
MOOS2133 | 73.146 | 73.0577 | -0.0883 |  | SEAN1831 | 69.004 | 68.9159 | -0.0881
MOOS2144 | 75.0263 | 74.9391 | -0.0872 |  | SEAN1799 | 72.581 | 72.4931 | -0.0879
MOOS2159 | 74.1575 | 74.0719 | -0.0856 |  | SEAN1788 | 73.5417 | 73.4539 | -0.0878
|  | Average | -0.088424 |  |  |  | Average | -0.090686
|  | SD | 0.0017632 |  |  |  | SD | 0.0040231
|  | RSD | -1.994% |  |  |  | RSD | -4.436%
|  | Max | -0.0853 |  |  |  | Max | -0.0856
|  | Min | -0.0944 |  |  |  | Min | -0.1089
Cleaning bottle test
Date | 2023-03-28
Date | 2023-04-03
|  | After | After | After | After | [7 empty] | 2023-04-03
Bottle | Empty | EtOH wash | water wash 1 | water wash 2 | water wash 3 |  | EtOH | H2O - 1 | H2O - 2 | H2O - 3 | H2O - 4 |  | empty | after 2 h | after 4h |  | H2O - 5 | H2O - 6 | H2O - 7
ID | g | g | g | g | g |  | g | g | g | g | g |  | g | g | g |  | g | g | g
1 | 11.7805 | 12.0046 | 11.7993 | 11.7896 | 11.8997 |  | 0.2241 | 0.0188 | 0.0091 | 0.1192 | 0.0231 |  | 11.7867 | 11.7832 | 11.7817 |  | 0.0062 | 0.0027 | 0.0012
2 | 11.7815 | 11.9737 | 11.7878 | 11.7858 | 11.7927 |  | 0.1922 | 0.0063 | 0.0043 | 0.0112 | 0.010133 |  | 11.7835 | 11.7817 | 11.7819 |  | 0.002 | 0.0002 | 0.0004
3 | 11.8214 | 12.0062 | 11.8447 | 11.8276 | 11.8309 |  | 0.1848 | 0.0233 | 0.0062 | 0.0095 | 0.013 |  | 11.8238 | 11.8222 | 11.822 |  | 0.0024 | 0.0008 | 0.0006
Standard deviation of the residual water |  |  | 0.028886 |  |  |  |  | w(EtOH)-1 | w(EtOH)-2 | w(EtOH)-3 | w(EtOH)-4
Added hom component |  |  | 0.1155% |  |  |  |  | µg/g | µg/g | µg/g | ng/g
Dilution |  |  | 0.0706% |  |  |  |  | 8964 | 6.7409 | 0.032141 | 0.014849
[8 empty] | 7688 | 1.9374 | 0.00086794 | 0.0001759
[8 empty] | 7392 | 6.8893 | 0.002618 | 0.00068067
One wash with EtOH sterilize the bottle
three washes with 25 mL water removes the EtOH
Then the bottle is stored with 50 mL of water for a few days both upright and upside down for removing all impurities
After the cleaning the bottle is emptied but not dried
2023-04-03: after the bottle has been kept full of water
Screening nitrate on the starting material left at RT for 9 months aftern irradiation: almost full recovery and homogenous
Date | 2023-03-13
Starting from the MOOS-4 candidate box 39 after irradiation
The box was pulled from the cold room on 2023-06-08 and left at RT in B5 till today
Five units of the material were analyzed
Unit | Vial | Sample | IS | mA | mB | Area | Area | rAB | Hom
| g | g | g | g | g | 46 | 47
MOOS2155 | 11.4593 | 13.4935 | 13.6919 | 2.0342 | 0.1984 | 4991.1755 | 5538.811 | 0.90113 | 0.087889
MOOS2169 | 11.4451 | 13.4731 | 13.6722 | 2.028 | 0.1991 | 5729.7895 | 6340.7905 | 0.90364 | 0.088715
MOOS2140 | 11.4054 | 13.4305 | 13.6301 | 2.0251 | 0.1996 | 5744.4295 | 6498.502 | 0.88396 | 0.087126
MOOS2165 | 11.5293 | 13.557 | 13.7565 | 2.0277 | 0.1995 | 5404.0055 | 6102.7715 | 0.8855 | 0.087122
MOOS2121 | 11.4349 | 13.4625 | 13.662 | 2.0276 | 0.1995 | 5647.9498 | 6336.6359 | 0.89132 | 0.087699
0.7497%
The same internal standard used for the SALT-1 project was used for this screening in same proportion
SALT-1 solution isotope ratio: |  | 1.0512
SALT-1 solution mass fraction: |  | 1.1427 | mg/kg
Nitrate in MOOS-4 before irradiation: |  |  | 1.0599 | mg/kg
Nitrate in MOOS-4 after irradiation: |  |  | 0.97082 | mg/kg
The depletion is marginal, after sterilization the material seems to be stable even at RT
Likely the depletion was cause by nitro/denitro bacteria during transportation (days that the material from cold room to irradiation, N was the only analyte that failed the long term stability)
Homogeneiety is good
Screening phosphate on the starting material left at RT for 9 months after irradiation: full recovery and homogenous
Date | 2023-03-14
Starting from the MOOS-4 candidate box 39 after irradiation
The box was pulled from the cold room on 2023-06-08 and left at RT in B5 till today
Five units of the material were analyzed
the SALT-1 was the same solution used for the CCQM and given to Kenny at a certain point
used a 5 cm cuvette zerod against GenPure water
Unit | Vial | Sample | Chemistry | mA | mF | Area | Area
ID | g | g | g | g | g | 890 nm | norm
MOOS2167 | 7.3152 | 13.4468 | 14.2413 | 6.1316 | 6.9261 | 0.27095 | 0.30606
MOOS2124 | 7.3053 | 13.398 | 14.1999 | 6.0927 | 6.8946 | 0.26966 | 0.30515
MOOS2117 | 7.3443 | 13.4401 | 14.2436 | 6.0958 | 6.8993 | 0.27046 | 0.30611
MOOS2163 | 7.3235 | 13.423 | 14.2267 | 6.0995 | 6.9032 | 0.26931 | 0.30479
MOOS2116 | 7.3214 | 13.435 | 14.2383 | 6.1136 | 6.9169 | 0.27339 | 0.30931
SALT-1 | 7.3355 | 13.3315 | 14.135 | 5.996 | 6.7995 | 0.091846 | 0.10415
MOOS-3 | 7.2834 | 13.3537 | 14.1582 | 6.0703 | 6.8748 | 0.17533 | 0.19857
BLK | 7.3207 | 13.5368 | 14.3387 | 6.2161 | 7.018 | -0.002385 | -0.0026927
Homogeneiety |  |  |  |  | average | 0.27075
0.5832%
2021-10-07 | MOOS-3
Signal | w(PO4)
0.17252 | 0.16354
Before irradiation the phosphate mass fraction was:
0.24514 | mg/kg
Today the mass fraction of phoshate is roughly:
0.25666 | mg/kg
Area norm measured before irradiation was:
0.28549
Calcualtion of phosphate increase after irradiation
Ratio 1 | 1.047
Ratio 2 | 1.0728
After irradiation the phosphate is stable even at room temperature with light
It was observed a slight increase of phosphate likely due to the irradiation process (+5-7 %)
Within box homogeneiety is good
Screening nitrite on the starting material left at RT for 9 months after irradiation: nitrite formation
Date | 2023-03-14
Starting from the MOOS-4 candidate box 39 after irradiation
The box was pulled from the cold room on 2023-06-08 and left at RT in B5 till today
Five units of the material were analyzed
used a 1 cm cuvette zerod against GenPure water
Unit | Vial | Sample | Chemistry | mA | mF | Area | Area
ID | g | g | g | g | g | 541 nm | norm
MOOS2167 | 6.1318 | 11.2231 | 11.4266 | 5.0913 | 5.2948 | 0.080771 | 0.083999
MOOS2124 | 6.1477 | 11.2492 | 11.4504 | 5.1015 | 5.3027 | 0.090395 | 0.09396
MOOS2117 | 6.1308 | 11.2229 | 11.4233 | 5.0921 | 5.2925 | 0.085665 | 0.089036
MOOS2163 | 6.1376 | 11.2248 | 11.4258 | 5.0872 | 5.2882 | 0.085456 | 0.088832
MOOS2116 | 6.1114 | 11.1877 | 11.3887 | 5.0763 | 5.2773 | 0.081848 | 0.085089
MOOS-3 | 6.0968 | 11.2061 | 11.4068 | 5.1093 | 5.31 | 0.17366 | 0.18048
BLK | 6.1075 | 11.1502 | 11.3512 | 5.0427 | 5.2437 | 0.000931 | 0.00096811
Homogeneiety |  |  |  |  | average | 0.084827
4.45%
The nitrite concentration is roughly half of that obtained for MOOS-3
0.077228 | mg/kg nitrite
After irradiation there is formation of nitrite (before irradiation it was undetected)
This also explains why the nitrite mass fraction of the previous MOOS was so high
The homogeneiety is not great but the re-mixing will help
Irradiation of the nitrite spike solution eliminates nitrite
Date | 2023-03-14
Starting from the spike prepared as describe in line:
77
It was observed that the spiking solution of nitrite is not good anymore
Unit | Vial | LNSW | Sample | Chemistry | mA | mF | Area | Area
ID | g | g | g | g | g | g | 541 nm | norm
Irradiated | 6.1253 | 11.2993 | 11.3631 | 11.5638 | 0.0638 | 5.4385 | 0.011214 | 0.95591
Non_irradiated | 6.1061 | 11.1689 | 11.233 | 11.4338 | 0.0641 | 5.3277 | 0.13851 | 11.512
After irradiation, 90 % of nitrite was lost
However, this is not a concern as the irradiation allowed formation of nitrite in-situ
Brief short-term stability of the MOOS-4 starting material
Date | 2023-03-15
Starting from the MOOS-4 candidate box 39 after irradiation
The box was pulled from the cold room on 2023-06-08 and left at RT in B5 till today
Five units of the material were opend and half of their content was transferred into a bottle just cleaned with GenPure water
The two half unit MOOS2122 were incubated in the 40 °C oven
The non-sterile half of unit 2152 was kept uncapped in B3 for 5 hours (MOOS2152T)
The other units (MOOS2159, MOOS2137 and MOOS2142) were kept in room B5 (RT in presence of light)
"T" = unit trasfered into another bottles (not pre-cleaned)
MOOS2164R | reference unit, open just the day of analysis after sterilization
Date | 2023-03-27
1. Determination of nitrate by ID GC-MS
Unit | Vial | Sample | IS | mA | mB | Area | Area | rAB | Hom |  | Unit | Ratio
ID | g | g | g | g | g | 46 | 47 |  |  |  | ID | "T" vs not "T"
MOOS2142 | 11.4592 | 13.4968 | 13.6529 | 2.0376 | 0.1561 | 5603.477 | 4991.9025 | 1.1225 | 0.085995 |  | MOOS2142 | 1.0023
MOOS2142T | 11.4671 | 13.5431 | 13.7419 | 2.076 | 0.1988 | 5379.1645 | 5976.292 | 0.90008 | 0.086193 |  | MOOS2159 | 0.9984
MOOS2159 | 11.44 | 13.4862 | 13.6844 | 2.0462 | 0.1982 | 5235.573 | 5652.2612 | 0.92628 | 0.089722 |  | MOOS2137 | 1.0006
MOOS2159T | 11.4758 | 13.5153 | 13.7143 | 2.0395 | 0.199 | 5600.249 | 6100.0964 | 0.91806 | 0.089578 |  | MOOS2152 | 0.99656
MOOS2137 | 11.614 | 13.6467 | 13.8452 | 2.0327 | 0.1985 | 6060.8226 | 6544.5625 | 0.92609 | 0.090435 |  | MOOS2122 | 1.0089
MOOS2137T | 11.4252 | 13.465 | 13.6633 | 2.0398 | 0.1983 | 5371.919 | 5771.4505 | 0.93077 | 0.090486
MOOS2152 | 11.5173 | 13.567 | 13.7662 | 2.0497 | 0.1992 | 5101.4225 | 5699.81 | 0.89502 | 0.086982
MOOS2152T | 11.4874 | 13.5194 | 13.7181 | 2.032 | 0.1987 | 5279.755 | 5955.99 | 0.88646 | 0.086683
MOOS2122 | 11.9807 | 14.0255 | 14.2238 | 2.0448 | 0.1983 | 5390.1895 | 5960.8796 | 0.90426 | 0.087693
MOOS2122T | 11.6341 | 13.662 | 13.8604 | 2.0279 | 0.1984 | 6020.7566 | 6657.815 | 0.90431 | 0.088474
MOOS2164R | 11.4788 | 13.5075 | 13.706 | 2.0287 | 0.1985 | 5290.5154 | 5946.8318 | 0.88964 | 0.087047
BLK | 11.4538 | 13.5744 | 13.5744 | 2.1206 | 0 | 11.6375 | 0
IS = 15NO3 used for the SALT-1 certification
0.1 mL of 1 % sulfamic acid
0.05 mL of triethyloxonium tetrafluoroborate solution (1 g reagent + 1 g cold water)
HS analysis with SALT1.M method
2. Determination of nitrite by UV-vis
Unit | Vial | Sample | Chemistry | mA | mF | Area | Area |  | Unit | Ratio
ID | g | g | g | g | g | 541 nm | norm |  | ID | "T" vs not "T"
MOOS2142 | 6.1173 | 11.2052 | 11.409 | 5.0879 | 5.2917 | 0.08778 | 0.091296 |  | MOOS2142 | 0.97208
MOOS2142T | 6.1443 | 11.2311 | 11.4327 | 5.0868 | 5.2884 | 0.085364 | 0.088747 |  | MOOS2159 | 0.99657
MOOS2159 | 6.1375 | 11.247 | 11.4481 | 5.1095 | 5.3106 | 0.085799 | 0.089176 |  | MOOS2137 | 0.99904
MOOS2159T | 6.0953 | 11.1845 | 11.3852 | 5.0892 | 5.2899 | 0.085498 | 0.08887 |  | MOOS2152 | 1.0053
MOOS2137 | 6.1374 | 11.2269 | 11.4278 | 5.0895 | 5.2904 | 0.091011 | 0.094604 |  | MOOS2122 | 0.98638
MOOS2137T | 6.1374 | 11.2236 | 11.4247 | 5.0862 | 5.2873 | 0.090918 | 0.094513
MOOS2152 | 6.1089 | 11.1964 | 11.3975 | 5.0875 | 5.2886 | 0.086925 | 0.090361
MOOS2152T | 6.1132 | 11.2084 | 11.4091 | 5.0952 | 5.2959 | 0.087398 | 0.090841
MOOS2122 | 6.1444 | 11.2269 | 11.4283 | 5.0825 | 5.2839 | 0.086635 | 0.090068
MOOS2122T | 6.1287 | 11.2253 | 11.4262 | 5.0966 | 5.2975 | 0.085472 | 0.088841
MOOS2164R | 6.1392 | 11.2328 | 11.4334 | 5.0936 | 5.2942 | 0.083739 | 0.087037
BLK | 6.156 | 11.4044 | 11.6058 | 5.2484 | 5.4498 | 8.7e-05 | 9.0339e-05
used a 1 cm cuvette zerod against GenPure water
3. Determination of phosphate by UV-vis
Unit | Vial | Sample | Chemistry | mA | mF | Area | Area |  | Unit | Ratio
ID | g | g | g | g | g | 890 nm | norm |  | ID | "T" vs not "T"
MOOS2142 | 6.1512 | 12.2566 | 13.0428 | 6.1054 | 6.8916 | 0.27522 | 0.31067 |  | MOOS2142 | 1.0003
MOOS2142T | 6.1697 | 12.2776 | 13.0634 | 6.1079 | 6.8937 | 0.27533 | 0.31076 |  | MOOS2159 | 1.0052
MOOS2159 | 6.1407 | 12.2678 | 13.0534 | 6.1271 | 6.9127 | 0.27778 | 0.3134 |  | MOOS2137 | 1.0092
MOOS2159T | 6.1142 | 12.2248 | 13.0112 | 6.1106 | 6.897 | 0.27911 | 0.31503 |  | MOOS2152 | 1.0016
MOOS2137 | 6.1004 | 12.1999 | 12.9869 | 6.0995 | 6.8865 | 0.27786 | 0.31371 |  | MOOS2122 | 0.99186
MOOS2137T | 6.1018 | 12.2225 | 13.008 | 6.1207 | 6.9062 | 0.2806 | 0.31661
MOOS2152 | 6.1016 | 12.207 | 12.9928 | 6.1054 | 6.8912 | 0.27459 | 0.30993
MOOS2152T | 6.1086 | 12.213 | 12.9991 | 6.1044 | 6.8905 | 0.27503 | 0.31044
MOOS2122 | 6.1275 | 12.2296 | 13.0158 | 6.1021 | 6.8883 | 0.28004 | 0.31613
MOOS2122T | 6.1069 | 12.2164 | 13.0024 | 6.1095 | 6.8955 | 0.27781 | 0.31355
MOOS2164R | 6.1189 | 12.218 | 13.0846 | 6.0991 | 6.9657 | 0.2735 | 0.31236
BLK | 6.1338 | 12.259 | 13.0536 | 6.1252 | 6.9198 | -0.001916 | -0.0021646
used a 5 cm cuvette zerod against GenPure water
4. Determination of silicate by UV-vis
Unit | Vial | Sample | Chemistry | mA | mF | Area | Area |  | Unit | Ratio
ID | g | g | g | g | g | 810 nm | norm |  | ID | "T" vs not "T"
MOOS2142 | 6.1106 | 9.162 | 9.4824 | 3.0514 | 3.3718 | 0.11997 | 0.13257 |  | MOOS2142 | 0.9979
MOOS2142T | 6.1462 | 9.1873 | 9.5095 | 3.0411 | 3.3633 | 0.11962 | 0.13229 |  | MOOS2159 | 1.0038
MOOS2159 | 6.1539 | 9.1631 | 9.486 | 3.0092 | 3.3321 | 0.11855 | 0.13127 |  | MOOS2137 | 1.0036
MOOS2159T | 6.0972 | 9.1489 | 9.4716 | 3.0517 | 3.3744 | 0.11916 | 0.13176 |  | MOOS2152 | 1.0015
MOOS2137 | 6.134 | 9.1784 | 9.5016 | 3.0444 | 3.3676 | 0.11968 | 0.13239 |  | MOOS2122 | 0.97945
MOOS2137T | 6.1541 | 9.222 | 9.5458 | 3.0679 | 3.3917 | 0.12019 | 0.13287
MOOS2152 | 6.1337 | 9.1835 | 9.5089 | 3.0498 | 3.3752 | 0.12036 | 0.1332
MOOS2152T | 6.1531 | 9.1907 | 9.5136 | 3.0376 | 3.3605 | 0.12058 | 0.1334
MOOS2122 | 6.1414 | 9.2025 | 9.5256 | 3.0611 | 3.3842 | 0.12253 | 0.13546
MOOS2122T | 6.0858 | 9.1144 | 9.4385 | 3.0286 | 3.3527 | 0.11985 | 0.13268
MOOS2164R | 6.1651 | 9.2186 | 9.5419 | 3.0535 | 3.3768 | 0.12042 | 0.13317
BLK | 6.0982 | 9.3387 | 9.6637 | 3.2405 | 3.5655 | 0.012381 | 0.013623
used a 1 cm cuvette zerod against GenPure water
5. Comparing nutrient data after transfer in unclean bottle
the ratios are beterrn the response in the "T" bottle against irradiated bottle (e.g. signal in MOOS2142T / MOOS2142)
Unit | Nitrate | Nitrite | Phosphate | Silicate
ID | ratio | ratio | ratio | ratio
MOOS2142 | 100.2% | 97.21% | 100% | 99.79%
MOOS2159 | 99.84% | 99.66% | 100.5% | 100.4%
MOOS2137 | 100.1% | 99.9% | 100.9% | 100.4%
MOOS2152 | 99.66% | 100.5% | 100.2% | 100.2%
MOOS2122 | 100.9% | 98.64% | 99.19% | 97.95%
No loss of nutrient in this preliminary stability study
Screening of nutrient on the 7 boxes before reblending
Date | 2023-04-12
Measuring nutrients in the seven boxes selected for reblending
Box | Unit
N | N
2 | MOOS0062
7 | MOOS0357
24 | MOOS1299
28 | MOOS1550
29 | MOOS1570
38 | MOOS2080
54 | MOOS2984
REF (Box 39) | MOOS2156
1. Determination of nitrate by ID GC-MS
Unit | Vial | Sample | IS | mA | mB | Area | Area | rAB | Hom
ID | g | g | g | g | g | 46 | 47
2 | 11.538 | 13.566 | 13.7665 | 2.028 | 0.2005 | 6655.2225 | 7221.961 | 0.92153 | 0.091107
7 | 11.6255 | 13.6541 | 13.8533 | 2.0286 | 0.1992 | 6466.859 | 6924.9246 | 0.93385 | 0.0917
24 | 11.5264 | 13.5545 | 13.7536 | 2.0281 | 0.1991 | 6247.0862 | 7016.358 | 0.89036 | 0.087407
28 | 11.4704 | 13.4819 | 13.6811 | 2.0115 | 0.1992 | 5911.3167 | 6441.1736 | 0.91774 | 0.090884
29 | 11.5359 | 13.557 | 13.7566 | 2.0211 | 0.1996 | 6045.6825 | 6851.3478 | 0.88241 | 0.087145
38 | 11.5075 | 13.507 | 13.706 | 1.9995 | 0.199 | 5524.2526 | 6188.878 | 0.89261 | 0.088837
54 | 11.9535 | 13.9574 | 14.1569 | 2.0039 | 0.1995 | 5561.402 | 6672.142 | 0.83353 | 0.082982
REF | 11.4614 | 13.4613 | 13.6608 | 1.9999 | 0.1995 | 5330.9224 | 6058.0745 | 0.87997 | 0.087781
BLK | 11.4987 | 13.5251 | 13.5251 | 2.0264 | 0 | 21.7 | 0 | #DIV/0! | #DIV/0!
IS = 15NO3 used for the SALT-1 certification | [7 empty] | Average | 0.088581
0.1 mL of 1 % sulfamic acid | [7 empty] | SD | 0.0030624
0.05 mL of triethyloxonium tetrafluoroborate solution (1 g reagent + 1 g cold water) | [7 empty] | RSD | 3.457%
HS analysis with SALT1.M method
2. Determination of nitrite by UV-vis
Unit | Vial | Sample | Chemistry | mA | mF | Area | Area
ID | g | g | g | g | g | 541 nm | norm
2 | 6.1228 | 11.1775 | 11.3794 | 5.0547 | 5.2566 | 0.082601 | 0.0859
7 | 6.1367 | 11.1999 | 11.4011 | 5.0632 | 5.2644 | 0.081693 | 0.084939
24 | 6.0974 | 11.1627 | 11.3631 | 5.0653 | 5.2657 | 0.089514 | 0.093055
28 | 6.1209 | 11.179 | 11.3802 | 5.0581 | 5.2593 | 0.091417 | 0.095053
29 | 6.1156 | 11.179 | 11.38 | 5.0634 | 5.2644 | 0.10352 | 0.10763
38 | 6.0951 | 11.1571 | 11.3577 | 5.062 | 5.2626 | 0.078948 | 0.082077
54 | 6.1164 | 11.1856 | 11.3871 | 5.0692 | 5.2707 | 0.058268 | 0.060584
REF | 6.1448 | 11.2215 | 11.4224 | 5.0767 | 5.2776 | 0.0853 | 0.088676
BLK | 6.1353 | 11.2602 | 11.4612 | 5.1249 | 5.3259 | -7.2e-05 | -7.4824e-05
used a 1 cm cuvette zerod against GenPure water | [5 empty] | Average | 0.087034
[6 empty] | SD | 0.01446
[6 empty] | RSD | 16.61%
3. Determination of phosphate by UV-vis
Unit | Vial | Sample | Chemistry | mA | mF | Area | Area
ID | g | g | g | g | g | 890 nm | norm
2 | 6.1303 | 12.2162 | 12.9995 | 6.0859 | 6.8692 | 0.27688 | 0.31252
7 | 6.1364 | 12.2177 | 13.0048 | 6.0813 | 6.8684 | 0.27803 | 0.31401
24 | 6.1281 | 12.2223 | 13.0085 | 6.0942 | 6.8804 | 0.272 | 0.30709
28 | 6.1475 | 12.2498 | 13.0361 | 6.1023 | 6.8886 | 0.27596 | 0.31151
29 | 6.0861 | 12.201 | 12.9868 | 6.1149 | 6.9007 | 0.27634 | 0.31185
38 | 6.1378 | 12.2486 | 13.0366 | 6.1108 | 6.8988 | 0.27143 | 0.30643
54 | 6.145 | 12.245 | 13.0313 | 6.1 | 6.8863 | 0.26111 | 0.29477
REF | 6.1296 | 12.2331 | 13.0201 | 6.1035 | 6.8905 | 0.27586 | 0.31143
BLK | 6.1177 | 12.1894 | 12.9759 | 6.0717 | 6.8582 | -0.001975 | -0.0022308
used a 5 cm cuvette zerod against GenPure water | [5 empty] | Average | 0.30831
[6 empty] | SD | 0.0066018
[6 empty] | RSD | 2.141%
4. Determination of silicate by UV-vis
Unit | Vial | Sample | Chemistry | mA | mF | Area | Area
ID | g | g | g | g | g | 810 nm | norm
2 | 6.1364 | 9.1554 | 9.48 | 3.019 | 3.3436 | 0.12251 | 0.13568
7 | 6.1352 | 9.1594 | 9.4831 | 3.0242 | 3.3479 | 0.12242 | 0.13552
24 | 6.1187 | 9.1474 | 9.4709 | 3.0287 | 3.3522 | 0.12385 | 0.13708
28 | 6.1346 | 9.1652 | 9.4886 | 3.0306 | 3.354 | 0.12466 | 0.13796
29 | 6.1355 | 9.1555 | 9.4786 | 3.02 | 3.3431 | 0.12232 | 0.13541
38 | 6.1344 | 9.1529 | 9.4763 | 3.0185 | 3.3419 | 0.1232 | 0.1364
54 | 6.1441 | 9.1752 | 9.4995 | 3.0311 | 3.3554 | 0.12599 | 0.13946
REF | 6.1196 | 9.1509 | 9.4747 | 3.0313 | 3.3551 | 0.12623 | 0.13972
BLK | 6.1215 | 9.4172 | 9.7411 | 3.2957 | 3.6196 | 0.012948 | 0.014221
used a 1 cm cuvette zerod against GenPure water | [5 empty] | Average | 0.13679
[6 empty] | SD | 0.0015006
[6 empty] | RSD | 1.097%
The seven boxes are suitable for reblending
The small differences in the BOX 54 are cosistent with the data acquired during the first processing
Since the bottling was done over two days the second day some depletion in nutriet was observed expecially N and a little P
Order of reblending
54
29
2
7
38
28
24
EtOH interference on nitrite
Date | 2023-04-19
Determination of nitrite by UV-vis
Unit | Vial | Sample | EtOH | Chemistry | mA | mF | Area | Area |  | EtOH
ID | g | g | g | g | g | g | 541 nm | norm |  | µg/g
NO-1 | 6.1517 | 11.2302 | 11.2302 | 11.4327 | 5.0785 | 5.281 | 0.085172 | 0.088568 |  | 0
NO-2 | 6.1297 | 11.2235 | 11.2235 | 11.4246 | 5.0938 | 5.2949 | 0.085828 | 0.089216 |  | 0
NO-3 | 6.1273 | 11.2263 | 11.2263 | 11.4272 | 5.099 | 5.2999 | 0.085331 | 0.088693 |  | 0
Y-1 | 6.1212 | 11.2175 | 11.221 | 11.4223 | 5.0963 | 5.3011 | 0.085643 | 0.089085 |  | 686.3014
Y-2 | 6.1221 | 11.2133 | 11.2169 | 11.4174 | 5.0912 | 5.2953 | 0.085479 | 0.088906 |  | 706.6028
Y-3 | 6.0888 | 11.1768 | 11.1805 | 11.3813 | 5.088 | 5.2925 | 0.085586 | 0.089026 |  | 726.6728
used a 1 cm cuvette zerod against GenPure water | [6 empty] | Average | 0.088826
using MOOS 2126 | [6 empty] | SD | 0.00034395
[7 empty] | RSD | 0.3872%
[7 empty] | Average | 0.089005
[7 empty] | SD | 9.1193e-05
ppm level of EtOH do not interfere with nitrite assay | [6 empty] | RSD | 0.1025%
Level of EtOH in seawater 2-33 nM ~ 0.45 ppb
https://doi.org/10.1029/2010GL045534
Preparation of the 15NO3- internal standard with exact-matching
Date | 2024-01-02
15NO3- standard solution E PAGLIANO 31-03-2017
CIL K(15NO3) 15N 99 % NLM-765-1
Lot#: I1-12511L
K15NO3 | 0.1985 | g
water | 127.6409 | g
w(15NO3-) | ~ 959 | μg/g
Internal standard solution for SALT-1 project
Dilution from previous solution
~ 959 | 0.631 | mL
+ water | 50 | mL
w(15NO3-) | ~ 11.76 | μg/g
Box | Unit | Unit | Nitrate
N | N | N (blended) | rAB
2 | MOOS0062 | 49 | 0.92153
7 | MOOS0357 | 49 | 0.93385
24 | MOOS1299 | 49 | 0.89036
28 | MOOS1550 | 48 | 0.91774
29 | MOOS1570 | 49 | 0.88241
38 | MOOS2080 | 49 | 0.89261
54 | MOOS2984 | 37 | 0.83353
weighted average |  |  | 0.89821
This ratio was obtained by using the SALT-1 internal standard
Date | 2024-01-09
Preparation of the solution for the MOOS-4 certification 50 mL
Dilution from 15NO3- standard solution E PAGLIANO 31-03-2017
Bottle | 11.8034 | g
~ 959 | 12.3735 | g | 0.56677 | mL
+ water | 62.4846 | g | 50 | mL
w(15NO3-) | ~ 10.79
Preparation of the 15NO2- and 15NO3- internal standard with exact-matching for simultaneous determination
Date | 2024-01-25
15NO3- standard solution E PAGLIANO 31-03-2017
CIL K(15NO3) 15N 99 % NLM-765-1
Lot#: I1-12511L
K15NO3 | 0.1985 | g
water | 127.6409 | g
w(15NO3-) | ~ 959 | μg/g
15NO2- standard solution E PAGLIANO 2019/04/04
CIL NaNO2 15N 99 % NLM-658-1
Na15NO2 | 0.188 | g
Sol pH 11 | 124.7191 | g
w(15NO2-) | 1012.103 | μg/g
Preparation of the solution for the MOOS-4 certification 100 mL
Bottle | 21.8876 | g
~ 959 15NO3 | 23.0127 | g | 0.5667 x2 | mL
~1012 15NO2 | 23.0937 | g | 0.081 | mL
Sol pH 11 | 123.1084 | g | 100 | mL | Gen pure water pH corrected with NaOH 10 N
w(15NO3-) | ~ 10.66
w(15NO2-) | ~ 0.8099
Standardization NO2- Sigma Aldrich (verified 2024-04-22)
1.a Preparation of a 0.75 N solution of H2SO4
Date | 2024-01-31
97.96 mL of GenPure water + 2.04 mL of concentrated sulfuric acid
1.b Preparation of a 0.1 N solution of KMnO4
Date | 2024-01-31
0.06 g of KMnO4 in 20 g soluitiom with GenPure water
2. Preparation of internal standard 15NO3-
Date | 2024-02-05
Dilution from 15NO3- standard solution E PAGLIANO 31-03-2017
Bottle | 8.2029 | g
~ 959 | 11.1911 | g | 3 | mL
+ water | 37.9785 | g | 28.7 | mL
w(15NO3-) | ~ 96.24
3. Preparation of sample dilution from the Sigma Aldrich standard
Date | 2024-02-06
| load | unload | mass
| g | g | g
Bottle | 7.3195 | -0.0001 | 7.3195
+ standard | 8.0674 | -0.0001 | 8.0674
+ water | 17.2278 | 0 | 17.2278
w(NO2-) μg/g |  |  | 75.558
ID: NO2-SAMPLE
| Origin |  | m0 = vial |  | m1 = m0 + NO2 standard |  | m2 = m1 + DIW | [7 empty] | Final
ID solution | w(NO2) | u[w(NO2)] | load | unload | load | unload | load | unload | m0 | m1 | m2 | u(m0) | u(m1) | u(m2) | w(NO2) | u[w(NO2)] | Target | Traciability | Prep. Date | Exp. Date | t | Rel Hum | Analyst | Balance
unique | µg/g | µg/g | g | g | g | g | g | g | g | g | g | g | g | g | µg/g | µg/g | % |  | yyyymmdd | yyyymmdd | °C | % | initials | model
NO2-SAMPLE | 1001 | 2 | 7.3195 | -0.0001 | 8.0674 | -0.0001 | 17.2278 | 0 | 7.3195 | 8.0674 | 17.2278 | 0.0002 | 0.0002 | 0.0002 | 75.558 | 0.15366 | N/A | SIGMA | 20240206 | 20240207 | 21.8 | 16.5 | EP | MS204S
In a 50 mL bottle
Product no.: 67276
Lot no.: BCCK5026
Nitrite 1001 mg kg-1 ± 4 mg kg-1 1000 mg L-1 ± 4 mg L-1
4. Preparation of the sample dilution from the Sigma Aldrich standard with nitrite elimination by sulfamic acid
Date | 2024-02-06
| load | unload | mass
| g | g | g
Bottle | 7.3131 | 0.0001 | 7.3131
+ standard | 8.0589 | 0 | 8.0589
+ 1% sulfamic | 11.059 | 0.0002 | 11.0589
+ water | 17.4444 | 0 | 17.4444
ID: NO2-SAMPLE-SULF
5. Preparation of blank sulfamic acid
Date | 2024-02-06
| load | unload | mass
| g | g | g
Bottle | 7.3127 | -0.0001 | 7.3128
+ 1% sulfamic | 10.3142 | 0.0001 | 10.3141
+ water | 17.208 | 0.0001 | 17.2079
ID: BLANK-SULF
6. Preparation of the primary standard from NIST
| Origin |  | m0 = mass of vial |  | m1 = m0 + mass of standard |  | m2 = m1 + mass of DIW | [7 empty] | Final
ID solution | w(NO3) | u[w(NO3)] | load | unload | load | unload | load | unload | m0 | m1 | m2 | u(m0) | u(m1) | u(m2) | w(NO3) | u[w(NO3)] | Target | Traciability | Prep. Date | Exp. Date | t | Rel Hum | Analyst | Balance
unique | µg/g | µg/g | g | g | g | g | g | g | g | g | g | g | g | g | µg/g | µg/g | % |  | yyyymmdd | yyyymmdd | °C | % | initials | model
ST1 | 1000.6 | 0 | 7.314 | -0.0002 | 7.8157 | 0 | 17.3206 | -0.0001 | 7.3141 | 7.8157 | 17.3207 | 0.0002 | 0.0002 | 0.0002 | 50.1572 | 0.028318 | N/A | NIST | 20240205 | 20240405 | 21.8 | 17.2 | EP | MS204S
ST2 | 1000.6 | 0 | 7.3093 | -0.0003 | 8.3121 | 0 | 17.1971 | -0.0001 | 7.3095 | 8.3121 | 17.1972 | 0.0002 | 0.0002 | 0.0002 | 101.4646 | 0.028769 | N/A | NIST | 20240205 | 20240405 | 21.8 | 17.2 | EP | MS204S
ST3 | 1000.6 | 0 | 7.3148 | 0 | 8.3202 | 0.0001 | 12.1916 | 0 | 7.3148 | 8.3201 | 12.1916 | 0.0002 | 0.0002 | 0.0002 | 206.2732 | 0.059253 | N/A | NIST | 20240205 | 20240405 | 21.8 | 17.2 | EP | MS204S
7. Preparation of the sample blends
|  | m0 = mass of vial |  |  | m1 = m0 + mass of sample/ST |  |  | m2 = m1 + IntST
|  | load | unload | mass | load | unload | mass | load | unload | mass | mA | mB
ID | From | g | g | g | g | g | g | g | g | g | g | g
AB-1 | ST1 | 2.6949 | -0.0001 | 2.6949 | 3.7022 | 0 | 3.7022 | 4.7108 | 0 | 4.7108 | 1.0072 | 1.0086
AB-2A | ST2 | 2.7043 | 0 | 2.7043 | 3.7098 | 0.0001 | 3.7098 | 4.7071 | 0 | 4.7071 | 1.0055 | 0.99735
AB-2B | ST2 | 2.6849 | 0 | 2.6849 | 3.6787 | 0 | 3.6787 | 4.6734 | 0 | 4.6734 | 0.9938 | 0.9947
AB-2C | ST2 | 2.6567 | 0.0001 | 2.6566 | 3.6488 | -0.0002 | 3.6489 | 4.6432 | 0.0001 | 4.6432 | 0.99225 | 0.99425
AB-3 | ST3 | 2.6991 | -0.0002 | 2.6992 | 3.7051 | 0.0001 | 3.705 | 4.6993 | 0 | 4.6993 | 1.0058 | 0.99425
BLK1 | GenPure | 2.7241 | 0 | 2.7241 | 3.73 | 0 | 3.73 | 4.7385 | -0.0003 | 4.7386 | 1.0059 | 1.0086
BLK2 | GenPure | 2.7248 | 0.0001 | 2.7248 | 3.7317 | 0 | 3.7317 | 4.7314 | -0.0002 | 4.7315 | 1.0069 | 0.9998
BLK3 | GenPure | 2.7113 | 0.0001 | 2.7113 | 3.718 | 0 | 3.718 | 4.7137 | 0 | 4.7137 | 1.0067 | 0.9957
S1 | Sample | 2.678 | -0.0001 | 2.678 | 3.6845 | 0.0001 | 3.6845 | 4.6802 | 0 | 4.6802 | 1.0064 | 0.99575
S2 | Sample | 2.6865 | 0 | 2.6865 | 3.6934 | 0 | 3.6934 | 4.6896 | 0 | 4.6896 | 1.0069 | 0.9962
S3 | Sample | 2.7041 | -0.0001 | 2.7041 | 3.7124 | 0 | 3.7124 | 4.709 | 0.0001 | 4.7089 | 1.0083 | 0.99655
SULF-BLK1 | BLK sulf | 2.7041 | -0.0001 | 2.7041 | 3.7142 | -0.0001 | 3.7142 | 4.7226 | -0.0001 | 4.7226 | 1.0101 | 1.0084
SULF-BLK2 | BLK sulf | 2.6843 | 0 | 2.6843 | 3.6937 | 0 | 3.6937 | 4.6946 | 0 | 4.6946 | 1.0094 | 1.0009
SULF-BLK3 | BLK sulf | 2.7144 | 0 | 2.7144 | 3.7236 | -0.0002 | 3.7237 | 4.7238 | 0 | 4.7238 | 1.0093 | 1.0001
SULF-S1 | Sample sulf | 2.6796 | -0.0001 | 2.6797 | 3.6892 | 0 | 3.6892 | 4.6887 | 0.0001 | 4.6886 | 1.0095 | 0.99945
SULF-S2 | Sample sulf | 2.6737 | -0.0001 | 2.6738 | 3.6833 | 0 | 3.6833 | 4.6836 | 0.0002 | 4.6835 | 1.0095 | 1.0002
SULF-S3 | Sample sulf | 2.732 | 0.0001 | 2.732 | 3.7413 | -0.0001 | 3.7413 | 4.742 | -0.0001 | 4.742 | 1.0094 | 1.0007
8. Oxidation of the sample and dilution
Prepare 11 vials with 1 mL of 0.75 N sulfuric acid + 0.1 mL of 0.1 N potassium permanganate
Prepare 6 vials with 1.1 mL of water
Incubate at 40 °C 15 min
Add 0.2 mL fo the previous blend (6 SULF blends just water dilution, no KMnO4)
Immersion of the pipette tip
Mix and wait 10 minutes
9. Analysis by GC-MS
In 17 combiPAL vial 10 mL transfer 1.5 mL of water + 0.5 mL of the digested sample + 50 uL of oxonium solution in water
10. Difference between Isotopic composition of sample (NO2 Sigma) and standard (NO3 NIST)
Sample | rAB/(V/V) |  | t-Test: Paired Two Sample for Means
Sigma_1 | 85.9391
Sigma_2 | 83.5038 |  |  | Variable 1 | Variable 2
Sigma_3 | 84.722 |  | Mean | 84.0116 | 83.218
Sigma_4 | 84.6215 |  | Variance | 1.4425 | 0.29766
Sigma_5 | 82.2031 |  | Observations | 7 | 7
Sigma_6 | 83.7705 |  | Pearson Correlation | 0.22331
Sigma_7 | 83.3213 |  | Hypothesized Mean Difference | 0
NIST_1 | 83.8391 |  | df | 6
NIST_2 | 84.1322 |  | t Stat | 1.7451
NIST_3 | 82.6831 |  | P(T<=t) one-tail | 0.065793
NIST_4 | 83.0453 |  | t Critical one-tail | 1.9432
NIST_5 | 83.0079 |  | P(T<=t) two-tail | 0.13159
NIST_6 | 82.8263 |  | t Critical two-tail | 2.4469
NIST_7 | 82.9923 |  | No significantive difference alfa = 0.05
Same oxidation process as reported above
11. Quantitative conversion of nitrite by permanganate
File | NO2 (A at 45) | average | stdev
AB2A_I.D | 22.3117
AB2B_I.D | 13.7652 | 17.2137 | 4.5057
AB2C_I.D | 15.5642
S1_I.D | 22.8758
S2_I.D | 23.6908 | 22.0187 | 2.228
S3_I.D | 19.4895
WATER_I.D | 21.3963
Difference |  | 4.805 | 5.0264
In practice there is no difference between the nitrite signal observed in the sample and in the nitrate standard
Therefore the nitrite conversion by permanganate was quantitative
12. No endogenous nitrate found in the Sigma nitrite standard
ID | rAB | u(rAB) | (mB/mA)*rAB | average | stdev
SULF_BLK1 | 0.017604 | 0.00012888 | 0.017575
SULF_BLK2 | 0.017121 | 0.0001114 | 0.016977 | 0.016678 | 0.001078
SULF_BLK3 | 0.015624 | 0.00017707 | 0.015482
SULF_S1 | 0.013885 | 0.00010776 | 0.013746
SULF_S2 | 0.01728 | 0.00022307 | 0.01712 | 0.015947 | 0.0019077
SULF_S3 | 0.017123 | 0.00012189 | 0.016976
Difference |  |  |  | -0.00073049 | 0.0021912
There is no detectable nitrate in the Sigma Aldrich standard BCCK5026
12. Quantitation of nitrate in the Sigma Aldrich standard BCCK5026 after KMnO4 conversion (exact-matching, blank-matching)
Sample ID | rA*B | u(rA*B) | Date | wA* | u(wA*) | mA* | mB | wA*(mA*/mB) | u(wA*(mA*/mB))
AB1 | 0.52821 | 0.00017515 | 2024-02-06 14:26 | 50.1572 | 0.028318 | 1.0072 | 1.0086 | 50.0901 | 0.034568
AB2A | 1.0496 | 0.00099849 | 2024-02-06 14:39 | 101.4646 | 0.028769 | 1.0055 | 0.99735 | 102.2887 | 0.050106
AB2B | 1.044 | 0.0007231 | 2024-02-06 14:53 | 101.4646 | 0.028769 | 0.9938 | 0.9947 | 101.3728 | 0.049895
AB2C | 1.0389 | 0.00043831 | 2024-02-06 15:07 | 101.4646 | 0.028769 | 0.99225 | 0.99425 | 101.2605 | 0.049873
AB3 | 2.1057 | 0.00096165 | 2024-02-06 15:21 | 206.2732 | 0.059253 | 1.0058 | 0.99425 | 208.6798 | 0.10277
[8 empty] | Gauss's method (GUM's Linear Approximation) |  |  |  | Monte Carlo method |  |  | Ratio wA | Ratio u(wA)
Sample ID | rAB | u(rAB) | Date | mA | mB | m = mB/mA | u(m) | wA | u(wA) | uR(wA) | R^2 | wA | u(wA) | uR(wA) | GS/MT | GS/MT
S1 | 1.0555 | 0.00040941 | 2024-02-06 15:35 | 1.0064 | 0.99575 | 0.98942 | 0.00039536 | 101.6901 | 0.13837 | 0.1361% | 0.99999 | 101.6883 | 0.078634 | 0.07733% | 1 | 1.7596
S2 | 1.0613 | 0.00041885 | 2024-02-06 15:49 | 1.0069 | 0.9962 | 0.98937 | 0.00039515 | 102.2581 | 0.13937 | 0.1363% | 0.99999 | 102.2631 | 0.077884 | 0.07616% | 0.99995 | 1.7895
S3 | 1.0647 | 0.00088072 | 2024-02-06 16:03 | 1.0083 | 0.99655 | 0.9884 | 0.00039443 | 102.4909 | 0.15936 | 0.1555% | 0.99999 | 102.4922 | 0.11052 | 0.1078% | 0.99999 | 1.4419
Consensus value | [7 empty] | 102.1464 | 0.27566 | 0.2699%
The 95 % coverage interval (DerSimonian-Laird, NIST consensus builder, 2024-04-22) is 101.7 to 102.6
13. Including the uncertainty from the primary standard
uR (analysis) | 0.0026986
uR (standard) | 0.00091548
uR | 0.0028497
14. Calculation of the nitrite mass fraction in the Sigma Aldrich standard BCCK5026
Solution | w(NO2-) | uw(NO2-)
ID | μg/g | μg/g
NO2-SAMPLE | 75.7892 | 0.21598
BCCK5026 | 1004.0619 | 2.8865
15. Consistency IDMS results vs value reported by Sigma Aldrich
| CRM |  | Measurement |  |  | Comparability |  |  |  |  | Ratio
Sample | μ | unc | x | s | n | IRMM | ISO | Wald | t-test | Mood, B=2 | x/μ
NO2-SAMPLE | 1001 | 2 | 1004.0619 | 2.8865 | 3 | TRUE | TRUE | TRUE | TRUE | TRUE | 1.0031
Development of novel method for the determination of nitrite by GC-MS following derivatization with triethyloxonium
1. Reactivity: NO2- + Et3O+ → EtONO + Et2O
2. Full scan NCI (methane) mass spectrum of the EtONO derivative (from nitrite of natural isotopic composition)
Date | 20240122
#"- Scan (2.283 min) NO2.D Subtract "
#Point | X(Thompsons) | Y(Counts) | Y(norm)
0 | 14.6 | 0.83696 | 0.0007022%
1 | 15.2 | 4.913 | 0.004122%
2 | 19 | 22.2065 | 0.01863%
3 | 19.7 | 0.81522 | 0.0006839%
4 | 20.4 | 0.81522 | 0.0006839%
5 | 20.9 | 1.8478 | 0.00155%
6 | 23 | 0.91304 | 0.000766%
7 | 25.1 | 194.3044 | 0.163%
8 | 28.4 | 12.75 | 0.0107%
9 | 29.3 | 1.7391 | 0.001459%
10 | 30.2 | 115.6739 | 0.09705%
11 | 31.1 | 8447.8809 | 7.088%
12 | 32 | 103.6304 | 0.08694%
13 | 33 | 120.5109 | 0.1011%
14 | 34.3 | 0.93478 | 0.0007843%
15 | 35.3 | 0.84783 | 0.0007113%
16 | 35.8 | 0.83696 | 0.0007022%
17 | 36.3 | 0.82609 | 0.0006931%
18 | 37.2 | 0.48913 | 0.0004104%
19 | 38.1 | 7.163 | 0.00601%
20 | 39 | 8.5543 | 0.007177%
21 | 40.5 | 4.6848 | 0.00393%
22 | 40.9 | 9.9022 | 0.008308%
23 | 43.1 | 33287.4141 | 27.93%
24 | 44.2 | 1148.674 | 0.9637%
25 | 45.1 | 119193.0312 | 100%
26 | 46.1 | 4156.2935 | 3.487%
27 | 47.2 | 253.0544 | 0.2123%
28 | 47.9 | 20.7391 | 0.0174%
29 | 48.4 | 0.69565 | 0.0005836%
30 | 49.3 | 0.72826 | 0.000611%
31 | 50.2 | 0.73913 | 0.0006201%
32 | 51 | 0.61957 | 0.0005198%
33 | 54.1 | 4.7826 | 0.004012%
34 | 55.2 | 16.5543 | 0.01389%
35 | 57 | 211.7717 | 0.1777%
36 | 58.4 | 5.8587 | 0.004915%
37 | 58.9 | 34.7609 | 0.02916%
38 | 59.3 | 29.7391 | 0.02495%
39 | 60 | 11.5435 | 0.009685%
40 | 60.8 | 2.5543 | 0.002143%
41 | 61.4 | 0.80435 | 0.0006748%
42 | 61.9 | 7.7935 | 0.006539%
43 | 63.1 | 313.0109 | 0.2626%
44 | 64.3 | 1.75 | 0.001468%
45 | 64.8 | 3.7065 | 0.00311%
46 | 65.3 | 19.75 | 0.01657%
47 | 67.2 | 15.0435 | 0.01262%
48 | 69.1 | 25.1957 | 0.02114%
49 | 70.1 | 14.6739 | 0.01231%
50 | 71.2 | 56.1739 | 0.04713%
51 | 72 | 5.087 | 0.004268%
52 | 73 | 4.5543 | 0.003821%
53 | 74.1 | 4.3587 | 0.003657%
54 | 74.9 | 147.4891 | 0.1237%
55 | 76.2 | 6.913 | 0.0058%
56 | 77.2 | 23.7935 | 0.01996%
57 | 77.9 | 5.7174 | 0.004797%
58 | 79.2 | 17.4674 | 0.01465%
59 | 80.2 | 2.5543 | 0.002143%
60 | 80.9 | 44.2717 | 0.03714%
61 | 81.9 | 3.3152 | 0.002781%
62 | 82.7 | 0.33696 | 0.0002827%
63 | 83.3 | 5.6304 | 0.004724%
64 | 83.8 | 11.087 | 0.009302%
65 | 84.9 | 21.3152 | 0.01788%
66 | 85.4 | 26.7609 | 0.02245%
67 | 86.7 | 3.3478 | 0.002809%
68 | 87.8 | 1.6413 | 0.001377%
69 | 88.9 | 24.5435 | 0.02059%
70 | 89.6 | 9.3587 | 0.007852%
71 | 91 | 117.4891 | 0.09857%
72 | 91.7 | 26.5217 | 0.02225%
73 | 92.9 | 23.3696 | 0.01961%
74 | 93.8 | 1.337 | 0.001122%
75 | 95.1 | 16.6087 | 0.01393%
76 | 97 | 17.3261 | 0.01454%
77 | 98 | 4.2283 | 0.003547%
78 | 98.8 | 43.4674 | 0.03647%
79 | 100 | 6.337 | 0.005317%
80 | 100.8 | 2.5 | 0.002097%
81 | 102 | 8.5109 | 0.00714%
82 | 102.5 | 0.5 | 0.0004195%
83 | 103.9 | 1.2174 | 0.001021%
84 | 105.1 | 7.8152 | 0.006557%
85 | 106.2 | 9.5109 | 0.007979%
86 | 107.1 | 15.6196 | 0.0131%
87 | 108.9 | 23.1739 | 0.01944%
88 | 109.6 | 8.4565 | 0.007095%
89 | 111 | 24.5326 | 0.02058%
90 | 112.1 | 9.6848 | 0.008125%
91 | 112.8 | 20.3913 | 0.01711%
92 | 114.5 | 0.70652 | 0.0005928%
93 | 115 | 12.0978 | 0.01015%
94 | 116.2 | 4.8261 | 0.004049%
95 | 117.1 | 53.413 | 0.04481%
96 | 119.5 | 1.8043 | 0.001514%
3. Full scan NCI (methane) mass spectrum of the EtONO derivative (from 15N-nitrite)
Date | 20240122
#"- Scan (2.265-2.268 min | 2 Scans) 15NO2.D Subtract "
#Point | X(Thompsons) | Y(Counts) | Y(norm)
0 | 16.8 | 30.5 | 0.04492%
1 | 17.8 | 0.375 | 0.0005522%
2 | 18.2 | 0.3 | 0.0004418%
3 | 18.5 | 0.375 | 0.0005522%
4 | 19.1 | 3.875 | 0.005706%
5 | 19.4 | 11.975 | 0.01763%
6 | 20 | 0.45 | 0.0006627%
7 | 21.8 | 1.475 | 0.002172%
8 | 22.5 | 0.425 | 0.0006259%
9 | 23.3 | 0.35 | 0.0005154%
10 | 25 | 85.325 | 0.1257%
11 | 25.2 | 49.675 | 0.07315%
12 | 25.9 | 5.275 | 0.007768%
13 | 26.7 | 0.2 | 0.0002945%
14 | 27 | 5.65 | 0.00832%
15 | 28.8 | 2.325 | 0.003424%
16 | 29.5 | 0.275 | 0.000405%
17 | 30 | 2.7 | 0.003976%
18 | 31.2 | 62.15 | 0.09152%
19 | 32.1 | 5541.3999 | 8.161%
20 | 33 | 34.275 | 0.05047%
21 | 33.9 | 11.8 | 0.01738%
22 | 34.1 | 21.9 | 0.03225%
23 | 34.9 | 0.925 | 0.001362%
24 | 35.6 | 0.9 | 0.001325%
25 | 36.6 | 1.275 | 0.001878%
26 | 38.9 | 14.25 | 0.02099%
27 | 39.1 | 10.875 | 0.01602%
28 | 39.8 | 0.2 | 0.0002945%
29 | 41.1 | 14.775 | 0.02176%
30 | 41.4 | 4.575 | 0.006737%
31 | 41.8 | 24.4 | 0.03593%
32 | 43.1 | 19349.25 | 28.49%
33 | 44.2 | 799.775 | 1.178%
34 | 45.1 | 67905.0781 | 100%
35 | 46.1 | 1959.3 | 2.885%
36 | 47.1 | 633.2 | 0.9325%
37 | 48 | 1.65 | 0.00243%
38 | 48.8 | 0.275 | 0.000405%
39 | 49.8 | 0.375 | 0.0005522%
40 | 51 | 0.25 | 0.0003682%
41 | 51.5 | 0.4 | 0.0005891%
42 | 52.2 | 0.5 | 0.0007363%
43 | 53.1 | 0.45 | 0.0006627%
44 | 54 | 0.95 | 0.001399%
45 | 54.7 | 0.775 | 0.001141%
46 | 55.6 | 1.325 | 0.001951%
47 | 56.8 | 19.35 | 0.0285%
48 | 57.2 | 85 | 0.1252%
49 | 57.8 | 1.1 | 0.00162%
50 | 58.3 | 3.9 | 0.005743%
51 | 58.8 | 2.025 | 0.002982%
52 | 59.7 | 0.4 | 0.0005891%
53 | 60.9 | 1.375 | 0.002025%
54 | 61.4 | 12.5 | 0.01841%
55 | 62.1 | 4.725 | 0.006958%
56 | 63 | 95.725 | 0.141%
57 | 63.6 | 4.85 | 0.007142%
58 | 64.1 | 0.225 | 0.0003313%
59 | 64.3 | 2.4 | 0.003534%
60 | 65.1 | 1.375 | 0.002025%
61 | 65.3 | 2.85 | 0.004197%
62 | 65.9 | 0.275 | 0.000405%
63 | 66.3 | 0.375 | 0.0005522%
64 | 66.9 | 7.75 | 0.01141%
65 | 67.1 | 2.75 | 0.00405%
66 | 67.9 | 0.375 | 0.0005522%
67 | 69.1 | 7.35 | 0.01082%
68 | 69.7 | 0.45 | 0.0006627%
69 | 70.9 | 1.65 | 0.00243%
70 | 71.2 | 22.275 | 0.0328%
71 | 72.1 | 3.025 | 0.004455%
72 | 72.5 | 2.725 | 0.004013%
73 | 72.9 | 1.775 | 0.002614%
74 | 74 | 0.75 | 0.001104%
75 | 74.8 | 0.4 | 0.0005891%
76 | 75 | 2.175 | 0.003203%
77 | 75.9 | 106.95 | 0.1575%
78 | 76.1 | 51.975 | 0.07654%
79 | 76.9 | 2 | 0.002945%
80 | 77.9 | 0.275 | 0.000405%
81 | 78.2 | 0.975 | 0.001436%
82 | 79 | 7.175 | 0.01057%
83 | 79.7 | 0.775 | 0.001141%
84 | 80.4 | 0.55 | 0.00081%
85 | 80.9 | 18.25 | 0.02688%
86 | 81.6 | 0.75 | 0.001104%
87 | 82.9 | 2.025 | 0.002982%
88 | 85 | 1.35 | 0.001988%
89 | 85.3 | 14.85 | 0.02187%
90 | 86 | 4.475 | 0.00659%
91 | 86.9 | 2.025 | 0.002982%
92 | 88.8 | 5.65 | 0.00832%
93 | 89.7 | 0.375 | 0.0005522%
94 | 90.6 | 1.2 | 0.001767%
95 | 90.9 | 70.7 | 0.1041%
96 | 91.1 | 11.9 | 0.01752%
97 | 92.1 | 0.85 | 0.001252%
98 | 92.9 | 8.9 | 0.01311%
99 | 93.7 | 2.1 | 0.003093%
100 | 94.6 | 1.45 | 0.002135%
101 | 95.2 | 3.85 | 0.00567%
102 | 95.4 | 1 | 0.001473%
103 | 96.3 | 1.225 | 0.001804%
104 | 96.8 | 1.55 | 0.002283%
105 | 97.1 | 3.475 | 0.005117%
106 | 98.8 | 15.225 | 0.02242%
107 | 99.5 | 0.875 | 0.001289%
108 | 100.6 | 0.475 | 0.0006995%
109 | 100.9 | 2.025 | 0.002982%
110 | 104.3 | 0.375 | 0.0005522%
111 | 104.6 | 3.6 | 0.005302%
112 | 104.8 | 2.35 | 0.003461%
113 | 105.2 | 4.725 | 0.006958%
114 | 105.9 | 2.9 | 0.004271%
115 | 106.3 | 3.125 | 0.004602%
116 | 107 | 0.1 | 0.0001473%
117 | 108.1 | 10.7 | 0.01576%
118 | 109.8 | 13.1 | 0.01929%
119 | 110 | 4.525 | 0.006664%
120 | 111.3 | 8.4 | 0.01237%
121 | 112.9 | 0.95 | 0.001399%
122 | 113.2 | 8.525 | 0.01255%
123 | 113.7 | 0.4 | 0.0005891%
124 | 116.2 | 0.1 | 0.0001473%
125 | 117.3 | 3.425 | 0.005044%
126 | 117.9 | 10.325 | 0.01521%
127 | 119.1 | 5.45 | 0.008026%
128 | 119.7 | 3.625 | 0.005338%
4. The PCI mode is not sensitive enough for the determination of nitrite and nitrate in seawater
Date | 20240123
5. The former derivatization method witn NH3 ammonia buffer and N18O2- internal standard is not promissing for the determination of low level nitrite and nitrate in seawater
Date | 20240123
6. Optimal injection volume is 0.75 mL of headspace using a 1 mm liner
Date | 20240124
7. Blank and reproduciability nitrite m/z 31 MOOS-4 candidate (n = 3)
Date | 20240124
Area | S/N
1863.38 | 452
1859.78 | 442
1978.84 | 314.5
3.563% | 442
8. Experiment in PCI to access the separation of CH3OCH3 and EtONO
On the m/z 31 the interference CH3OCH3 is separated. When run in NCI its area is negible
Validation figures nitrite and nitrate method and preliminary data on MOOS-4 and SEAN-1
Date | 2024-02-14
1. Preparation of primary standards
mass
g
Bottle | 7.3149
+ NO2 | 8.2867
+ water | 17.2585
w(NO2−) | 7.3844 | μg/g
ID | 1:10 NO2
mass
g
Bottle | 8.1617
+ 1:10 NO2 | 8.564
+ NO3 | 8.8672
+ water | 37.3061
w(NO2−) | 0.10193 | μg/g
w(NO3−) | 1.0556 | μg/g
ID | MIX, CAL1
mass
g
Bottle | 7.2913
+ mix | 14.7717
+ water | 17.2906
w(NO2−) | 0.076254 | μg/g
w(NO3−) | 0.78967 | μg/g
ID | CAL2
mass
g
Bottle | 7.3514
+ mix | 12.3234
+ water | 17.3636
w(NO2−) | 0.050619 | μg/g
w(NO3−) | 0.52419 | μg/g
ID | CAL3
mass
g
Bottle | 7.3282
+ mix | 9.8146
+ water | 17.2686
w(NO2−) | 0.025496 | μg/g
w(NO3−) | 0.26403 | μg/g
ID | CAL4
mass
g
Bottle | 7.3217
+ mix | 8.3253
+ water | 17.3638
w(NO2−) | 0.010187 | μg/g
w(NO3−) | 0.10549 | μg/g
ID | CAL5
mass
g
Bottle | 7.3043
+ mix | 7.4061
+ water | 17.4647
w(NO2−) | 0.0010213 | μg/g
w(NO3−) | 0.010576 | μg/g
ID | LOD
2. Preparation of sample blends
SEAN1797 | 72.416 | mass of the unit in g
MOOS2157 | 72.9111 | mass of the unit in g
For the preparation of the "blank matched" experiment, 2 mL of water was evaporated oven overnight in the 110 °C oven in B5
Thereafter, 2 mL of water was transferred in the vial and derivatized with 50 μL of Et3OBF4 solution
For the preparation of the "blank" experiment, 2 mL of water was transferred in the vial and derivatized with 50 μL of Et3OBF4 solution
For the LOD experiment 2 mL of LOD solution in water was transferred in the vial and derivatized
Name | Vial | Sample | IS | mA | mB
CAL1 | 11.587 | 13.5985 | 13.7985 | 2.0115 | 0.2
CAL2 | 11.7432 | 13.7561 | 13.9553 | 2.0129 | 0.1992
CAL3 | 11.7925 | 13.8047 | 14.0037 | 2.0122 | 0.199
CAL4 | 11.7523 | 13.7669 | 13.9664 | 2.0146 | 0.1995
CAL5 | 11.7535 | 13.7682 | 13.9674 | 2.0147 | 0.1992
SEAN1797 | 11.7923 | 13.8481 | 14.0473 | 2.0558 | 0.1992
MOOS2157 | 11.6889 | 13.7468 | 13.9461 | 2.0579 | 0.1993
3. Integration results
Sample | Nitrite |  | Nitrate
Data File | Area 31 | Area 32 | Area 46 | Area 47
VOID.D
BLK_M1.D | 33.1798 |  | 63.3455
BLK_1.D | 4.0593 |  | 31.2491
BLK_M2.D | 26.4855 |  | 39.108
BLK_2.D | 3.9274 |  | 30.6688
BLK_M3.D | 26.03 |  | 41.947
BLK_3.D | 3.6349 |  | 29.6838
LOD_1.D | 44.2184 |  | 255.805
LOD_2.D | 46.1822 |  | 282.1934
LOD_3.D | 45.3955 |  | 284.542 | 4.11
CAL_5.D | 420.0154 | 3048.9851 | 2196.9592 | 18759.6954
CAL_4.D | 983.7853 | 3057.1758 | 4876.6445 | 18735.4776
CAL_3.D | 1962.4017 | 3148.3171 | 9719.639 | 19328.18
CAL_2.D | 2713.2296 | 2931.6985 | 12998.8232 | 17489.1105
CAL_1.D | 3559.5854 | 2906.59 | 17380.055 | 17406.2999
SEAN1797.D | 878.6904 | 2375.5602 | 4811.443 | 9018.527
MOOS2157.D | 2229.7591 | 2308.6759 | 8910.6326 | 8814.154
4. Evaluation of signal suppression due to the high-chloride matrix
Nitrite | 22.41%
Nitrate | 51.39%
Isotope dilution quantitation deals with these effects
5. Evaluation of blanks
| Nitrite | Nitrate |  |  | Nitrite | Nitrate
BLK_M1.D | 33.1798 | 63.3455 |  | BLK_1.D | 4.0593 | 31.2491
BLK_M2.D | 26.4855 | 39.108 |  | BLK_2.D | 3.9274 | 30.6688
BLK_M3.D | 26.03 | 41.947 |  | BLK_3.D | 3.6349 | 29.6838
The blank increase in the "blank matching" experiment is not justified. Formation of nitrite could be due to higher NOx present at high temperature.
Nitrite blank is very low and no blank matching is needed.
Relevance of the reagent blank on determination of nitrite in MOOS-4
Between | 0.1419% | 0.1737%
Relevance of the reagent blank on determination of nitrate in MOOS-4
Between | 0.3568% | 0.5402%
Difference between blank matching and non blank matching design for nitrate
BLK_M | 48.1335 | 7.65
BLK | 30.5339 | 0.45688
Difference | 17.5996 | 7.6637
Between | 0.1305% | 0.1975%
6. Evaluation of the limit of instrumental detection for nitrite
3 blanks and 3 standards LOD
w(NO2−) | 0.0010213
Signal | 41.3915
S/N | 13.442 | Area, peak-to-peak from drift
LOD water | 0.22793 | ng/g
LOD water | 4.9544 | nmol/L
LOD seawater | 0.29376 | ng/g
LOD seawater | 6.3853 | nmol/L
7. Evaluation of the instrumental limit of detection for nitrate
3 blanks and 3 standards LOD
w(NO3−) | 0.010576
Signal | 243.6462
S/N | 218.0712 | Area, peak-to-peak from drift
LOD water | 0.1455 | ng/g
LOD water | 2.3465 | nmol/L
LOD seawater | 0.29933 | ng/g
LOD seawater | 4.8275 | nmol/L
8. Evaluation of the calibration plot for nitrite
| w(NO2−) | Signal 31 | wA*(mA*/mB) | rA*B
CAL_5 | 0.010187 | 420.0154 | 0.10303 | 0.13776
CAL_4 | 0.025496 | 983.7853 | 0.25747 | 0.3218
CAL_3 | 0.050619 | 1962.4017 | 0.51183 | 0.62332
CAL_2 | 0.076254 | 2713.2296 | 0.77054 | 0.92548
CAL_1 | 0.10193 | 3559.5854 | 1.0252 | 1.2247
9. Evaluation of the calibration plot for nitrate
| w(NO3−) | Signal 46 | wA*(mA*/mB) | rA*B
CAL_5 | 0.10549 | 2196.9592 | 1.067 | 0.11711
CAL_4 | 0.26403 | 4876.6445 | 2.6663 | 0.26029
CAL_3 | 0.52419 | 9719.639 | 5.3004 | 0.50287
CAL_2 | 0.78967 | 12998.8232 | 7.9795 | 0.74325
CAL_1 | 1.0556 | 17380.055 | 10.6164 | 0.99849
10. Preliminar evaluation of nitrite in MOOS-4 and SEAN-1
Sample | mA | mB | m = mB/mA | rAB | u[rAB] | wA | u[wA] | uR[wA]
SEAN1797.D | 2.0558 | 0.1992 | 0.096897 | 0.36989 | 0.0018494 | 0.028893 | 0.0001628 | 0.5635%
MOOS2157.D | 2.0579 | 0.1993 | 0.096846 | 0.96582 | 0.0048291 | 0.077862 | 0.00041086 | 0.5277%
11. Preliminar evaluation of nitrate in MOOS-4 and SEAN-1
Sample | mA | mB | m = mB/mA | rAB | u[rAB] | wA | u[wA] | uR[wA]
SEAN1797.D | 2.0558 | 0.1992 | 0.096897 | 0.53351 | 0.0026675 | 0.54972 | 0.0042419 | 0.7716%
MOOS2157.D | 2.0579 | 0.1993 | 0.096846 | 1.0109 | 0.0050547 | 1.0422 | 0.0067067 | 0.6435%
Purification of oxonium reagent for nitrate measurement
Date | 2024-03-05
1 g of Et3OBF4 was transferred into a teflon vial
It was capped with a cap with two openings which allowed introduction of Argon and evaquation of the overpressure into a bottle containing water
In a fumehood the headspace of the vial was flushed with Ar for 5 minutes
Then the vial under was placed in a hot block at 50 °C and kept under Argon flushing
Then the vial was removed from the hot block and kept under Argon flushing for another 10 minutes
The reagent was then capped with a solid cap and kept at -20 °C
The PAL3 autosampler vials where cleanned by rinsing three times in water and then stored full of water for one week (dryed the day of analysis under N2)
[The w1d1 nitrate measurement was performed using vials cleaned in HCl 10 %. It is likely that the higher blank was due to the use of HCl for cleaning. Use only water for cleaning]
The purification of the Et3OBF4 reagent improved the blank by a factor of 10
Study matrix effects MOOS-4 by photometry (nitrite, phosphate, silicate; verified 2024-05-03)
Date | 2024-05-02
Analyte | Value | SD | RSD | Unit
Nitrite | 0.075565 | 0.0002452 | 0.3245% | mg/kg
Phosphate | 0.26357 | 0.00096944 | 0.3678% | mg/kg
Silicate | 0.42171 | 0.0052795 | 1.252% | mg/kg
2. Target mass fraction in LNSW
Analyte | Value | Unit
Nitrite | 0.075565 | mg/kg
Phosphate | 0.26357 | mg/kg
Silicate | 0.37953 | mg/kg
signal SiO2 LNWS |  | 0.012451
signal SiO2 MOOS-4 |  | 0.12448
3. Preparation mix dilution nitrite, phosphate and silicate in LNSW (to be used also as spike)
| Origin |  | Origin |  | Origin |  | m0 = vial |  | m1 = m0 + NO2 standard |  | m2 = m1 + PO4 standard |  | m3 = m2 + SiO2 standard |  | m4 = m3 + LNSW | [11 empty] | Final |  | Final |  | Final |  | Target | Target | Target
ID solution | w(NO2) | u[w(NO2)] | w(PO4) | u[w(PO4)] | w(SiO2) | u[w(SiO2)] | load | unload | load | unload | load | unload | load | unload | load | unload | m0 | m1 | m2 | m3 | m4 | u(m0) | u(m1) | u(m2) | u(m3) | u(m4) | w(NO2) | u[w(NO2)] | w(PO4) | u[w(PO4)] | w(SiO2) | u[w(SiO2)] | NO2 | PO4 | SiO2 | Traciability | Prep. Date | Exp. Date | t | Rel Hum | Analyst | Balance
unique | µg/g | µg/g | µg/g | µg/g | µg/g | µg/g | g | g | g | g | g | g | g | g | g | g | g | g | g | g | g | g | g | g | g | g | µg/g | µg/g | µg/g | µg/g | µg/g | µg/g | % | % | % |  | yyyymmdd | yyyymmdd | °C | % | initials | model
NPS-1 | 1004.0619 | 0 | 1000.5 | 0 | 744.4628 | 0.20097 | 11.7655 | -0.0001 | 12.1433 | -0.0001 | 13.4793 | 0 | 16.0454 | 0 | 61.7557 | 0.0002 | 11.7655 | 12.1433 | 13.4793 | 16.0454 | 61.7556 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.002 | 7.5882 | 0.0056891 | 26.7377 | 0.005762 | 38.2149 | 0.011248 | 100.4% | 101.4% | 100.7% | NIST | 20240502 | 20240503 | 22.3 | 46.1 | EP | MS204S
Silicate from Sw2d1-D 2024-03-05
In a 50 mL bottle
4. Preparation of the LNSW standard matched to the MOOS-4
| Origin |  | Origin |  | Origin |  | m0 = vial |  | m1 = m0 + standard |  | m2 = m1 + LNSW | [7 empty] | Final |  | Final |  | Final |  | Target
ID solution | w(NO2) | u[w(NO2)] | w(PO4) | u[w(PO4)] | w(SiO2) | u[w(SiO2)] | load | unload | load | unload | load | unload | m0 | m1 | m2 | u(m0) | u(m1) | u(m2) | w(NO2) | u[w(NO2)] | w(PO4) | u[w(PO4)] | w(SiO2) | u[w(SiO2)] | NO2 | PO4 | SiO2 | Traciability | Prep. Date | Exp. Date | t | Rel Hum | Analyst | Balance
unique | µg/g | µg/g | µg/g | µg/g | µg/g | µg/g | g | g | g | g | g | g | g | g | g | g | g | g | µg/g | µg/g | µg/g | µg/g | µg/g | µg/g | % | % | % |  | yyyymmdd | yyyymmdd | °C | % | initials | model
NPS-MOOS | 7.5882 | 0.0056891 | 26.7377 | 0.005762 | 38.2149 | 0.011248 | 21.7402 | 0 | 22.767 | 0 | 124.8505 | 0 | 21.7402 | 22.767 | 124.8505 | 0.0002 | 0.0002 | 0.002 | 0.075565 | 6.0375e-05 | 0.26626 | 9.3267e-05 | 0.38055 | 0.00015359 | 100% | 101% | 100.3% | NIST | 20240502 | 20240503 | 23 | 45.1 | EP | MS204S
In a 100 mL bottle
5. Preparation of the blends
Sample | Vial | Sample | Spike | mSAMPLE | mSPIKE | wSAMPLE | wSPIKE
ID | g | g | g | g | g | g | g
LNSW_0 | 8.1777 | 23.5246 | 23.5246 | 15.3469 | 0 | 1 | 0
LNSW_1 | 8.2073 | 23.5508 | 23.7078 | 15.3435 | 0.157 | 0.98987 | 0.010129
LNSW_2 | 8.1863 | 23.5401 | 23.8507 | 15.3538 | 0.3106 | 0.98017 | 0.019828
LNSW_3 | 8.1575 | 23.508 | 23.9741 | 15.3505 | 0.4661 | 0.97053 | 0.029469
MOOS_0 | 8.1419 | 23.432 | 23.432 | 15.2901 | 0 | 1 | 0
MOOS_1 | 8.1497 | 23.4403 | 23.5969 | 15.2906 | 0.1566 | 0.98986 | 0.010138
MOOS_2 | 8.1565 | 23.4565 | 23.7672 | 15.3 | 0.3107 | 0.9801 | 0.019903
MOOS_3 | 8.1125 | 23.3768 | 23.8433 | 15.2643 | 0.4665 | 0.97034 | 0.029655
LNSW = NPS-MOOS standard
MOOS = MOOS-4 CRM
SPIKE = NPS-1 solution
6. Nitrite sample preparation
Sample | Vial | Blend | Chemistry | mBLEND | mSAMPLE | mSPIKE | mTOTAL | Raw signal | Grav | Density | w(NO2−)ADD |  | LINEST LNSW |  |  | LINEST MOOS-4
ID | g | g | g | g | g | g | g | A(541) | A(541) | A(541) | µg/g
LNSW_N0 | 7.3092 | 12.4533 | 12.6545 | 5.1441 | 5.1441 | 0 | 5.3453 | 0.083995 | 0.08728 | 0.085217 | 0 |  | 1.1708 | 0.085332 |  | 1.182 | 0.085916
LNSW_N1 | 7.3243 | 12.4631 | 12.6635 | 5.1388 | 5.0868 | 0.052049 | 5.3392 | 0.16804 | 0.17638 | 0.17222 | 0.073974 |  | 0.0016072 | 0.00021675 |  | 0.00018671 | 2.5307e-05
LNSW_N2 | 7.2996 | 12.4484 | 12.6494 | 5.1488 | 5.0467 | 0.10209 | 5.3498 | 0.24606 | 0.26083 | 0.25467 | 0.14481 |  | 1 | 0.00025749 |  | 1 | 3.0098e-05
LNSW_N3 | 7.2998 | 12.449 | 12.6502 | 5.1492 | 4.9975 | 0.15174 | 5.3504 | 0.32271 | 0.3455 | 0.33733 | 0.21521 |  | 530681.1397 | 2 |  | 40077297.7811 | 2
MOOS_N0 | 7.2995 | 12.4365 | 12.6378 | 5.137 | 5.137 | 0 | 5.3383 | 0.084394 | 0.087701 | 0.085898 | 0 |  | 0.035185 | 1.326e-07 |  | 0.036307 | 1.8118e-09
MOOS_N1 | 7.3269 | 12.4452 | 12.6462 | 5.1183 | 5.0664 | 0.051888 | 5.3193 | 0.16864 | 0.17706 | 0.17342 | 0.074021
MOOS_N2 | 7.2919 | 12.4324 | 12.6336 | 5.1405 | 5.0382 | 0.10231 | 5.3417 | 0.24819 | 0.26314 | 0.25773 | 0.14534
MOOS_N3 | 7.3191 | 12.4348 | 12.6357 | 5.1157 | 4.964 | 0.15171 | 5.3166 | 0.32585 | 0.349 | 0.34183 | 0.21653
| Value | Uncertainty
α(NO2−) | 1.0096 | 0.001395
7. Phosphate sample preparation
Sample | Vial | Blend | Chemistry | mBLEND | mSAMPLE | mSPIKE | mTOTAL | Raw signal | Grav | Density | w(PO43−)ADD |  | LINEST LNSW |  |  | LINEST MOOS-4
ID | g | g | g | g | g | g | g | A(890) | A(890) | A(890) | µg/g
LNSW_P0 | 7.326 | 13.4983 | 14.2828 | 6.1723 | 6.1723 | 0 | 6.9568 | 0.27951 | 0.31504 | 0.30759 | 0 |  | 1.288 | 0.30988 |  | 1.2909 | 0.30827
LNSW_P1 | 7.32 | 13.498 | 14.2845 | 6.178 | 6.1154 | 0.062575 | 6.9645 | 0.55881 | 0.6364 | 0.62136 | 0.24023 |  | 0.0067357 | 0.0029499 |  | 0.011474 | 0.0050495
LNSW_P2 | 7.3673 | 13.5409 | 14.3272 | 6.1736 | 6.0512 | 0.12241 | 6.9599 | 0.81784 | 0.94065 | 0.91842 | 0.47027 |  | 0.99995 | 0.0035043 |  | 0.99984 | 0.0060052
LNSW_P3 | 7.3208 | 13.4949 | 14.2823 | 6.1741 | 5.9922 | 0.18194 | 6.9615 | 1.0643 | 1.2365 | 1.2073 | 0.69881 |  | 36562.4479 | 2 |  | 12657.1065 | 2
MOOS_P0 | 7.3119 | 13.4725 | 14.2606 | 6.1606 | 6.1606 | 0 | 6.9487 | 0.27619 | 0.31152 | 0.30511 | 0 |  | 0.44898 | 2.456e-05 |  | 0.45644 | 7.2125e-05
MOOS_P1 | 7.3161 | 13.4695 | 14.256 | 6.1534 | 6.091 | 0.062382 | 6.9399 | 0.55584 | 0.63331 | 0.62028 | 0.24034
MOOS_P2 | 7.3382 | 13.5115 | 14.298 | 6.1733 | 6.0504 | 0.12287 | 6.9598 | 0.81985 | 0.94307 | 0.92368 | 0.47202
MOOS_P3 | 7.2856 | 13.4451 | 14.2335 | 6.1595 | 5.9768 | 0.18266 | 6.9479 | 1.0636 | 1.2364 | 1.211 | 0.70294
| Value | Uncertainty
α(PO43−) | 1.0023 | 0.010336
No matrix effect, α is not statistically different then 1
8. Silicate sample preparation
Sample | Vial | Blend | Chemistry | mBLEND | mSAMPLE | mSPIKE | mTOTAL | Raw signal | Grav | Density | w(SiO2)ADD |  | LINEST LNSW |  |  | LINEST MOOS-4
ID | g | g | g | g | g | g | g | A(810) | A(810) | A(810) | µg/g
LNSW_Si0 | 7.3489 | 10.4231 | 10.7471 | 3.0742 | 3.0742 | 0 | 3.3982 | 0.12779 | 0.14126 | 0.13792 | 0 |  | 0.36918 | 0.13679 |  | 0.36822 | 0.13603
LNSW_Si1 | 7.3384 | 10.4088 | 10.7317 | 3.0704 | 3.0393 | 0.031099 | 3.3933 | 0.2425 | 0.27075 | 0.26435 | 0.35024 |  | 0.0020039 | 0.0012799 |  | 0.0018823 | 0.0012077
LNSW_Si2 | 7.3052 | 10.3836 | 10.7076 | 3.0784 | 3.0174 | 0.06104 | 3.4024 | 0.35417 | 0.39937 | 0.38993 | 0.68558 |  | 0.99994 | 0.0015206 |  | 0.99995 | 0.0014361
LNSW_Si3 | 7.3302 | 10.4117 | 10.7347 | 3.0815 | 2.9907 | 0.090809 | 3.4045 | 0.46216 | 0.5261 | 0.51367 | 1.0193 |  | 33939.8982 | 2 |  | 38270.7 | 2
MOOS_Si0 | 7.286 | 10.3598 | 10.6834 | 3.0738 | 3.0738 | 0 | 3.3974 | 0.12651 | 0.13983 | 0.13695 | 0 |  | 0.078476 | 4.6244e-06 |  | 0.07893 | 4.1248e-06
MOOS_Si1 | 7.345 | 10.4359 | 10.7592 | 3.0909 | 3.0596 | 0.031335 | 3.4142 | 0.2419 | 0.26994 | 0.26439 | 0.35073
MOOS_Si2 | 7.3327 | 10.4003 | 10.7239 | 3.0676 | 3.0065 | 0.061054 | 3.3912 | 0.35136 | 0.39631 | 0.38816 | 0.68801
MOOS_Si3 | 7.3352 | 10.3943 | 10.7178 | 3.0591 | 2.9684 | 0.090718 | 3.3826 | 0.46098 | 0.52531 | 0.51451 | 1.0249
| Value | Uncertainty
α(SiO2) | 0.99742 | 0.0074369
No matrix effect, α is not statistically different then 1
Temperature effect on the determination of nitrite, phosphate and silicate by UV-vis
Date | 2024-05-07
1. Preparation of the LNSW standard matched to the MOOS-4
| Origin |  | Origin |  | Origin |  | m0 = vial |  | m1 = m0 + standard |  | m2 = m1 + LNSW | [7 empty] | Final |  | Final |  | Final |  | Target
ID solution | w(NO2) | u[w(NO2)] | w(PO4) | u[w(PO4)] | w(SiO2) | u[w(SiO2)] | load | unload | load | unload | load | unload | m0 | m1 | m2 | u(m0) | u(m1) | u(m2) | w(NO2) | u[w(NO2)] | w(PO4) | u[w(PO4)] | w(SiO2) | u[w(SiO2)] | NO2 | PO4 | SiO2 | Traciability | Prep. Date | Exp. Date | t | Rel Hum | Analyst | Balance
unique | µg/g | µg/g | µg/g | µg/g | µg/g | µg/g | g | g | g | g | g | g | g | g | g | g | g | g | µg/g | µg/g | µg/g | µg/g | µg/g | µg/g | % | % | % |  | yyyymmdd | yyyymmdd | °C | % | initials | model
NPS-MOOS | 7.5882 | 0.0056891 | 26.7377 | 0.005762 | 38.2149 | 0.011248 | 21.7464 | -0.0001 | 22.7805 | 0 | 124.1941 | 0 | 21.7465 | 22.7805 | 124.1941 | 0.0002 | 0.0002 | 0.002 | 0.076591 | 6.1144e-05 | 0.26988 | 9.4126e-05 | 0.38572 | 0.00015517 | 101.4% | 102.4% | 101.6% | NIST | 20240506 | 20240507 | 21.7 | 49.5 | EP | MS204S
In a 100 mL bottle
50 mL in the fridge and 50 mL at room temperature for one day
Measure 3 times each nutrient at each temperature
18 vials total
2. Nitrite sample preparation
Sample | Vial | Sample | Chemistry | mSAMPLE | mTOTAL | Abs | Abs (norm)
ID | g | g | g | g | g | 541nm | 541nm
RT_N1 | 7.293 | 12.4159 | 12.6153 | 5.1229 | 5.3223 | 0.085787 | 0.089126 |  | Average | 0.088973
RT_N2 | 7.3152 | 12.437 | 12.6303 | 5.1218 | 5.3151 | 0.085791 | 0.089029 |  | SD | 0.0001878
RT_N3 | 7.3397 | 12.4591 | 12.6512 | 5.1194 | 5.3115 | 0.085553 | 0.088763 |  | RSD | 0.2111%
Fridge_N1 | 7.3484 | 12.511 | 12.7027 | 5.1626 | 5.3543 | 0.085043 | 0.088201 |  | Average | 0.088409
Fridge_N2 | 7.2995 | 12.317 | 12.5088 | 5.0175 | 5.2093 | 0.084657 | 0.087893 |  | SD | 0.00064638
Fridge_N3 | 7.3265 | 12.2825 | 12.4742 | 4.956 | 5.1477 | 0.085815 | 0.089134 |  | RSD | 0.7311%
t-Test: Paired Two Sample for Means
|  | Variable 1 | Variable 2
Mean |  | 0.088973 | 0.088409
Variance |  | 3.5269e-08 | 4.178e-07
Observations |  | 3 | 3
Pearson Correlation |  | -0.87643
Hypothesized Mean Difference |  | 0
df |  | 2
t Stat |  | 1.1957
P(T<=t) one-tail |  | 0.17718
t Critical one-tail |  | 2.92
P(T<=t) two-tail |  | 0.35437
t Critical two-tail |  | 4.3027
F<FCRITICAL no difference for nitrite when the color development is carried out in the sample at 4 °C or room temperature (21-22 °C)
3. Phosphate sample preparation
Sample | Vial | Sample | Chemistry | mSAMPLE | mTOTAL | Abs | Abs (norm)
ID | g | g | g | g | g | 890nm | 890nm
RT_P1 | 7.311 | 13.43 | 14.2172 | 6.119 | 6.9062 | 0.28332 | 0.31977 |  | Average | 0.32033
RT_P2 | 7.3359 | 13.4721 | 14.2646 | 6.1362 | 6.9287 | 0.28377 | 0.32042 |  | SD | 0.00051653
RT_P3 | 7.3068 | 13.4353 | 14.2273 | 6.1285 | 6.9205 | 0.28408 | 0.32079 |  | RSD | 0.1612%
Fridge_P1 | 7.3121 | 13.4863 | 14.2794 | 6.1742 | 6.9673 | 0.28616 | 0.32291 |  | Average | 0.32252
Fridge_P2 | 7.3331 | 13.3102 | 14.1044 | 5.9771 | 6.7713 | 0.28465 | 0.32248 |  | SD | 0.00037644
Fridge_P3 | 7.3517 | 13.2251 | 14.0183 | 5.8734 | 6.6666 | 0.28383 | 0.32217 |  | RSD | 0.1167%
t-Test: Paired Two Sample for Means
|  | Variable 1 | Variable 2
Mean |  | 0.32033 | 0.32252
Variance |  | 2.668e-07 | 1.4171e-07
Observations |  | 3 | 3
Pearson Correlation |  | -0.99817
Hypothesized Mean Difference |  | 0
df |  | 2
t Stat |  | -4.2479
P(T<=t) one-tail |  | 0.025599
t Critical one-tail |  | 2.92
P(T<=t) two-tail |  | 0.051199
t Critical two-tail |  | 4.3027
F<FCRITICAL no difference for phosphate when the color development is carried out in the sample at 4 °C or room temperature (21-22 °C)
4. Silicate sample preparation
Sample | Vial | Sample | Chemistry | mSAMPLE | mTOTAL | Abs | Abs (norm)
ID | g | g | g | g | g | 810nm | 890nm
RT_Si1 | 7.327 | 10.389 | 10.7112 | 3.062 | 3.3842 | 0.12907 | 0.14265 |  | Average | 0.14187
RT_Si2 | 7.307 | 10.3606 | 10.6845 | 3.0536 | 3.3775 | 0.12767 | 0.14121 |  | SD | 0.00072517
RT_Si3 | 7.3517 | 10.4184 | 10.7419 | 3.0667 | 3.3902 | 0.12823 | 0.14176 |  | RSD | 0.5111%
Fridge_Si1 | 7.3272 | 10.4126 | 10.7362 | 3.0854 | 3.409 | 0.10326 | 0.1141 |  | Average | 0.10892
Fridge_Si2 | 7.3144 | 10.3148 | 10.6388 | 3.0004 | 3.3244 | 0.098004 | 0.10859 |  | SD | 0.0050168
Fridge_Si3 | 7.3518 | 10.3073 | 10.6319 | 2.9555 | 3.2801 | 0.093779 | 0.10408 |  | RSD | 4.606%
Signal suppression at 4°C:
-23.23%
For silicate it is important to carry out the derivatization at constant room temperaure
Silicate interference on phosphate
Date | 2024-05-06
| Origin |  | m0 = vial |  | m1 = m0 + standard |  | m2 = m1 + DIW | [7 empty] | Final
ID solution | w(SiO2) | u[w(SiO2)] | load | unload | load | unload | load | unload | m0 | m1 | m2 | u(m0) | u(m1) | u(m2) | w(SiO2) | u[w(SiO2)] | Target | Traciability | Prep. Date | Exp. Date | t | Rel Hum | Analyst | Balance
unique | µg/g | µg/g | g | g | g | g | g | g | g | g | g | g | g | g | µg/g | µg/g | % |  | yyyymmdd | yyyymmdd | °C | % | initials | model
SP-SiO2 | 744.4628 | 0.20097 | 8.1937 | -0.0002 | 9.2017 | 0 | 32.6558 | 0 | 8.1938 | 9.2017 | 32.6558 | 0.0002 | 0.0002 | 0.0002 | 30.6739 | 0.011949 | N/A | NIST | 20240506 | 20240507 | 21.6 | 49.4 | EP | MS204S
In a 30 mL bottle |  |  |  |  | 1 |  | 24
Spike with silicate the LNSW and than measure the phosphate signal in a 5 cm cuvette
Sample | Vial | DIW | SP-SiO2 | DIW | SP-SiO2 | Chemistry | mDIW | mSPIKE | mSAMPLE | w(SiO2) | mTOTAL | Abs | Abs (norm) | Abs (norm)
ID | g | g | g | g | g | g | g | g | g | µg/g | g | 890nm | 890nm | 890nm
0 | 6.1271 | 9.093 | 9.093 | 12.0726 | 12.0726 | 12.8617 | 5.9455 | 0 | 5.9455 | 0 | 6.7346 | -0.000756 | -0.00085634 | 0
0.1 | 6.1249 | 9.0978 | 9.1079 | 12.0744 | 12.084 | 12.878 | 5.9394 | 0.0197 | 5.9591 | 0.1014 | 6.7531 | -0.000761 | -0.0008624 | -6.0589e-06
0.4 | 6.1425 | 9.1174 | 9.1566 | 12.1185 | 12.1581 | 12.9538 | 5.9368 | 0.0788 | 6.0156 | 0.40181 | 6.8113 | -0.000758 | -0.00085826 | -1.9247e-06
1 | 6.1394 | 9.1039 | 9.2028 | 12.1622 | 12.2613 | 13.0578 | 5.9239 | 0.198 | 6.1219 | 0.99208 | 6.9184 | -0.00068 | -0.00076847 | 8.7865e-05
1.6 | 6.1499 | 9.1091 | 9.2675 | 12.2174 | 12.3754 | 13.1733 | 5.9091 | 0.3164 | 6.2255 | 1.5589 | 7.0234 | -0.000789 | -0.00089012 | -3.3785e-05
2.2 | 6.1271 | 9.0785 | 9.2968 | 12.2385 | 12.4569 | 13.2538 | 5.8931 | 0.4367 | 6.3298 | 2.1162 | 7.1267 | -0.00058 | -0.00065302 | 0.00020332
3 | 7.3029 | 10.2434 | 10.5422 | 13.4724 | 13.7711 | 14.5678 | 5.8707 | 0.5975 | 6.4682 | 2.8335 | 7.2649 | -0.000679 | -0.00076263 | 9.3704e-05
At 3.0 mg/L silicate, the effect on phosphate is only 0.0001 abs, whereas at 0.5 mg/kg the effect is negligible
Uncertainty estimated from the performance verification on the instrument
Validation of silicate by LC-ICPMS: calibration plot, analytical range, matrix effect, LOD, LOQ, signal repeatability
Date | 20240502
[6 empty] | x-axis | y-axis | y-axis
Vial | Sample | Int ST | mA | mB | wA | wA*(mA/mB) | r28/30 | r28/30
g | g | g | g | g | ng/g | ng/g | average | ST dev
0.9423 | 1.4385 | 1.9354 | 0.4962 | 0.4969 | 0 | 0 | 0 | 0
0.9306 | 1.4282 | 1.9261 | 0.4976 | 0.4979 | 42.6397 | 42.614 | 0.12144 | 0.033017
0.9422 | 1.4396 | 1.9379 | 0.4974 | 0.4983 | 206.555 | 206.1819 | 0.44778 | 0.0039692
0.9309 | 1.4281 | 1.9263 | 0.4972 | 0.4982 | 413.2076 | 412.3782 | 0.91565 | 0.050157
0.9314 | 1.4276 | 1.9267 | 0.4962 | 0.4991 | 853.1011 | 848.1442 | 1.727 | 0.010273
Fit IDMS
parameters | value | uncertainty |  | covariance
a0 | 0.006304 | 0.017444 | a0a1 | -1.6076e-06
a1 | 0.0023271 | 0.00013156 | a0a2 | -7.3505e-07
a2 | 0.00017117 | 7.3212e-05 | a1a2 | 9.3239e-09
R^2 | 0.99947
x | 0 | 4 | 8 | 12 | 16 | 20 | 24 | 28 | 32 | 36 | 40 | 44 | 48 | 52 | 56 | 60 | 64 | 68 | 72 | 76 | 80 | 84 | 88 | 92 | 96 | 100 | 104 | 108 | 112 | 116 | 120 | 124 | 128 | 132 | 136 | 140 | 144 | 148 | 152 | 156 | 160 | 164 | 168 | 172 | 176 | 180 | 184 | 188 | 192 | 196 | 200 | 204 | 208 | 212 | 216 | 220 | 224 | 228 | 232 | 236 | 240 | 244 | 248 | 252 | 256 | 260 | 264 | 268 | 272 | 276 | 280 | 284 | 288 | 292 | 296 | 300 | 304 | 308 | 312 | 316 | 320 | 324 | 328 | 332 | 336 | 340 | 344 | 348 | 352 | 356 | 360 | 364 | 368 | 372 | 376 | 380 | 384 | 388 | 392 | 396 | 400 | 404 | 408 | 412 | 416 | 420 | 424 | 428 | 432 | 436 | 440 | 444 | 448 | 452 | 456 | 460 | 464 | 468 | 472 | 476 | 480 | 484 | 488 | 492 | 496 | 500 | 504 | 508 | 512 | 516 | 520 | 524 | 528 | 532 | 536 | 540 | 544 | 548 | 552 | 556 | 560 | 564 | 568 | 572 | 576 | 580 | 584 | 588 | 592 | 596 | 600 | 604 | 608 | 612 | 616 | 620 | 624 | 628 | 632 | 636 | 640 | 644 | 648 | 652 | 656 | 660 | 664 | 668 | 672 | 676 | 680 | 684 | 688 | 692 | 696 | 700 | 704 | 708 | 712 | 716 | 720 | 724 | 728 | 732 | 736 | 740 | 744 | 748 | 752 | 756 | 760 | 764 | 768 | 772 | 776 | 780 | 784 | 788 | 792 | 796 | 800 | 804 | 808 | 812 | 816 | 820 | 824 | 828 | 832 | 836 | 840 | 844 | 848 | 852 | 856 | 860 | 864 | 868 | 872 | 876 | 880 | 884 | 888 | 892 | 896 | 900 | 904 | 908 | 912 | 916 | 920
y | 0.006304 | 0.015606 | 0.024896 | 0.034172 | 0.043437 | 0.052688 | 0.061927 | 0.071153 | 0.080367 | 0.089568 | 0.098757 | 0.10793 | 0.1171 | 0.12625 | 0.13539 | 0.14451 | 0.15363 | 0.16273 | 0.17182 | 0.1809 | 0.18996 | 0.19901 | 0.20805 | 0.21708 | 0.2261 | 0.2351 | 0.24409 | 0.25307 | 0.26204 | 0.271 | 0.27994 | 0.28887 | 0.29779 | 0.3067 | 0.3156 | 0.32448 | 0.33335 | 0.34221 | 0.35106 | 0.3599 | 0.36872 | 0.37753 | 0.38634 | 0.39513 | 0.4039 | 0.41267 | 0.42142 | 0.43017 | 0.4389 | 0.44762 | 0.45633 | 0.46502 | 0.47371 | 0.48238 | 0.49105 | 0.4997 | 0.50834 | 0.51696 | 0.52558 | 0.53419 | 0.54278 | 0.55136 | 0.55993 | 0.56849 | 0.57704 | 0.58558 | 0.59411 | 0.60262 | 0.61113 | 0.61962 | 0.6281 | 0.63657 | 0.64503 | 0.65348 | 0.66192 | 0.67035 | 0.67876 | 0.68717 | 0.69556 | 0.70395 | 0.71232 | 0.72068 | 0.72903 | 0.73737 | 0.7457 | 0.75402 | 0.76233 | 0.77062 | 0.77891 | 0.78718 | 0.79545 | 0.8037 | 0.81195 | 0.82018 | 0.8284 | 0.83661 | 0.84481 | 0.853 | 0.86118 | 0.86935 | 0.87751 | 0.88566 | 0.8938 | 0.90193 | 0.91004 | 0.91815 | 0.92625 | 0.93433 | 0.94241 | 0.95048 | 0.95853 | 0.96658 | 0.97461 | 0.98264 | 0.99065 | 0.99866 | 1.0066 | 1.0146 | 1.0226 | 1.0306 | 1.0385 | 1.0465 | 1.0544 | 1.0623 | 1.0702 | 1.0781 | 1.086 | 1.0939 | 1.1018 | 1.1097 | 1.1175 | 1.1254 | 1.1332 | 1.141 | 1.1488 | 1.1566 | 1.1644 | 1.1722 | 1.18 | 1.1878 | 1.1955 | 1.2033 | 1.211 | 1.2187 | 1.2264 | 1.2341 | 1.2418 | 1.2495 | 1.2572 | 1.2649 | 1.2725 | 1.2802 | 1.2878 | 1.2955 | 1.3031 | 1.3107 | 1.3183 | 1.3259 | 1.3335 | 1.3411 | 1.3486 | 1.3562 | 1.3637 | 1.3713 | 1.3788 | 1.3863 | 1.3938 | 1.4013 | 1.4088 | 1.4163 | 1.4238 | 1.4312 | 1.4387 | 1.4461 | 1.4536 | 1.461 | 1.4684 | 1.4758 | 1.4832 | 1.4906 | 1.498 | 1.5054 | 1.5127 | 1.5201 | 1.5274 | 1.5348 | 1.5421 | 1.5494 | 1.5567 | 1.564 | 1.5713 | 1.5786 | 1.5859 | 1.5932 | 1.6004 | 1.6077 | 1.6149 | 1.6221 | 1.6294 | 1.6366 | 1.6438 | 1.651 | 1.6582 | 1.6654 | 1.6725 | 1.6797 | 1.6868 | 1.694 | 1.7011 | 1.7083 | 1.7154 | 1.7225 | 1.7296 | 1.7367 | 1.7438 | 1.7508 | 1.7579 | 1.765 | 1.772 | 1.7791 | 1.7861 | 1.7931 | 1.8001 | 1.8072 | 1.8142 | 1.8212 | 1.8281 | 1.8351 | 1.8421 | 1.849 | 1.856
LOD estimation and range as SiO2
w(SiO2) | 42.6397 | ng/g
Signal | 7308.0533 | Height
S/N | 21.5 | Height, RSM x 1
LOD | 5.9497 | ng/g
LOD | 99.0237 | nmol/L
LOQ | 19.8324 | ng/g
Range | 20-850 | ng/g
Overlap between 1 blank and 1 signal at 42.6 ng/g SiO2
Matrix effect estimated on the signal of the internal standard 30Si and repeatability of the analytical signal within MOOS-4
| Area 30Si | Area 30Si |  | Area 28Si
| in MOOS-4 | in DIW |  | in MOOS-4
| 1332744 | 1333334 |  | 1350612
| 1336728 | 1361576 |  | 1343316
| 1371992 | 1386333 |  | 1386241
| 1329318 | 1347631 |  | 1288786
| 1374279 | 1354108 |  | 1393242
| 1362962 | 1362538 |  | 1289474
1358967
1321080
1355947
1376305
1286542
1324463
1352025
1325209
1333384
1312039
1312904
1283438
1289958
1338270
Average | 1351337.1667 | 1335802.55 |  | 1341945.1667
SD | 20649.7647 | 29099.1085 |  | 45276.0411
RSD | 1.528% | 2.178% |  | 3.374%
| Value | Uncertainty
Ratio | 1.0116 | 0.026919
Conclusion: no detectable matrix suppression, LC separate the matrix efficiently
𝑦=(0.0063+0.00233∙𝑥)/(1+0.00017∙𝑥)
𝑹^𝟐=𝟎.𝟗𝟗𝟗𝟓

## UNITS
Index
Line | Description
13 | Red series: characterization
33 | Green series: short-term stability and homogeneiety
52 | Blue series: internal use and archive
67 | Weighing the units of starting material after irradiation
416 | Cleaning of the bottles
425 | Weighing the units after reblending and rebottling
Red series: characterization
Unit | Unit | Analyzed | Notes
number | ID | date
1 | MOOS001a | 2024-02-20 | CHAR week 1 GC-MS, UV-vis Enea, density
26.7143 | MOOS027a | 2024-03-20 | CHAR Silicate LC-ICPMS Zuzana
52.4286 | MOOS052a | 2024-03-05 | CHAR week 2 GC-MS, UV-vis Enea, density
78.1429 | MOOS078a | 2024-02-20 | CHAR week 1 GC-MS, UV-vis Enea, density
103.8571 | MOOS104a | 2024-05-02 | Matrix effect evaluation
129.5714 | MOOS130a | 2024-03-05 | CHAR week 2 GC-MS, UV-vis Enea, density
155.2857 | MOOS155a | 2024-02-20 | CHAR week 1 GC-MS, UV-vis Enea, density
181 | MOOS181a | 2024-03-20 | CHAR Silicate LC-ICPMS Zuzana
206.7143 | MOOS207a | 2024-03-05 | CHAR week 2 GC-MS, UV-vis Enea, density
232.4286 | MOOS232a | 2024-02-20 | CHAR week 1 GC-MS, UV-vis Enea, density
258.1429 | MOOS258a | 2024-05-02 | Matrix effect evaluation
283.8571 | MOOS284a | 2024-03-05 | CHAR week 2 GC-MS, UV-vis Enea, density
309.5714 | MOOS310a | 2024-02-20 | CHAR week 1 GC-MS, UV-vis Enea, density
335.2857 | MOOS335a | 2024-03-20 | CHAR Silicate LC-ICPMS Zuzana
360 | MOOS360a | 2024-03-05 | CHAR week 2 GC-MS, UV-vis Enea, density
Green series: short-term stability and homogeneiety
Unit | Unit | Storage, t | From | Until | Days | Analyzed | Notes
number | ID | °C | date | date | n. | date
9 | MOOS009a | 21 | 2023-05-03 | 2023-05-15 | 12 | 2024-01-15 | see SHORT_HOM tab
47 | MOOS047a | 21 | 2023-05-03 | 2023-05-30 | 27 | 2024-01-15 | see SHORT_HOM tab
85 | MOOS085a | 4 | # ref | # ref | # ref | 2024-01-15 | see SHORT_HOM tab
123 | MOOS123a | 40 | 2023-05-03 | 2023-05-15 | 12 | 2024-01-15 | see SHORT_HOM tab
161 | MOOS161a | 40 | 2023-05-03 | 2023-05-30 | 27 | 2024-01-15 | see SHORT_HOM tab
199 | MOOS199a | 21 | 2023-05-03 | 2023-06-07 | 35 | 2024-01-15 | see SHORT_HOM tab
237 | MOOS237a | 4 | # ref | # ref | # ref | 2024-01-15 | see SHORT_HOM tab
275 | MOOS275a | 21 | 2023-05-03 | 2023-06-15 | 43 | 2024-01-15 | see SHORT_HOM tab
313 | MOOS313a | 40 | 2023-05-03 | 2023-06-07 | 35 | 2024-01-15 | see SHORT_HOM tab
351 | MOOS351a | 40 | 2023-05-03 | 2023-06-15 | 43 | 2024-01-15 | see SHORT_HOM tab
RT = room temperature ~ 21°C in lab B5 (CTC CombiPAL drawer (no light))
Thermostat = 40 °C in lab B5
Reference conditions: fridge of B5 lab (4 °C)
Blue series: internal use and archive
Unit | Unit | Analyzed | Notes
number | ID | date
18 | MOOS018a
53 | MOOS053a
88 | MOOS088a
124 | MOOS124a
159 | MOOS159a
194 | MOOS194a
229 | MOOS229a
264 | MOOS264a
299 | MOOS299a
333 | MOOS333a
Weighing the units of starting material after irradiation
Date | 2023-04-11
Box numbers of the starting material:
7 | 54 | 2 | 24 | 28 | 29 | 38
Wipe with EtOH each unit and discard the moldy boxes
Box | Unit | Unit mass | Unit mass | Delta mass | Sample
N | ID | 2023-04-11 | 2017-03-09 | g | g
7 | MOOS0384 | 71.2861 | 71.3289 | -0.0428 | 59.4624 |  | Max | -0.0401
7 | MOOS0341 | 72.3586 | 72.3998 | -0.0412 | 60.5349 |  | Min | -0.0785
7 | MOOS0347 | 75.027 | 75.068 | -0.041 | 63.2033 |  | Average | -0.045011
7 | MOOS0352 | 74.7826 | 74.8249 | -0.0423 | 62.9589 |  | SD | 0.0042986
7 | MOOS0360 | 74.2822 | 74.3248 | -0.0426 | 62.4585 |  | RSD | -9.55%
7 | MOOS0371 | 75.3718 | 75.4148 | -0.043 | 63.5481
7 | MOOS0344 | 73.0668 | 73.1098 | -0.043 | 61.2431 |  | Empty (g) | 11.8237
7 | MOOS0353 | 73.7026 | 73.7451 | -0.0425 | 61.8789 |  | Total (g) | 20491.105
7 | MOOS0364 | 77.0161 | 77.0585 | -0.0424 | 65.1924 |  | Total (mL) | 19894.2767
7 | MOOS0365 | 74.5026 | 74.5453 | -0.0427 | 62.6789
7 | MOOS0385 | 73.4866 | 73.5284 | -0.0418 | 61.6629 |  | Transp. | -0.1272%
7 | MOOS0370 | 75.0223 | 75.0639 | -0.0416 | 63.1986
7 | MOOS0348 | 72.6951 | 72.7358 | -0.0407 | 60.8714 |  | N. new units | 397.8855
7 | MOOS0343 | 72.7547 | 72.796 | -0.0413 | 60.931 |  | (50 mL)
7 | MOOS0359 | 72.8968 | 72.9389 | -0.0421 | 61.0731
7 | MOOS0355 | 75.3496 | 75.3914 | -0.0418 | 63.5259
7 | MOOS0372 | 76.2653 | 76.3081 | -0.0428 | 64.4416
7 | MOOS0381 | 74.2027 | 74.2458 | -0.0431 | 62.379
7 | MOOS0374 | 75.5013 | 75.5445 | -0.0432 | 63.6776
7 | MOOS0367 | 74.9715 | 75.0139 | -0.0424 | 63.1478
7 | MOOS0351 | 74.6568 | 74.6989 | -0.0421 | 62.8331
7 | MOOS0340 | 75.6008 | 75.6412 | -0.0404 | 63.7771
7 | MOOS0350 | 75.0607 | 75.1028 | -0.0421 | 63.237
7 | MOOS0369 | 75.0846 | 75.1265 | -0.0419 | 63.2609
7 | MOOS0345 | 73.788 | 73.8331 | -0.0451 | 61.9643
7 | MOOS0357 | 74.9781 | 75.0203 | -0.0422 | 63.1544
7 | MOOS0354 | 72.4005 | 72.4419 | -0.0414 | 60.5768
7 | MOOS0346 | 75.5081 | 75.5511 | -0.043 | 63.6844
7 | MOOS0363 | 75.193 | 75.2349 | -0.0419 | 63.3693
7 | MOOS0380 | 74.2159 | 74.2584 | -0.0425 | 62.3922
7 | MOOS0342 | 75.7819 | 75.8232 | -0.0413 | 63.9582
7 | MOOS0338 | 75.5699 | 75.6134 | -0.0435 | 63.7462
7 | MOOS0361 | 74.3029 | 74.3461 | -0.0432 | 62.4792
7 | MOOS0379 | 75.6071 | 75.648 | -0.0409 | 63.7834
7 | MOOS0356 | 72.942 | 72.9831 | -0.0411 | 61.1183
7 | MOOS0366 | 75.9674 | 76.01 | -0.0426 | 64.1437
7 | MOOS0376 | 75.2187 | 75.2611 | -0.0424 | 63.395
7 | MOOS0336 | 74.1343 | 74.1766 | -0.0423 | 62.3106
7 | MOOS0373 | 74.0818 | 74.1248 | -0.043 | 62.2581
7 | MOOS0382 | 75.9639 | 76.0071 | -0.0432 | 64.1402
7 | MOOS0362 | 75.7501 | 75.7915 | -0.0414 | 63.9264
7 | MOOS0339 | 77.0249 | 77.066 | -0.0411 | 65.2012
7 | MOOS0349 | 75.0952 | 75.1389 | -0.0437 | 63.2715
7 | MOOS0358 | 74.3018 | 74.345 | -0.0432 | 62.4781
7 | MOOS0334 | 73.2707 | 73.3121 | -0.0414 | 61.447
7 | MOOS0333 | 73.1939 | 73.2349 | -0.041 | 61.3702
7 | MOOS0337 | 74.7984 | 74.8403 | -0.0419 | 62.9747
7 | MOOS0383 | 76.4016 | 76.4454 | -0.0438 | 64.5779
7 | MOOS0335 | 73.4184 | 73.4609 | -0.0425 | 61.5947
7 | MOOS0368 | 74.3847 | 74.4278 | -0.0431 | 62.561
54 | MOOS2967 | 75.1986 | 75.2431 | -0.0445 | 63.3749
54 | MOOS2978 | 73.9307 | 73.9766 | -0.0459 | 62.107
54 | MOOS2987 | 72.3088 | 72.3562 | -0.0474 | 60.4851
54 | MOOS2979 | 70.282 | 70.3277 | -0.0457 | 58.4583
54 | MOOS2982 | 72.5843 | 72.6278 | -0.0435 | 60.7606
54 | MOOS2962 | 72.9108 | 72.9557 | -0.0449 | 61.0871
54 | MOOS2948 | 71.8069 | 71.8473 | -0.0404 | 59.9832
54 | MOOS2961 | 73.8369 | 73.8823 | -0.0454 | 62.0132
54 | MOOS2949 | 74.2883 | 74.3302 | -0.0419 | 62.4646
54 | MOOS2960 | 73.4306 | 73.4752 | -0.0446 | 61.6069
54 | MOOS2957 | 72.8567 | 72.8984 | -0.0417 | 61.033
54 | MOOS2968 | 71.4747 | 71.5194 | -0.0447 | 59.651
54 | MOOS2973 | 73.3364 | 73.3794 | -0.043 | 61.5127
54 | MOOS2966 | 71.5212 | 71.5646 | -0.0434 | 59.6975
54 | MOOS2983 | 71.4618 | 71.5068 | -0.045 | 59.6381
54 | MOOS2954 | 72.8092 | 72.8517 | -0.0425 | 60.9855
54 | MOOS2952 | 71.8003 | 71.8437 | -0.0434 | 59.9766
54 | MOOS2980 | 72.5718 | 72.6174 | -0.0456 | 60.7481
54 | MOOS2981 | 72.7278 | 72.7692 | -0.0414 | 60.9041
54 | MOOS2971 | 73.0995 | 73.1408 | -0.0413 | 61.2758
54 | MOOS2964 | 72.4006 | 72.4421 | -0.0415 | 60.5769
54 | MOOS2984 | 72.8626 | 72.9085 | -0.0459 | 61.0389
54 | MOOS2951 | 72.4186 | 72.4625 | -0.0439 | 60.5949
54 | MOOS2963 | 71.1914 | 71.2364 | -0.045 | 59.3677
54 | MOOS2985 | 74.144 | 74.191 | -0.047 | 62.3203
54 | MOOS2969 | 72.3932 | 72.4389 | -0.0457 | 60.5695
54 | MOOS2986 | 71.363 | 71.4094 | -0.0464 | 59.5393
54 | MOOS2974 | 72.5091 | 72.5538 | -0.0447 | 60.6854
54 | MOOS2977 | 71.9019 | 71.9483 | -0.0464 | 60.0782
54 | MOOS2959 | 71.5268 | 71.5698 | -0.043 | 59.7031
54 | MOOS2947 | 71.5386 | 71.5806 | -0.042 | 59.7149
54 | MOOS2975 | 73.1983 | 73.2444 | -0.0461 | 61.3746
54 | MOOS2970 | 73.3272 | 73.3724 | -0.0452 | 61.5035
54 | MOOS2956 | 72.0227 | 72.0643 | -0.0416 | 60.199
54 | MOOS2953 | 73.7147 | 73.7601 | -0.0454 | 61.891
54 | MOOS2965 | 73.9884 | 74.0324 | -0.044 | 62.1647
54 | MOOS2958 | 73.1472 | 73.19 | -0.0428 | 61.3235
54 | MOOS2950 | 73.8754 | 73.9191 | -0.0437 | 62.0517
2 | MOOS0068 | 71.8534 | 71.8985 | -0.0451 | 60.0297
2 | MOOS0083 | 72.368 | 72.4092 | -0.0412 | 60.5443
2 | MOOS0113 | 73.0182 | 73.0605 | -0.0423 | 61.1945
2 | MOOS0110 | 72.6119 | 72.6565 | -0.0446 | 60.7882
2 | MOOS0107 | 71.2777 | 71.3209 | -0.0432 | 59.454
2 | MOOS0063 | 73.673 | 73.7146 | -0.0416 | 61.8493
2 | MOOS0064 | 74.297 | 74.3397 | -0.0427 | 62.4733
2 | MOOS0070 | 74.8505 | 74.8947 | -0.0442 | 63.0268
2 | MOOS0103 | 70.8876 | 70.9317 | -0.0441 | 59.0639
2 | MOOS0092 | 73.5078 | 73.551 | -0.0432 | 61.6841
2 | MOOS0058 | 74.8524 | 74.8965 | -0.0441 | 63.0287
2 | MOOS0108 | 72.8198 | 72.8628 | -0.043 | 60.9961
2 | MOOS0093 | 72.9879 | 73.0297 | -0.0418 | 61.1642
2 | MOOS0111 | 72.5273 | 72.5707 | -0.0434 | 60.7036
2 | MOOS0076 | 73.0431 | 73.0863 | -0.0432 | 61.2194
2 | MOOS0073 | 70.3471 | 70.3895 | -0.0424 | 58.5234
2 | MOOS0088 | 74.4641 | 74.5099 | -0.0458 | 62.6404
2 | MOOS0089 | 72.2675 | 72.3093 | -0.0418 | 60.4438
2 | MOOS0114 | 71.2256 | 71.2678 | -0.0422 | 59.4019
2 | MOOS0102 | 73.0751 | 73.1187 | -0.0436 | 61.2514
2 | MOOS0071 | 75.0991 | 75.1436 | -0.0445 | 63.2754
2 | MOOS0099 | 74.606 | 74.6492 | -0.0432 | 62.7823
2 | MOOS0065 | 74.0572 | 74.0999 | -0.0427 | 62.2335
2 | MOOS0112 | 73.5877 | 73.6313 | -0.0436 | 61.764
2 | MOOS0100 | 71.9425 | 71.9862 | -0.0437 | 60.1188
2 | MOOS0062 | 76.2499 | 76.2946 | -0.0447 | 64.4262
2 | MOOS0085 | 76.2049 | 76.2502 | -0.0453 | 64.3812
2 | MOOS0090 | 75.2652 | 75.309 | -0.0438 | 63.4415
2 | MOOS0097 | 71.5245 | 71.5663 | -0.0418 | 59.7008
2 | MOOS0104 | 73.1297 | 73.1747 | -0.045 | 61.306
2 | MOOS0082 | 72.4242 | 72.4669 | -0.0427 | 60.6005
2 | MOOS0109 | 69.6251 | 69.6675 | -0.0424 | 57.8014
2 | MOOS0077 | 71.0012 | 71.045 | -0.0438 | 59.1775
2 | MOOS0105 | 70.8847 | 70.9295 | -0.0448 | 59.061
2 | MOOS0066 | 73.649 | 73.6928 | -0.0438 | 61.8253
2 | MOOS0086 | 71.1876 | 71.2325 | -0.0449 | 59.3639
2 | MOOS0087 | 74.1004 | 74.1456 | -0.0452 | 62.2767
2 | MOOS0067 | 72.6343 | 72.6784 | -0.0441 | 60.8106
2 | MOOS0084 | 72.7872 | 72.8299 | -0.0427 | 60.9635
2 | MOOS0095 | 75.1279 | 75.1724 | -0.0445 | 63.3042
2 | MOOS0079 | 74.3788 | 74.4239 | -0.0451 | 62.5551
2 | MOOS0115 | 70.3672 | 70.409 | -0.0418 | 58.5435
2 | MOOS0081 | 73.5508 | 73.5926 | -0.0418 | 61.7271
2 | MOOS0074 | 74.9647 | 75.0089 | -0.0442 | 63.141
2 | MOOS0106 | 71.4481 | 71.4905 | -0.0424 | 59.6244
2 | MOOS0096 | 73.0812 | 73.1249 | -0.0437 | 61.2575
2 | MOOS0072 | 72.0474 | 72.0888 | -0.0414 | 60.2237
2 | MOOS0091 | 71.1679 | 71.2112 | -0.0433 | 59.3442
2 | MOOS0080 | 75.7992 | 75.8447 | -0.0455 | 63.9755
2 | MOOS0098 | 71.1462 | 71.1897 | -0.0435 | 59.3225
24 | MOOS1336 | 72.4364 | 72.492 | -0.0556 | 60.6127
24 | MOOS1332 | 75.358 | 75.401 | -0.043 | 63.5343
24 | MOOS1304 | 73.0952 | 73.1379 | -0.0427 | 61.2715
24 | MOOS1327 | 71.6626 | 71.7055 | -0.0429 | 59.8389
24 | MOOS1285 | 72.7231 | 72.7655 | -0.0424 | 60.8994
24 | MOOS1287 | 70.6199 | 70.6606 | -0.0407 | 58.7962
24 | MOOS1302 | 72.6395 | 72.6826 | -0.0431 | 60.8158
24 | MOOS1293 | 71.4986 | 71.5414 | -0.0428 | 59.6749
24 | MOOS1317 | 74.8002 | 74.8424 | -0.0422 | 62.9765
24 | MOOS1334 | 72.8468 | 72.8902 | -0.0434 | 61.0231
24 | MOOS1335 | 72.7424 | 72.7862 | -0.0438 | 60.9187
24 | MOOS1282 | 74.2211 | 74.2685 | -0.0474 | 62.3974
24 | MOOS1329 | 72.6273 | 72.67 | -0.0427 | 60.8036
24 | MOOS1316 | 72.8052 | 72.848 | -0.0428 | 60.9815
24 | MOOS1297 | 73.35 | 73.393 | -0.043 | 61.5263
24 | MOOS1324 | 73.0384 | 73.0801 | -0.0417 | 61.2147
24 | MOOS1290 | 74.0511 | 74.094 | -0.0429 | 62.2274
24 | MOOS1279 | 73.6569 | 73.6984 | -0.0415 | 61.8332
24 | MOOS1318 | 72.4969 | 72.5407 | -0.0438 | 60.6732
24 | MOOS1306 | 73.5404 | 73.5824 | -0.042 | 61.7167
24 | MOOS1312 | 74.3666 | 74.412 | -0.0454 | 62.5429
24 | MOOS1313 | 74.845 | 74.8878 | -0.0428 | 63.0213
24 | MOOS1298 | 73.9816 | 74.0304 | -0.0488 | 62.1579
24 | MOOS1331 | 75.0876 | 75.1307 | -0.0431 | 63.2639
24 | MOOS1308 | 72.6343 | 72.6771 | -0.0428 | 60.8106
24 | MOOS1299 | 73.4063 | 73.4492 | -0.0429 | 61.5826
24 | MOOS1319 | 73.6601 | 73.7047 | -0.0446 | 61.8364
24 | MOOS1303 | 73.4548 | 73.4974 | -0.0426 | 61.6311
24 | MOOS1281 | 75.1574 | 75.1993 | -0.0419 | 63.3337
24 | MOOS1330 | 72.4912 | 72.5351 | -0.0439 | 60.6675
24 | MOOS1301 | 72.2362 | 72.2785 | -0.0423 | 60.4125
24 | MOOS1315 | 73.8534 | 73.8956 | -0.0422 | 62.0297
24 | MOOS1310 | 72.4966 | 72.5403 | -0.0437 | 60.6729
24 | MOOS1321 | 72.5207 | 72.5646 | -0.0439 | 60.697
24 | MOOS1328 | 73.5915 | 73.6356 | -0.0441 | 61.7678
24 | MOOS1311 | 73.2976 | 73.3421 | -0.0445 | 61.4739
24 | MOOS1307 | 73.3207 | 73.3631 | -0.0424 | 61.497
24 | MOOS1291 | 72.9798 | 73.0211 | -0.0413 | 61.1561
24 | MOOS1305 | 71.8629 | 71.9047 | -0.0418 | 60.0392
24 | MOOS1323 | 72.3199 | 72.3632 | -0.0433 | 60.4962
24 | MOOS1288 | 71.6358 | 71.6782 | -0.0424 | 59.8121
24 | MOOS1309 | 72.6121 | 72.6557 | -0.0436 | 60.7884
24 | MOOS1283 | 72.5999 | 72.642 | -0.0421 | 60.7762
24 | MOOS1296 | 74.3934 | 74.4373 | -0.0439 | 62.5697
24 | MOOS1286 | 71.3089 | 71.3517 | -0.0428 | 59.4852
24 | MOOS1325 | 72.5602 | 72.604 | -0.0438 | 60.7365
24 | MOOS1294 | 73.5874 | 73.6308 | -0.0434 | 61.7637
24 | MOOS1326 | 71.4711 | 71.5143 | -0.0432 | 59.6474
24 | MOOS1292 | 73.9548 | 73.9975 | -0.0427 | 62.1311
24 | MOOS1280 | 72.0911 | 72.133 | -0.0419 | 60.2674
28 | MOOS1519 | 72.2661 | 72.3094 | -0.0433 | 60.4424
28 | MOOS1542 | 74.3807 | 74.4263 | -0.0456 | 62.557
28 | MOOS1523 | 72.6413 | 72.6854 | -0.0441 | 60.8176
28 | MOOS1543 | 76.2669 | 76.3118 | -0.0449 | 64.4432
28 | MOOS1551 | 76.067 | 76.1126 | -0.0456 | 64.2433
28 | MOOS1555 | 73.0387 | 73.0859 | -0.0472 | 61.215
28 | MOOS1534 | 72.5313 | 72.5763 | -0.045 | 60.7076
28 | MOOS1527 | 73.3909 | 73.4339 | -0.043 | 61.5672
28 | MOOS1525 | 73.302 | 73.345 | -0.043 | 61.4783
28 | MOOS1533 | 73.3718 | 73.4279 | -0.0561 | 61.5481
28 | MOOS1544 | 73.6405 | 73.6868 | -0.0463 | 61.8168
28 | MOOS1532 | 73.1931 | 73.2456 | -0.0525 | 61.3694
28 | MOOS1563 | 75.571 | 75.6171 | -0.0461 | 63.7473
28 | MOOS1529 | 73.1501 | 73.2128 | -0.0627 | 61.3264
28 | MOOS1540 | 75.9528 | 75.998 | -0.0452 | 64.1291
28 | MOOS1521 | 74.3119 | 74.3565 | -0.0446 | 62.4882
28 | MOOS1511 | 73.8345 | 73.8763 | -0.0418 | 62.0108
28 | MOOS1553 | 73.2281 | 73.2747 | -0.0466 | 61.4044
28 | MOOS1516 | 72.7092 | 72.7526 | -0.0434 | 60.8855
28 | MOOS1545 | 73.2347 | 73.2811 | -0.0464 | 61.411
28 | MOOS1546 | 71.5447 | 71.5901 | -0.0454 | 59.721
28 | MOOS1600 | 73.5992 | 73.6777 | -0.0785 | 61.7755
28 | MOOS1512 | 74.0665 | 74.1091 | -0.0426 | 62.2428
28 | MOOS1505 | 75.0635 | 75.1052 | -0.0417 | 63.2398
28 | MOOS1550 | 73.0522 | 73.0987 | -0.0465 | 61.2285
28 | MOOS1522 | 74.897 | 74.9419 | -0.0449 | 63.0733
28 | MOOS1562 | 75.3396 | 75.3862 | -0.0466 | 63.5159
28 | MOOS1549 | 75.2364 | 75.2835 | -0.0471 | 63.4127
28 | MOOS1536 | 71.7213 | 71.7655 | -0.0442 | 59.8976
28 | MOOS1538 | 72.3168 | 72.3621 | -0.0453 | 60.4931
28 | MOOS1557 | 76.3743 | 76.4227 | -0.0484 | 64.5506
28 | MOOS1556 | 72.1458 | 72.1928 | -0.047 | 60.3221
28 | MOOS1520 | 72.8714 | 72.9151 | -0.0437 | 61.0477
28 | MOOS1548 | 74.3052 | 74.3529 | -0.0477 | 62.4815
28 | MOOS1554 | 75.844 | 75.8913 | -0.0473 | 64.0203
28 | MOOS1558 | 72.2198 | 72.267 | -0.0472 | 60.3961
28 | MOOS1559 | 75.713 | 75.7589 | -0.0459 | 63.8893
28 | MOOS1561 | 73.2293 | 73.2763 | -0.047 | 61.4056
28 | MOOS1514 | 72.7828 | 72.8265 | -0.0437 | 60.9591
28 | MOOS1524 | 72.3776 | 72.4208 | -0.0432 | 60.5539
28 | MOOS1535 | 74.9697 | 75.0137 | -0.044 | 63.146
28 | MOOS1537 | 73.9669 | 74.0121 | -0.0452 | 62.1432
28 | MOOS1530 | 73.6258 | 73.6698 | -0.044 | 61.8021
28 | MOOS1510 | 72.1652 | 72.2062 | -0.041 | 60.3415
28 | MOOS1518 | 74.6103 | 74.6529 | -0.0426 | 62.7866
28 | MOOS1539 | 72.746 | 72.7922 | -0.0462 | 60.9223
28 | MOOS1541 | 77.2119 | 77.2582 | -0.0463 | 65.3882
28 | MOOS1547 | 74.2932 | 74.3406 | -0.0474 | 62.4695
28 | MOOS1513 | 72.4724 | 72.5157 | -0.0433 | 60.6487
29 | MOOS1564 | 74.5129 | 74.5622 | -0.0493 | 62.6892
29 | MOOS1507 | 74.352 | 74.3968 | -0.0448 | 62.5283
29 | MOOS1568 | 75.4578 | 75.5054 | -0.0476 | 63.6341
29 | MOOS1613 | 71.9301 | 71.982 | -0.0519 | 60.1064
29 | MOOS1592 | 73.5856 | 73.6356 | -0.05 | 61.7619
29 | MOOS1579 | 74.6099 | 74.6607 | -0.0508 | 62.7862
29 | MOOS1573 | 72.7377 | 72.7864 | -0.0487 | 60.914
29 | MOOS1591 | 75.7536 | 75.802 | -0.0484 | 63.9299
29 | MOOS1586 | 75.7445 | 75.7942 | -0.0497 | 63.9208
29 | MOOS1612 | 73.446 | 73.4985 | -0.0525 | 61.6223
29 | MOOS1596 | 72.8914 | 72.9424 | -0.051 | 61.0677
29 | MOOS1581 | 75.182 | 75.2316 | -0.0496 | 63.3583
29 | MOOS1585 | 72.2094 | 72.2583 | -0.0489 | 60.3857
29 | MOOS1508 | 71.8283 | 71.8715 | -0.0432 | 60.0046
29 | MOOS1593 | 71.0282 | 71.0767 | -0.0485 | 59.2045
29 | MOOS1572 | 75.325 | 75.3755 | -0.0505 | 63.5013
29 | MOOS1606 | 71.2279 | 71.2765 | -0.0486 | 59.4042
29 | MOOS1580 | 75.1271 | 75.1785 | -0.0514 | 63.3034
29 | MOOS1566 | 73.8136 | 73.8631 | -0.0495 | 61.9899
29 | MOOS1610 | 72.3178 | 72.369 | -0.0512 | 60.4941
29 | MOOS1576 | 75.6314 | 75.6799 | -0.0485 | 63.8077
29 | MOOS1611 | 71.7892 | 71.841 | -0.0518 | 59.9655
29 | MOOS1575 | 73.7754 | 73.8224 | -0.047 | 61.9517
29 | MOOS1605 | 71.8272 | 71.8769 | -0.0497 | 60.0035
29 | MOOS1570 | 75.5714 | 75.6205 | -0.0491 | 63.7477
29 | MOOS1614 | 72.0485 | 72.0996 | -0.0511 | 60.2248
29 | MOOS1589 | 72.583 | 72.6332 | -0.0502 | 60.7593
29 | MOOS1597 | 70.4638 | 70.5142 | -0.0504 | 58.6401
29 | MOOS1608 | 72.1755 | 72.2238 | -0.0483 | 60.3518
29 | MOOS1587 | 75.535 | 75.5862 | -0.0512 | 63.7113
29 | MOOS1609 | 71.3258 | 71.3784 | -0.0526 | 59.5021
29 | MOOS1603 | 73.5531 | 73.6047 | -0.0516 | 61.7294
29 | MOOS1607 | 72.4526 | 72.5028 | -0.0502 | 60.6289
29 | MOOS1604 | 72.0717 | 72.1228 | -0.0511 | 60.248
29 | MOOS1571 | 75.2995 | 75.3496 | -0.0501 | 63.4758
29 | MOOS1594 | 72.9742 | 73.0244 | -0.0502 | 61.1505
29 | MOOS1578 | 73.7652 | 73.8157 | -0.0505 | 61.9415
29 | MOOS1601 | 71.9043 | 71.9572 | -0.0529 | 60.0806
29 | MOOS1574 | 74.8142 | 74.8631 | -0.0489 | 62.9905
29 | MOOS1599 | 72.2878 | 72.3402 | -0.0524 | 60.4641
29 | MOOS1567 | 74.3382 | 74.3888 | -0.0506 | 62.5145
29 | MOOS1615 | 72.3572 | 72.4091 | -0.0519 | 60.5335
29 | MOOS1577 | 75.1891 | 75.2381 | -0.049 | 63.3654
29 | MOOS1616 | 71.5763 | 71.6535 | -0.0772 | 59.7526
29 | MOOS1584 | 73.3234 | 73.3706 | -0.0472 | 61.4997
29 | MOOS1569 | 74.6688 | 74.7176 | -0.0488 | 62.8451
29 | MOOS1602 | 71.6449 | 71.6964 | -0.0515 | 59.8212
29 | MOOS1590 | 71.2419 | 71.2912 | -0.0493 | 59.4182
29 | MOOS1565 | 75.2807 | 75.3285 | -0.0478 | 63.457
29 | MOOS1588 | 71.4424 | 71.4935 | -0.0511 | 59.6187
38 | MOOS2067 | 75.4843 | 75.5357 | -0.0514 | 63.6606
38 | MOOS2074 | 75.4754 | 75.5165 | -0.0411 | 63.6517
38 | MOOS2111 | 72.6552 | 72.6957 | -0.0405 | 60.8315
38 | MOOS2066 | 73.5722 | 73.6136 | -0.0414 | 61.7485
38 | MOOS2058 | 76.7429 | 76.7917 | -0.0488 | 64.9192
38 | MOOS2069 | 74.9541 | 74.9994 | -0.0453 | 63.1304
38 | MOOS2096 | 71.1737 | 71.2218 | -0.0481 | 59.35
38 | MOOS2094 | 75.5465 | 75.5883 | -0.0418 | 63.7228
38 | MOOS2100 | 75.07 | 75.1122 | -0.0422 | 63.2463
38 | MOOS2097 | 74.1054 | 74.1528 | -0.0474 | 62.2817
38 | MOOS2115 | 73.244 | 73.2887 | -0.0447 | 61.4203
38 | MOOS2072 | 74.6055 | 74.6466 | -0.0411 | 62.7818
38 | MOOS2103 | 75.6641 | 75.7157 | -0.0516 | 63.8404
38 | MOOS2061 | 73.0074 | 73.0662 | -0.0588 | 61.1837
38 | MOOS2102 | 73.5623 | 73.6024 | -0.0401 | 61.7386
38 | MOOS2068 | 74.8197 | 74.8663 | -0.0466 | 62.996
38 | MOOS2055 | 72.3339 | 72.3742 | -0.0403 | 60.5102
38 | MOOS2079 | 72.6998 | 72.7428 | -0.043 | 60.8761
38 | MOOS2099 | 73.9512 | 73.9938 | -0.0426 | 62.1275
38 | MOOS2110 | 72.9652 | 73.0054 | -0.0402 | 61.1415
38 | MOOS2077 | 75.3152 | 75.3609 | -0.0457 | 63.4915
38 | MOOS2098 | 75.5082 | 75.5536 | -0.0454 | 63.6845
38 | MOOS2113 | 73.7893 | 73.8397 | -0.0504 | 61.9656
38 | MOOS2073 | 75.8945 | 75.9361 | -0.0416 | 64.0708
38 | MOOS2114 | 72.2117 | 72.2583 | -0.0466 | 60.388
38 | MOOS2080 | 74.8575 | 74.8988 | -0.0413 | 63.0338
38 | MOOS2095 | 75.1288 | 75.1702 | -0.0414 | 63.3051
38 | MOOS2070 | 74.4921 | 74.5365 | -0.0444 | 62.6684
38 | MOOS2075 | 75.092 | 75.1414 | -0.0494 | 63.2683
38 | MOOS2059 | 74.8434 | 74.8896 | -0.0462 | 63.0197
38 | MOOS2091 | 75.1198 | 75.1628 | -0.043 | 63.2961
38 | MOOS2064 | 73.3422 | 73.3842 | -0.042 | 61.5185
38 | MOOS2078 | 71.7183 | 71.7784 | -0.0601 | 59.8946
38 | MOOS2060 | 72.8347 | 72.8777 | -0.043 | 61.011
38 | MOOS2087 | 74.4397 | 74.4906 | -0.0509 | 62.616
38 | MOOS2082 | 75.4899 | 75.5312 | -0.0413 | 63.6662
38 | MOOS2105 | 71.7594 | 71.8045 | -0.0451 | 59.9357
38 | MOOS2106 | 74.1233 | 74.1675 | -0.0442 | 62.2996
38 | MOOS2112 | 73.1645 | 73.2062 | -0.0417 | 61.3408
38 | MOOS2090 | 73.1871 | 73.2319 | -0.0448 | 61.3634
38 | MOOS2101 | 74.9009 | 74.9425 | -0.0416 | 63.0772
38 | MOOS2089 | 74.2315 | 74.2779 | -0.0464 | 62.4078
38 | MOOS2093 | 73.9071 | 73.9488 | -0.0417 | 62.0834
38 | MOOS2071 | 73.9388 | 73.9814 | -0.0426 | 62.1151
38 | MOOS2088 | 75.1986 | 75.2473 | -0.0487 | 63.3749
38 | MOOS2056 | 75.6797 | 75.7207 | -0.041 | 63.856
38 | MOOS2104 | 73.5337 | 73.5819 | -0.0482 | 61.71
38 | MOOS2081 | 75.5518 | 75.5935 | -0.0417 | 63.7281
38 | MOOS2065 | 72.7573 | 72.7975 | -0.0402 | 60.9336
38 | MOOS2108 | 74.8447 | 74.8884 | -0.0437 | 63.021
Cleaning of the bottles
Date | 2023-04-24 and 2023-04-25
Each bottle was rinsed with EtOH 95 % (sterilization), following rinsing three times with GenPure water (B3) and finally filled with GenPure water (B3)
During cleaning the bottles were kept in both upside up and down position.
Date | 20230501
The bottle were emptyed and left for 2+ hours with the cap loose. Then the cap was tight and the bottles were weighed
Weighing the units after reblending and rebottling
weight of empty unit: bottle + cap + serial label
weight after clean: the bottle sterilized with EtOH and then rinsed with water. Bottle + cap + serial label + moisture that remains in the bottle after cleaning
weight after bottling. Bottle + cap + serial label + moisture that remains in the bottle after cleaning + MOOS-4 sample
Unit | Unit | Weight | Weight | Weight
|  | 20230324 | 20230501 | 20230503 |  | Dilution | Sample
number | ID | empty | after clean | after bottling |  | check | g
1 | MOOS001a | 11.811 | 11.8126 | 65.5573 |  | 0.0016 | 53.7463
2 | MOOS002a | 11.8278 | 11.8294 | 62.4633 |  | 0.0016 | 50.6355
3 | MOOS003a | 11.7929 | 11.7967 | 64.3758 |  | 0.0038 | 52.5829
4 | MOOS004a | 11.7909 | 11.7912 | 65.2383 |  | 0.0003 | 53.4474
5 | MOOS005a | 11.8212 | 11.8235 | 61.9519 |  | 0.0023 | 50.1307
6 | MOOS006a | 11.8321 | 11.8338 | 64.9339 |  | 0.0017 | 53.1018
7 | MOOS007a | 11.7895 | 11.7906 | 65.0391 |  | 0.0011 | 53.2496
8 | MOOS008a | 11.8291 | 11.8304 | 64.0177 |  | 0.0013 | 52.1886
9 | MOOS009a | 11.8069 | 11.8075 | 63.3898 |  | 0.0006 | 51.5829
10 | MOOS010a | 11.8216 | 11.8241 | 63.5721 |  | 0.0025 | 51.7505
11 | MOOS011a | 11.7922 | 11.7941 | 64.1538 |  | 0.0019 | 52.3616
12 | MOOS012a | 11.8629 | 11.8641 | 66.6514 |  | 0.0012 | 54.7885
13 | MOOS013a | 11.8092 | 11.8103 | 64.4987 |  | 0.0011 | 52.6895
14 | MOOS014a | 11.8432 | 11.8445 | 66.6284 |  | 0.0013 | 54.7852
15 | MOOS015a | 11.7746 | 11.7776 | 65.7414 |  | 0.003 | 53.9668
16 | MOOS016a | 11.7874 | 11.7898 | 65.3847 |  | 0.0024 | 53.5973
17 | MOOS017a | 11.8142 | 11.8164 | 64.4831 |  | 0.0022 | 52.6689
18 | MOOS018a | 11.7606 | 11.7619 | 66.837 |  | 0.0013 | 55.0764
19 | MOOS019a | 11.881 | 11.881 | 64.3562 |  | 0 | 52.4752
20 | MOOS020a | 11.773 | 11.7757 | 64.4525 |  | 0.0027 | 52.6795
21 | MOOS021a | 11.7975 | 11.7995 | 65.3382 |  | 0.002 | 53.5407
22 | MOOS022a | 11.8025 | 11.8045 | 65.2102 |  | 0.002 | 53.4077
23 | MOOS023a | 11.7968 | 11.7997 | 65.1572 |  | 0.0029 | 53.3604
24 | MOOS024a | 11.8259 | 11.8269 | 64.7467 |  | 0.001 | 52.9208
25 | MOOS025a | 11.8628 | 11.8643 | 64.6657 |  | 0.0015 | 52.8029
26 | MOOS026a | 11.7955 | 11.7969 | 64.2029 |  | 0.0014 | 52.4074
27 | MOOS027a | 11.7996 | 11.8017 | 64.1478 |  | 0.0021 | 52.3482
28 | MOOS028a | 11.8383 | 11.8395 | 64.0754 |  | 0.0012 | 52.2371
29 | MOOS029a | 11.8234 | 11.8267 | 64.3311 |  | 0.0033 | 52.5077
30 | MOOS030a | 11.8009 | 11.803 | 64.094 |  | 0.0021 | 52.2931
31 | MOOS031a | 11.81 | 11.8122 | 64.2968 |  | 0.0022 | 52.4868
32 | MOOS032a | 11.804 | 11.8051 | 64.4184 |  | 0.0011 | 52.6144
33 | MOOS033a | 11.7854 | 11.7868 | 65.8167 |  | 0.0014 | 54.0313
34 | MOOS034a | 11.8274 | 11.8293 | 65.1422 |  | 0.0019 | 53.3148
35 | MOOS035a | 11.8261 | 11.8287 | 65.4445 |  | 0.0026 | 53.6184
36 | MOOS036a | 11.7851 | 11.787 | 63.6595 |  | 0.0019 | 51.8744
37 | MOOS037a | 11.8025 | 11.8041 | 63.3527 |  | 0.0016 | 51.5502
38 | MOOS038a | 11.7606 | 11.7616 | 61.97 |  | 0.001 | 50.2094
39 | MOOS039a | 11.8353 | 11.8373 | 64.351 |  | 0.002 | 52.5157
40 | MOOS040a | 11.7984 | 11.801 | 64.2382 |  | 0.0026 | 52.4398
41 | MOOS041a | 11.8182 | 11.8202 | 66.0797 |  | 0.002 | 54.2615
42 | MOOS042a | 11.8057 | 11.808 | 64.3022 |  | 0.0023 | 52.4965
43 | MOOS043a | 11.7971 | 11.8 | 64.7288 |  | 0.0029 | 52.9317
44 | MOOS044a | 11.8085 | 11.8105 | 65.3671 |  | 0.002 | 53.5586
45 | MOOS045a | 11.8014 | 11.8038 | 65.4576 |  | 0.0024 | 53.6562
46 | MOOS046a | 11.7561 | 11.7568 | 65.0352 |  | 0.0007 | 53.2791
47 | MOOS047a | 11.7879 | 11.7888 | 64.4784 |  | 0.0009 | 52.6905
48 | MOOS048a | 11.7842 | 11.7863 | 63.7009 |  | 0.0021 | 51.9167
49 | MOOS049a | 11.7897 | 11.7915 | 64.9788 |  | 0.0018 | 53.1891
50 | MOOS050a | 11.859 | 11.8612 | 64.7262 |  | 0.0022 | 52.8672
51 | MOOS051a | 11.801 | 11.8046 | 64.5822 |  | 0.0036 | 52.7812
52 | MOOS052a | 11.7949 | 11.7967 | 66.2645 |  | 0.0018 | 54.4696
53 | MOOS053a | 11.8216 | 11.8227 | 64.4635 |  | 0.0011 | 52.6419
54 | MOOS054a | 11.8754 | 11.8768 | 63.8831 |  | 0.0014 | 52.0077
55 | MOOS055a | 11.8405 | 11.8414 | 64.4767 |  | 0.0009 | 52.6362
56 | MOOS056a | 11.7827 | 11.7846 | 65.9189 |  | 0.0019 | 54.1362
57 | MOOS057a | 11.8345 | 11.8355 | 65.2428 |  | 0.001 | 53.4083
58 | MOOS058a | 11.8309 | 11.8315 | 64.0815 |  | 0.0006 | 52.2506
59 | MOOS059a | 11.8126 | 11.8194 | 64.5075 |  | 0.0068 | 52.6949
60 | MOOS060a | 11.8624 | 11.8648 | 65.1956 |  | 0.0024 | 53.3332
61 | MOOS061a | 11.8415 | 11.8428 | 64.9218 |  | 0.0013 | 53.0803
62 | MOOS062a | 11.851 | 11.853 | 63.0604 |  | 0.002 | 51.2094
63 | MOOS063a | 11.7995 | 11.8022 | 65.0844 |  | 0.0027 | 53.2849
64 | MOOS064a | 11.8222 | 11.824 | 65.7562 |  | 0.0018 | 53.934
65 | MOOS065a | 11.814 | 11.8142 | 64.9286 |  | 0.0002 | 53.1146
66 | MOOS066a | 11.8292 | 11.831 | 63.7156 |  | 0.0018 | 51.8864
67 | MOOS067a | 11.8112 | 11.8139 | 64.0414 |  | 0.0027 | 52.2302
68 | MOOS068a | 11.8143 | 11.816 | 64.5383 |  | 0.0017 | 52.724
69 | MOOS069a | 11.8249 | 11.8253 | 64.457 |  | 0.0004 | 52.6321
70 | MOOS070a | 11.8237 | 11.8242 | 66.4567 |  | 0.0005 | 54.633
71 | MOOS071a | 11.8254 | 11.8272 | 66.1371 |  | 0.0018 | 54.3117
72 | MOOS072a | 11.8222 | 11.8241 | 64.3026 |  | 0.0019 | 52.4804
73 | MOOS073a | 11.8199 | 11.8208 | 64.858 |  | 0.0009 | 53.0381
74 | MOOS074a | 11.8112 | 11.8148 | 65.2358 |  | 0.0036 | 53.4246
75 | MOOS075a | 11.7797 | 11.7808 | 62.8327 |  | 0.0011 | 51.053
76 | MOOS076a | 11.838 | 11.841 | 63.6687 |  | 0.003 | 51.8307
77 | MOOS077a | 11.8345 | 11.8375 | 65.1951 |  | 0.003 | 53.3606
78 | MOOS078a | 11.8455 | 11.8528 | 62.6936 |  | 0.0073 | 50.8481
79 | MOOS079a | 11.8365 | 11.8384 | 62.3399 |  | 0.0019 | 50.5034
80 | MOOS080a | 11.8407 | 11.8419 | 63.6258 |  | 0.0012 | 51.7851
81 | MOOS081a | 11.8668 | 11.8711 | 64.188 |  | 0.0043 | 52.3212
82 | MOOS082a | 11.817 | 11.8193 | 62.6812 |  | 0.0023 | 50.8642
83 | MOOS083a | 11.8161 | 11.8183 | 62.8332 |  | 0.0022 | 51.0171
84 | MOOS084a | 11.787 | 11.7892 | 63.11 |  | 0.0022 | 51.323
85 | MOOS085a | 11.8168 | 11.8188 | 63.7758 |  | 0.002 | 51.959
86 | MOOS086a | 11.8659 | 11.8728 | 65.7586 |  | 0.0069 | 53.8927
87 | MOOS087a | 11.8317 | 11.84 | 63.5937 |  | 0.0083 | 51.762
88 | MOOS088a | 11.8436 | 11.8467 | 62.9883 |  | 0.0031 | 51.1447
89 | MOOS089a | 11.7881 | 11.7924 | 63.2537 |  | 0.0043 | 51.4656
90 | MOOS090a | 11.8587 | 11.8611 | 64.761 |  | 0.0024 | 52.9023
91 | MOOS091a | 11.8283 | 11.8296 | 64.8904 |  | 0.0013 | 53.0621
92 | MOOS092a | 11.8397 | 11.8432 | 64.1244 |  | 0.0035 | 52.2847
93 | MOOS093a | 11.8749 | 11.8771 | 65.1154 |  | 0.0022 | 53.2405
94 | MOOS094a | 11.8153 | 11.8189 | 64.7802 |  | 0.0036 | 52.9649
95 | MOOS095a | 11.8271 | 11.8296 | 62.8606 |  | 0.0025 | 51.0335
96 | MOOS096a | 11.7761 | 11.7784 | 64.0372 |  | 0.0023 | 52.2611
97 | MOOS097a | 11.8392 | 11.841 | 63.2577 |  | 0.0018 | 51.4185
98 | MOOS098a | 11.8246 | 11.8345 | 65.9839 |  | 0.0099 | 54.1593
99 | MOOS099a | 11.836 | 11.8373 | 64.721 |  | 0.0013 | 52.885
100 | MOOS100a | 11.8212 | 11.8234 | 64.6037 |  | 0.0022 | 52.7825
101 | MOOS101a | 11.8426 | 11.8446 | 66.1842 |  | 0.002 | 54.3416
102 | MOOS102a | 11.8342 | 11.8366 | 63.5274 |  | 0.0024 | 51.6932
103 | MOOS103a | 11.8481 | 11.8488 | 64.4592 |  | 0.0007 | 52.6111
104 | MOOS104a | 11.818 | 11.8197 | 62.8398 |  | 0.0017 | 51.0218
105 | MOOS105a | 11.8964 | 11.8976 | 62.8184 |  | 0.0012 | 50.922
106 | MOOS106a | 11.8469 | 11.849 | 63.1942 |  | 0.0021 | 51.3473
107 | MOOS107a | 11.8441 | 11.8483 | 65.9276 |  | 0.0042 | 54.0835
108 | MOOS108a | 11.8005 | 11.8029 | 63.7042 |  | 0.0024 | 51.9037
109 | MOOS109a | 11.8567 | 11.8595 | 63.8102 |  | 0.0028 | 51.9535
110 | MOOS110a | 11.8077 | 11.8103 | 63.2223 |  | 0.0026 | 51.4146
111 | MOOS111a | 11.8252 | 11.8293 | 65.0063 |  | 0.0041 | 53.1811
112 | MOOS112a | 11.7826 | 11.7863 | 63.8826 |  | 0.0037 | 52.1
113 | MOOS113a | 11.8502 | 11.8518 | 63.6615 |  | 0.0016 | 51.8113
114 | MOOS114a | 11.8175 | 11.8193 | 65.8683 |  | 0.0018 | 54.0508
115 | MOOS115a | 11.8434 | 11.8477 | 62.9716 |  | 0.0043 | 51.1282
116 | MOOS116a | 11.7723 | 11.7741 | 61.9938 |  | 0.0018 | 50.2215
117 | MOOS117a | 11.8049 | 11.8064 | 63.7558 |  | 0.0015 | 51.9509
118 | MOOS118a | 11.8204 | 11.8222 | 61.8083 |  | 0.0018 | 49.9879
119 | MOOS119a | 11.8626 | 11.864 | 59.4265 |  | 0.0014 | 47.5639
120 | MOOS120a | 11.8086 | 11.811 | 63.8144 |  | 0.0024 | 52.0058
121 | MOOS121a | 11.8857 | 11.8876 | 63.9855 |  | 0.0019 | 52.0998
122 | MOOS122a | 11.8591 | 11.8592 | 63.6373 |  | 0.0001 | 51.7782
123 | MOOS123a | 11.8388 | 11.845 | 63.5031 |  | 0.0062 | 51.6643
124 | MOOS124a | 11.8872 | 11.8891 | 63.5901 |  | 0.0019 | 51.7029
125 | MOOS125a | 11.8794 | 11.8825 | 63.6342 |  | 0.0031 | 51.7548
126 | MOOS126a | 11.8325 | 11.8335 | 64.045 |  | 0.001 | 52.2125
127 | MOOS127a | 11.7983 | 11.7982 | 64.0523 |  | -0.0001 | 52.254
128 | MOOS128a | 11.8343 | 11.8358 | 66.643 |  | 0.0015 | 54.8087
129 | MOOS129a | 11.879 | 11.8831 | 64.0217 |  | 0.0041 | 52.1427
130 | MOOS130a | 11.8357 | 11.8364 | 65.5789 |  | 0.0007 | 53.7432
131 | MOOS131a | 11.8218 | 11.8246 | 65.0202 |  | 0.0028 | 53.1984
132 | MOOS132a | 11.8005 | 11.8023 | 63.8801 |  | 0.0018 | 52.0796
133 | MOOS133a | 11.8359 | 11.8368 | 64.0965 |  | 0.0009 | 52.2606
134 | MOOS134a | 11.8518 | 11.8519 | 64.2626 |  | 0.0001 | 52.4108
135 | MOOS135a | 11.8384 | 11.8463 | 65.0184 |  | 0.0079 | 53.18
136 | MOOS136a | 11.8468 | 11.8503 | 63.4 |  | 0.0035 | 51.5532
137 | MOOS137a | 11.8391 | 11.8427 | 64.3847 |  | 0.0036 | 52.5456
138 | MOOS138a | 11.8142 | 11.814 | 63.7518 |  | -0.0002 | 51.9376
139 | MOOS139a | 11.8587 | 11.8608 | 63.9993 |  | 0.0021 | 52.1406
140 | MOOS140a | 11.8201 | 11.8261 | 62.1285 |  | 0.006 | 50.3084
141 | MOOS141a | 11.8325 | 11.8329 | 63.8702 |  | 0.0004 | 52.0377
142 | MOOS142a | 11.8187 | 11.8232 | 65.2385 |  | 0.0045 | 53.4198
143 | MOOS143a | 11.8373 | 11.8382 | 65.2963 |  | 0.0009 | 53.459
144 | MOOS144a | 11.8331 | 11.8348 | 64.5925 |  | 0.0017 | 52.7594
145 | MOOS145a | 11.8243 | 11.8262 | 62.9445 |  | 0.0019 | 51.1202
146 | MOOS146a | 11.8407 | 11.8468 | 64.9546 |  | 0.0061 | 53.1139
147 | MOOS147a | 11.8556 | 11.8563 | 65.344 |  | 0.0007 | 53.4884
148 | MOOS148a | 11.841 | 11.8434 | 63.1509 |  | 0.0024 | 51.3099
149 | MOOS149a | 11.8512 | 11.8543 | 62.9315 |  | 0.0031 | 51.0803
150 | MOOS150a | 11.8415 | 11.8417 | 62.9476 |  | 0.0002 | 51.1061
151 | MOOS151a | 11.8284 | 11.829 | 62.324 |  | 0.0006 | 50.4956
152 | MOOS152a | 11.8028 | 11.806 | 63.0109 |  | 0.0032 | 51.2081
153 | MOOS153a | 11.787 | 11.7896 | 64.0871 |  | 0.0026 | 52.3001
154 | MOOS154a | 11.8153 | 11.8175 | 63.8406 |  | 0.0022 | 52.0253
155 | MOOS155a | 11.7973 | 11.7997 | 65.5485 |  | 0.0024 | 53.7512
156 | MOOS156a | 11.7656 | 11.7713 | 66.2468 |  | 0.0057 | 54.4812
157 | MOOS157a | 11.7552 | 11.7592 | 64.5919 |  | 0.004 | 52.8367
158 | MOOS158a | 11.8544 | 11.8555 | 64.014 |  | 0.0011 | 52.1596
159 | MOOS159a | 11.8143 | 11.8168 | 64.5276 |  | 0.0025 | 52.7133
160 | MOOS160a | 11.8011 | 11.8034 | 63.1678 |  | 0.0023 | 51.3667
161 | MOOS161a | 11.7751 | 11.7775 | 65.4494 |  | 0.0024 | 53.6743
162 | MOOS162a | 11.8376 | 11.8396 | 63.9326 |  | 0.002 | 52.095
163 | MOOS163a | 11.8353 | 11.8377 | 63.7622 |  | 0.0024 | 51.9269
164 | MOOS164a | 11.7541 | 11.7556 | 63.2779 |  | 0.0015 | 51.5238
165 | MOOS165a | 11.8139 | 11.8157 | 63.8715 |  | 0.0018 | 52.0576
166 | MOOS166a | 11.8088 | 11.8105 | 63.1247 |  | 0.0017 | 51.3159
167 | MOOS167a | 11.7928 | 11.7949 | 64.1549 |  | 0.0021 | 52.3621
168 | MOOS168a | 11.7973 | 11.8005 | 63.7621 |  | 0.0032 | 51.9648
169 | MOOS169a | 11.8067 | 11.8086 | 62.2906 |  | 0.0019 | 50.4839
170 | MOOS170a | 11.7634 | 11.7644 | 64.4285 |  | 0.001 | 52.6651
171 | MOOS171a | 11.7883 | 11.792 | 64.1937 |  | 0.0037 | 52.4054
172 | MOOS172a | 11.8119 | 11.8139 | 65.6525 |  | 0.002 | 53.8406
173 | MOOS173a | 11.7981 | 11.8019 | 63.6495 |  | 0.0038 | 51.8514
174 | MOOS174a | 11.8049 | 11.8073 | 63.8728 |  | 0.0024 | 52.0679
175 | MOOS175a | 11.7912 | 11.7927 | 63.8668 |  | 0.0015 | 52.0756
176 | MOOS176a | 11.8191 | 11.8204 | 62.8493 |  | 0.0013 | 51.0302
177 | MOOS177a | 11.883 | 11.8853 | 64.4007 |  | 0.0023 | 52.5177
178 | MOOS178a | 11.77 | 11.7712 | 63.3002 |  | 0.0012 | 51.5302
179 | MOOS179a | 11.818 | 11.8205 | 65.6451 |  | 0.0025 | 53.8271
180 | MOOS180a | 11.8231 | 11.8279 | 64.2114 |  | 0.0048 | 52.3883
181 | MOOS181a | 11.8081 | 11.8084 | 65.6474 |  | 0.0003 | 53.8393
182 | MOOS182a | 11.8517 | 11.8533 | 63.0758 |  | 0.0016 | 51.2241
183 | MOOS183a | 11.8215 | 11.8233 | 64.1587 |  | 0.0018 | 52.3372
184 | MOOS184a | 11.8199 | 11.822 | 62.8784 |  | 0.0021 | 51.0585
185 | MOOS185a | 11.7349 | 11.7364 | 63.8742 |  | 0.0015 | 52.1393
186 | MOOS186a | 11.8363 | 11.8435 | 64.4694 |  | 0.0072 | 52.6331
187 | MOOS187a | 11.799 | 11.8003 | 64.7026 |  | 0.0013 | 52.9036
188 | MOOS188a | 11.8573 | 11.8588 | 65.2702 |  | 0.0015 | 53.4129
189 | MOOS189a | 11.8354 | 11.8366 | 65.171 |  | 0.0012 | 53.3356
190 | MOOS190a | 11.8052 | 11.8069 | 62.8784 |  | 0.0017 | 51.0732
191 | MOOS191a | 11.824 | 11.8257 | 63.1739 |  | 0.0017 | 51.3499
192 | MOOS192a | 11.8317 | 11.834 | 64.2225 |  | 0.0023 | 52.3908
193 | MOOS193a | 11.8142 | 11.8165 | 66.1391 |  | 0.0023 | 54.3249
194 | MOOS194a | 11.7981 | 11.8002 | 64.1113 |  | 0.0021 | 52.3132
195 | MOOS195a | 11.8486 | 11.8506 | 64.6392 |  | 0.002 | 52.7906
196 | MOOS196a | 11.8178 | 11.8183 | 65.1649 |  | 0.0005 | 53.3471
197 | MOOS197a | 11.8364 | 11.838 | 64.1649 |  | 0.0016 | 52.3285
198 | MOOS198a | 11.7756 | 11.776 | 64.265 |  | 0.0004 | 52.4894
199 | MOOS199a | 11.8276 | 11.8301 | 64.6243 |  | 0.0025 | 52.7967
200 | MOOS200a | 11.8135 | 11.8152 | 64.4302 |  | 0.0017 | 52.6167
201 | MOOS201a | 11.8429 | 11.8451 | 64.563 |  | 0.0022 | 52.7201
202 | MOOS202a | 11.8207 | 11.8229 | 64.7686 |  | 0.0022 | 52.9479
203 | MOOS203a | 11.8359 | 11.8389 | 63.767 |  | 0.003 | 51.9311
204 | MOOS204a | 11.756 | 11.7584 | 62.8129 |  | 0.0024 | 51.0569
205 | MOOS205a | 11.84 | 11.8415 | 64.1958 |  | 0.0015 | 52.3558
206 | MOOS206a | 11.7906 | 11.8015 | 64.9346 |  | 0.0109 | 53.144
207 | MOOS207a | 11.8754 | 11.8779 | 65.4132 |  | 0.0025 | 53.5378
208 | MOOS208a | 11.82 | 11.8219 | 64.9973 |  | 0.0019 | 53.1773
209 | MOOS209a | 11.8827 | 11.8849 | 64.4453 |  | 0.0022 | 52.5626
210 | MOOS210a | 11.8165 | 11.8189 | 64.9588 |  | 0.0024 | 53.1423
211 | MOOS211a | 11.8118 | 11.8148 | 65.8624 |  | 0.003 | 54.0506
212 | MOOS212a | 11.791 | 11.792 | 64.4549 |  | 0.001 | 52.6639
213 | MOOS213a | 11.839 | 11.8407 | 64.471 |  | 0.0017 | 52.632
214 | MOOS214a | 11.8369 | 11.8385 | 63.2192 |  | 0.0016 | 51.3823
215 | MOOS215a | 11.8345 | 11.8362 | 63.1862 |  | 0.0017 | 51.3517
216 | MOOS216a | 11.8267 | 11.8277 | 62.8193 |  | 0.001 | 50.9926
217 | MOOS217a | 11.799 | 11.8003 | 65.1932 |  | 0.0013 | 53.3942
218 | MOOS218a | 11.8816 | 11.8823 | 65.6887 |  | 0.0007 | 53.8071
219 | MOOS219a | 11.8492 | 11.8491 | 65.5442 |  | -0.0001 | 53.695
220 | MOOS220a | 11.8149 | 11.8169 | 65.1332 |  | 0.002 | 53.3183
221 | MOOS221a | 11.8008 | 11.8051 | 63.9597 |  | 0.0043 | 52.1589
222 | MOOS222a | 11.7973 | 11.7985 | 63.7536 |  | 0.0012 | 51.9563
223 | MOOS223a | 11.7772 | 11.7791 | 64.3262 |  | 0.0019 | 52.549
224 | MOOS224a | 11.8322 | 11.8342 | 65.1686 |  | 0.002 | 53.3364
225 | MOOS225a | 11.8292 | 11.8324 | 66.6259 |  | 0.0032 | 54.7967
226 | MOOS226a | 11.8374 | 11.8428 | 63.3764 |  | 0.0054 | 51.539
227 | MOOS227a | 11.8416 | 11.8453 | 64.2326 |  | 0.0037 | 52.391
228 | MOOS228a | 11.8086 | 11.8112 | 63.5519 |  | 0.0026 | 51.7433
229 | MOOS229a | 11.835 | 11.8372 | 63.8981 |  | 0.0022 | 52.0631
230 | MOOS230a | 11.8565 | 11.8588 | 62.0177 |  | 0.0023 | 50.1612
231 | MOOS231a | 11.8141 | 11.8173 | 65.3064 |  | 0.0032 | 53.4923
232 | MOOS232a | 11.8297 | 11.8318 | 64.0476 |  | 0.0021 | 52.2179
233 | MOOS233a | 11.8215 | 11.8229 | 63.8972 |  | 0.0014 | 52.0757
234 | MOOS234a | 11.8192 | 11.8205 | 63.1159 |  | 0.0013 | 51.2967
235 | MOOS235a | 11.8573 | 11.8597 | 63.4541 |  | 0.0024 | 51.5968
236 | MOOS236a | 11.8599 | 11.8618 | 64.7799 |  | 0.0019 | 52.92
237 | MOOS237a | 11.8197 | 11.8201 | 63.8877 |  | 0.0004 | 52.068
238 | MOOS238a | 11.8425 | 11.8448 | 65.3399 |  | 0.0023 | 53.4974
239 | MOOS239a | 11.8207 | 11.8268 | 64.9155 |  | 0.0061 | 53.0948
240 | MOOS240a | 11.8369 | 11.8396 | 66.0698 |  | 0.0027 | 54.2329
241 | MOOS241a | 11.7991 | 11.8014 | 62.2748 |  | 0.0023 | 50.4757
242 | MOOS242a | 11.8103 | 11.8124 | 62.4904 |  | 0.0021 | 50.6801
243 | MOOS243a | 11.7919 | 11.7938 | 65.279 |  | 0.0019 | 53.4871
244 | MOOS244a | 11.8305 | 11.8327 | 64.1229 |  | 0.0022 | 52.2924
245 | MOOS245a | 11.7904 | 11.793 | 65.1981 |  | 0.0026 | 53.4077
246 | MOOS246a | 11.8854 | 11.8874 | 64.622 |  | 0.002 | 52.7366
247 | MOOS247a | 11.8451 | 11.8472 | 65.1652 |  | 0.0021 | 53.3201
248 | MOOS248a | 11.8495 | 11.8515 | 65.4427 |  | 0.002 | 53.5932
249 | MOOS249a | 11.7982 | 11.8002 | 66.5674 |  | 0.002 | 54.7692
250 | MOOS250a | 11.804 | 11.8069 | 63.2019 |  | 0.0029 | 51.3979
251 | MOOS251a | 11.8829 | 11.8852 | 63.5251 |  | 0.0023 | 51.6422
252 | MOOS252a | 11.8527 | 11.8549 | 62.5509 |  | 0.0022 | 50.6982
253 | MOOS253a | 11.8423 | 11.8444 | 65.5832 |  | 0.0021 | 53.7409
254 | MOOS254a | 11.8511 | 11.8537 | 63.8141 |  | 0.0026 | 51.963
255 | MOOS255a | 11.8478 | 11.8494 | 63.9301 |  | 0.0016 | 52.0823
256 | MOOS256a | 11.8446 | 11.8468 | 64.343 |  | 0.0022 | 52.4984
257 | MOOS257a | 11.8302 | 11.8318 | 63.786 |  | 0.0016 | 51.9558
258 | MOOS258a | 11.8038 | 11.8059 | 64.5417 |  | 0.0021 | 52.7379
259 | MOOS259a | 11.8394 | 11.8427 | 65.3416 |  | 0.0033 | 53.5022
260 | MOOS260a | 11.8228 | 11.826 | 62.007 |  | 0.0032 | 50.1842
261 | MOOS261a | 11.8173 | 11.82 | 62.301 |  | 0.0027 | 50.4837
262 | MOOS262a | 11.8047 | 11.8065 | 64.767 |  | 0.0018 | 52.9623
263 | MOOS263a | 11.7716 | 11.7737 | 63.0078 |  | 0.0021 | 51.2362
264 | MOOS264a | 11.8495 | 11.8515 | 64.7965 |  | 0.002 | 52.947
265 | MOOS265a | 11.8193 | 11.8211 | 62.1867 |  | 0.0018 | 50.3674
266 | MOOS266a | 11.8242 | 11.8263 | 63.3662 |  | 0.0021 | 51.542
267 | MOOS267a | 11.8327 | 11.8377 | 64.3496 |  | 0.005 | 52.5169
268 | MOOS268a | 11.8782 | 11.8799 | 65.3925 |  | 0.0017 | 53.5143
269 | MOOS269a | 11.8401 | 11.8425 | 63.5724 |  | 0.0024 | 51.7323
270 | MOOS270a | 11.8132 | 11.8159 | 62.4641 |  | 0.0027 | 50.6509
271 | MOOS271a | 11.8022 | 11.8034 | 63.7244 |  | 0.0012 | 51.9222
272 | MOOS272a | 11.8397 | 11.8426 | 63.7883 |  | 0.0029 | 51.9486
273 | MOOS273a | 11.834 | 11.8358 | 63.2934 |  | 0.0018 | 51.4594
274 | MOOS274a | 11.794 | 11.7958 | 64.3489 |  | 0.0018 | 52.5549
275 | MOOS275a | 11.817 | 11.8183 | 63.1238 |  | 0.0013 | 51.3068
276 | MOOS276a | 11.7991 | 11.8011 | 61.6845 |  | 0.002 | 49.8854
277 | MOOS277a | 11.858 | 11.8682 | 65.6727 |  | 0.0102 | 53.8147
278 | MOOS278a | 11.8155 | 11.8171 | 63.524 |  | 0.0016 | 51.7085
279 | MOOS279a | 11.8578 | 11.8591 | 64.4034 |  | 0.0013 | 52.5456
280 | MOOS280a | 11.8355 | 11.8363 | 64.4658 |  | 0.0008 | 52.6303
281 | MOOS281a | 11.889 | 11.8911 | 62.228 |  | 0.0021 | 50.339
282 | MOOS282a | 11.8028 | 11.8047 | 63.7919 |  | 0.0019 | 51.9891
283 | MOOS283a | 11.8483 | 11.8501 | 66.3717 |  | 0.0018 | 54.5234
284 | MOOS284a | 11.8418 | 11.8443 | 65.2399 |  | 0.0025 | 53.3981
285 | MOOS285a | 11.8426 | 11.8442 | 65.4339 |  | 0.0016 | 53.5913
286 | MOOS286a | 11.8451 | 11.8466 | 65.3099 |  | 0.0015 | 53.4648
287 | MOOS287a | 11.833 | 11.8375 | 64.0909 |  | 0.0045 | 52.2579
288 | MOOS288a | 11.8745 | 11.8764 | 62.7782 |  | 0.0019 | 50.9037
289 | MOOS289a | 11.8075 | 11.8104 | 62.2948 |  | 0.0029 | 50.4873
290 | MOOS290a | 11.8111 | 11.8129 | 62.9897 |  | 0.0018 | 51.1786
291 | MOOS291a | 11.8392 | 11.8409 | 63.455 |  | 0.0017 | 51.6158
292 | MOOS292a | 11.8402 | 11.8454 | 63.1065 |  | 0.0052 | 51.2663
293 | MOOS293a | 11.7983 | 11.8009 | 66.5355 |  | 0.0026 | 54.7372
294 | MOOS294a | 11.8133 | 11.8142 | 64.4096 |  | 0.0009 | 52.5963
295 | MOOS295a | 11.8284 | 11.8301 | 64.1206 |  | 0.0017 | 52.2922
296 | MOOS296a | 11.845 | 11.8533 | 67.6923 |  | 0.0083 | 55.8473
297 | MOOS297a | 11.8369 | 11.8393 | 63.6815 |  | 0.0024 | 51.8446
298 | MOOS298a | 11.8361 | 11.8381 | 63.2843 |  | 0.002 | 51.4482
299 | MOOS299a | 11.8798 | 11.8811 | 63.1291 |  | 0.0013 | 51.2493
300 | MOOS300a | 11.8585 | 11.8601 | 62.8832 |  | 0.0016 | 51.0247
301 | MOOS301a | 11.8515 | 11.8577 | 63.5086 |  | 0.0062 | 51.6571
302 | MOOS302a | 11.8085 | 11.8131 | 62.5648 |  | 0.0046 | 50.7563
303 | MOOS303a | 11.8365 | 11.8389 | 63.8551 |  | 0.0024 | 52.0186
304 | MOOS304a | 11.8198 | 11.8215 | 64.8786 |  | 0.0017 | 53.0588
305 | MOOS305a | 11.8409 | 11.8438 | 65.7604 |  | 0.0029 | 53.9195
306 | MOOS306a | 11.8403 | 11.8477 | 63.8402 |  | 0.0074 | 51.9999
307 | MOOS307a | 11.8266 | 11.8287 | 63.3686 |  | 0.0021 | 51.542
308 | MOOS308a | 11.8206 | 11.8225 | 61.6123 |  | 0.0019 | 49.7917
309 | MOOS309a | 11.7815 | 11.7838 | 63.363 |  | 0.0023 | 51.5815
310 | MOOS310a | 11.8425 | 11.8444 | 63.305 |  | 0.0019 | 51.4625
311 | MOOS311a | 11.8246 | 11.8294 | 62.9859 |  | 0.0048 | 51.1613
312 | MOOS312a | 11.8211 | 11.8253 | 63.308 |  | 0.0042 | 51.4869
313 | MOOS313a | 11.8456 | 11.8485 | 63.4523 |  | 0.0029 | 51.6067
314 | MOOS314a | 11.8563 | 11.8584 | 62.3728 |  | 0.0021 | 50.5165
315 | MOOS315a | 11.8772 | 11.88 | 65.3597 |  | 0.0028 | 53.4825
316 | MOOS316a | 11.836 | 11.8373 | 64.0089 |  | 0.0013 | 52.1729
317 | MOOS317a | 11.8375 | 11.8449 | 63.8195 |  | 0.0074 | 51.982
318 | MOOS318a | 11.8049 | 11.8087 | 63.8064 |  | 0.0038 | 52.0015
319 | MOOS319a | 11.8887 | 11.8907 | 64.3755 |  | 0.002 | 52.4868
320 | MOOS320a | 11.8002 | 11.8099 | 64.2007 |  | 0.0097 | 52.4005
321 | MOOS321a | 11.827 | 11.834 | 64.0441 |  | 0.007 | 52.2171
322 | MOOS322a | 11.839 | 11.8427 | 63.78 |  | 0.0037 | 51.941
323 | MOOS323a | 11.8413 | 11.8469 | 64.0251 |  | 0.0056 | 52.1838
324 | MOOS324a | 11.8099 | 11.812 | 65.5513 |  | 0.0021 | 53.7414
325 | MOOS325a | 11.7878 | 11.7908 | 65.5874 |  | 0.003 | 53.7996
326 | MOOS326a | 11.819 | 11.8212 | 62.3714 |  | 0.0022 | 50.5524
327 | MOOS327a | 11.8838 | 11.8852 | 63.1689 |  | 0.0014 | 51.2851
328 | MOOS328a | 11.8122 | 11.8142 | 62.3101 |  | 0.002 | 50.4979
329 | MOOS329a | 11.8397 | 11.8411 | 63.7214 |  | 0.0014 | 51.8817
330 | MOOS330a | 11.7735 | 11.7767 | 63.9356 |  | 0.0032 | 52.1621
331 | MOOS331a | 11.8581 | 11.8616 | 63.3322 |  | 0.0035 | 51.4741
332 | MOOS332a | 11.7836 | 11.7857 | 65.6495 |  | 0.0021 | 53.8659
333 | MOOS333a | 11.8238 | 11.8254 | 63.0806 |  | 0.0016 | 51.2568
334 | MOOS334a | 11.7946 | 11.797 | 64.1144 |  | 0.0024 | 52.3198
335 | MOOS335a | 11.8573 | 11.8624 | 64.6865 |  | 0.0051 | 52.8292
336 | MOOS336a | 11.8239 | 11.8256 | 63.6577 |  | 0.0017 | 51.8338
337 | MOOS337a | 11.8056 | 11.808 | 64.1179 |  | 0.0024 | 52.3123
338 | MOOS338a | 11.7968 | 11.8002 | 64.4083 |  | 0.0034 | 52.6115
339 | MOOS339a | 11.808 | 11.8102 | 63.1648 |  | 0.0022 | 51.3568
340 | MOOS340a | 11.8168 | 11.818 | 64.2509 |  | 0.0012 | 52.4341
341 | MOOS341a | 11.8431 | 11.8439 | 62.3656 |  | 0.0008 | 50.5225
342 | MOOS342a | 11.8003 | 11.8025 | 65.4896 |  | 0.0022 | 53.6893
343 | MOOS343a | 11.842 | 11.8457 | 64.0185 |  | 0.0037 | 52.1765
344 | MOOS344a | 11.8566 | 11.8595 | 63.6043 |  | 0.0029 | 51.7477
345 | MOOS345a | 11.8196 | 11.8235 | 66.2228 |  | 0.0039 | 54.4032
346 | MOOS346a | 11.8511 | 11.8525 | 65.4781 |  | 0.0014 | 53.627
347 | MOOS347a | 11.8271 | 11.8297 | 63.8353 |  | 0.0026 | 52.0082
348 | MOOS348a | 11.8462 | 11.8477 | 65.819 |  | 0.0015 | 53.9728
349 | MOOS349a | 11.8059 | 11.8071 | 65.9169 |  | 0.0012 | 54.111
350 | MOOS350a | 11.8389 | 11.8465 | 65.882 |  | 0.0076 | 54.0431
351 | MOOS351a | 11.8199 | 11.8234 | 64.1701 |  | 0.0035 | 52.3502
352 | MOOS352a | 11.8146 | 11.8169 | 66.0993 |  | 0.0023 | 54.2847
353 | MOOS353a | 11.8304 | 11.8372 | 70.4571 |  | 0.0068 | 58.6267
354 | MOOS354a | 11.7804 | 11.7824 | 64.9875 |  | 0.002 | 53.2071
355 | MOOS355a | 11.7979 | 11.8005 | 62.9985 |  | 0.0026 | 51.2006
356 | MOOS356a | 11.8137 | 11.8176 | 63.3663 |  | 0.0039 | 51.5526
357 | MOOS357a | 11.8095 | 11.8139 | 64.4868 |  | 0.0044 | 52.6773
358 | MOOS358a | 11.8062 | 11.8129 | 63.3326 |  | 0.0067 | 51.5264
359 | MOOS359a | 11.8311 | 11.8332 | 65.0446 |  | 0.0021 | 53.2135
360 | MOOS360a | 11.8424 | 11.8442 | 66.4001 |  | 0.0018 | 54.5577
| Average | 11.8237 |  |  |  | 0.0024625 | 52.4057
| SD | 0.027552 |  |  |  | 0.0016783 | 1.168
| RSD | 0.233% |  |  |  | 68.15% | 2.229%
| Max | 11.8964 |  |  |  | 0.0109 | 58.6267
| Min | 11.7349 |  |  |  | -0.0002 | 47.5639
The 20 L carboy for the blending was cleaned with EtOH, rinsed three time in water and kept full or water overnight
The MOOS-4 starting material was blended inside the carboy
The content was manually homogenized followign the bottling
The Nalgene bottles were tight with the Kinnex Capper position 0 + 3 clicks on the right
Bottling of the MOOS-4 was done on 2023-05-03 in lab B3 (EP performing all operations)
Meter B5 065825 SPER SCIENTIFIC
At 09:20 AM: room temperature was 18.8 °C and RH 42 %
At 09:30 AM: started blending the MOOS-4 starting material
At 10:35 AM: finished re-blending and homogenization (room temperature was 18.6 °C and RH 42 %)
At 10:45 AM: started bottling (room temperature was 18.4 °C and RH 42 %)
At 12:50 AM: finished bottling (room temperature was 18.5 °C and RH 40 %)
The material was thermosealed in trilaminated bags and stored in the cold room (NRC M12, 2023-05-04)
Formulas:
Days = DAYS([Until],[Notes])

## SHORT_HOM
Index
Line | Description
17 | Weight of the units and transpiration loss
35 | Testing the transpiration loss at 40 °C in a trilaminate bag
59 | Randomization of injection order
90 | Short term stability nitrate by HS ID GC-MS (verified on 2024-04-10)
165 | Homogeneity of nitrate on the base of the stability data by ID HS GC-MS (verified on 2024-04-10)
233 | Short term stability nitrite by UV-Vis (verified 2024-04-10)
308 | Homogeneity of nitrite on the base of the stability data by UV-vis (verified on 2024-04-10)
376 | Short term stability phosphate by UV-Vis (verified on 2024-04-10)
456 | Homogeneity of phosphate on the base of the stability data by UV-vis (verified on 2024-04-10)
531 | Short term stability silicate by UV-Vis (verified on 2024-04-10)
Weight of the units and transpiration loss
Date | 2024-01-15
Unit | Storage, t | Days | Mass | Mass | Mass loss | Mass loss | Mass loss
ID | °C | n. | g | g | g | g | %
|  |  | 20240115 | 20230503 |  | sample
MOOS009a | 21 | 12 | 63.3798 | 63.3898 | -0.01 | 51.5829 | -0.01939%
MOOS047a | 21 | 27 | 64.4652 | 64.4784 | -0.0132 | 52.6905 | -0.02505%
MOOS085a | 4 | # ref | 63.768 | 63.7758 | -0.0078 | 51.959 | -0.01501%
MOOS123a | 40 | 12 | 63.4735 | 63.5031 | -0.0296 | 51.6643 | -0.05729%
MOOS161a | 40 | 27 | 65.3881 | 65.4494 | -0.0613 | 53.6743 | -0.1142%
MOOS199a | 21 | 35 | 64.6096 | 64.6243 | -0.0147 | 52.7967 | -0.02784%
MOOS237a | 4 | # ref | 63.8798 | 63.8877 | -0.0079 | 52.068 | -0.01517%
MOOS275a | 21 | 43 | 63.1091 | 63.1238 | -0.0147 | 51.3068 | -0.02865%
MOOS313a | 40 | 35 | 63.3868 | 63.4523 | -0.0655 | 51.6067 | -0.1269%
MOOS351a | 40 | 43 | 64.0913 | 64.1701 | -0.0788 | 52.3502 | -0.1505%
Testing the transpiration loss at 40 °C in a trilaminate bag
Date | 2024-01-23 | Start of the experiment
Date | 2024-03-26 | End of the experiment
| 63 | Number of days at 40-43 °C
|  |  | 2024-01-23 | 2024-01-23 | 2024-01-23 | 2024-03-26 | 2024-03-26 | 2024-03-26
Sample | Storage, t | Trilaminated | Mass load | Mass unload | Mass_0 | Mass load | Mass unload | Mass_1 | Delta mass | Relative
ID | °C | bag? | g | g | g | g | g | g | g | to 50 g (%)
1 | 40-43 | NO | 68.0614 | 0 | 68.0614 | 67.9274 | 0.0001 | 67.9274 | -0.13405 | -0.2681%
2 | 40-43 | NO | 66.026 | 0.0001 | 66.0259 | 65.8779 | 0.0001 | 65.8778 | -0.1481 | -0.2962%
3 | 40-43 | NO | 66.1269 | -0.0001 | 66.127 | 65.9728 | -0.0001 | 65.9729 | -0.1541 | -0.3082%
4 | 40-43 | YES | 66.3537 | 0 | 66.3537 | 66.3324 | 0.0001 | 66.3324 | -0.02135 | -0.0427%
5 | 40-43 | YES | 64.7405 | -0.0001 | 64.7405 | 64.7232 | 0.0001 | 64.7232 | -0.0174 | -0.0348%
6 | 40-43 | YES | 65.7961 | 0.0001 | 65.796 | 65.7778 | 0 | 65.7778 | -0.01825 | -0.0365%
Average delta mass relative to 50 g without trilaminate bag:
-0.2908%
Average delta mass relative to 50 g with trilaminate bag:
-0.038%
When the samples are kept in trilaminate bags the transpiration is reduced by a factor of 10
Randomization of injection order
Date | 2024-01-11
Used the Rand() function followed by the ordering A to Z
Random | Random | Random | Random
nitrate | nitrite | phosphate | silicate
BLK | BLK | BLK | BLK
BLK-IS | BLK-IS | BLK-IS | BLK-IS
MOOS237a1 | MOOS351a1 | MOOS199a2 | MOOS313a1
MOOS351a2 | MOOS313a2 | MOOS085a1 | MOOS009a2
MOOS161a1 | MOOS085a1 | MOOS085a2 | MOOS199a2
MOOS123a2 | MOOS237a2 | MOOS313a1 | MOOS047a2
MOOS161a2 | MOOS351a2 | MOOS047a2 | MOOS085a2
MOOS275a1 | MOOS047a2 | MOOS047a1 | MOOS085a1
MOOS009a2 | MOOS275a1 | MOOS351a1 | MOOS009a1
MOOS313a2 | MOOS009a2 | MOOS237a2 | MOOS161a1
MOOS199a1 | MOOS161a2 | MOOS313a2 | MOOS123a2
MOOS237a2 | MOOS313a1 | MOOS351a2 | MOOS275a2
MOOS047a2 | MOOS123a2 | MOOS275a2 | MOOS161a2
MOOS085a1 | MOOS199a1 | MOOS199a1 | MOOS237a2
MOOS199a2 | MOOS275a2 | MOOS237a1 | MOOS199a1
MOOS123a1 | MOOS123a1 | MOOS161a1 | MOOS351a2
MOOS009a1 | MOOS009a1 | MOOS123a1 | MOOS275a1
MOOS313a1 | MOOS085a2 | MOOS009a2 | MOOS237a1
MOOS275a2 | MOOS237a1 | MOOS009a1 | MOOS313a2
MOOS047a1 | MOOS161a1 | MOOS123a2 | MOOS351a1
MOOS351a1 | MOOS199a2 | MOOS161a2 | MOOS047a1
MOOS085a2 | MOOS047a1 | MOOS275a1 | MOOS123a1
Short term stability nitrate by HS ID GC-MS (verified on 2024-04-10)
Date | 2024-01-15
Sample | Vial | Sample | IS | mA | mB | rAB | rAB-NORM | Random | rAB-NORM
ID | g | g | g | g | g |  |  | nitrate | SD
BLK | 11.6247 | 13.6158 | 13.6159 | 1.9911 | 0.0001 |  |  | BLK
BLK-IS | 11.2155 | 13.1842 | 13.3835 | 1.9687 | 0.1993 |  |  | BLK-IS
MOOS009a1 | 11.4625 | 13.5021 | 13.7004 | 2.0396 | 0.1983 | 0.98799 | 0.096057 | MOOS237a1 | 0.06678%
MOOS009a2 | 11.4827 | 13.4789 | 13.6778 | 1.9962 | 0.1989 | 0.96314 | 0.095967 | MOOS351a2
MOOS047a1 | 11.4508 | 13.472 | 13.6704 | 2.0212 | 0.1984 | 0.97974 | 0.09617 | MOOS161a1 | 0.1365%
MOOS047a2 | 11.4905 | 13.4834 | 13.6813 | 1.9929 | 0.1979 | 0.96659 | 0.095985 | MOOS123a2
MOOS085a1 | 11.4681 | 13.4846 | 13.6834 | 2.0165 | 0.1988 | 0.97655 | 0.096275 | MOOS161a2 | 0.09677%
MOOS085a2 | 11.3853 | 13.3794 | 13.5779 | 1.9941 | 0.1985 | 0.96849 | 0.096407 | MOOS275a1
MOOS123a1 | 11.4635 | 13.4805 | 13.6794 | 2.017 | 0.1989 | 0.97738 | 0.096381 | MOOS009a2 | 0.1743%
MOOS123a2 | 11.5206 | 13.5173 | 13.7152 | 1.9967 | 0.1979 | 0.97483 | 0.096619 | MOOS313a2
MOOS161a1 | 11.5439 | 13.5604 | 13.7594 | 2.0165 | 0.199 | 0.97582 | 0.096299 | MOOS199a1 | 0.03379%
MOOS161a2 | 11.4995 | 13.4925 | 13.6909 | 1.993 | 0.1984 | 0.96782 | 0.096345 | MOOS237a2
MOOS199a1 | 11.5023 | 13.515 | 13.7134 | 2.0127 | 0.1984 | 0.97105 | 0.095721 | MOOS047a2 | 0.2768%
MOOS199a2 | 11.9895 | 13.9887 | 14.1877 | 1.9992 | 0.199 | 0.9654 | 0.096096 | MOOS085a1
MOOS237a1 | 11.4517 | 13.4695 | 13.6687 | 2.0178 | 0.1992 | 0.97656 | 0.096407 | MOOS199a2 | 0.2527%
MOOS237a2 | 11.4601 | 13.4535 | 13.6525 | 1.9934 | 0.199 | 0.96227 | 0.096063 | MOOS123a1
MOOS275a1 | 11.5154 | 13.5405 | 13.7397 | 2.0251 | 0.1992 | 0.97685 | 0.096089 | MOOS009a1 | 0.02145%
MOOS275a2 | 11.8887 | 13.8793 | 14.078 | 1.9906 | 0.1987 | 0.96234 | 0.09606 | MOOS313a1
MOOS313a1 | 11.5233 | 13.5402 | 13.7388 | 2.0169 | 0.1986 | 0.97966 | 0.096466 | MOOS275a2 | 0.1867%
MOOS313a2 | 11.4922 | 13.491 | 13.6897 | 1.9988 | 0.1987 | 0.96782 | 0.096211 | MOOS047a1
MOOS351a1 | 11.4319 | 13.4467 | 13.6456 | 2.0148 | 0.1989 | 0.97913 | 0.096659 | MOOS351a1 | 0.1058%
MOOS351a2 | 11.5062 | 13.5059 | 13.7045 | 1.9997 | 0.1986 | 0.9718 | 0.096515 | MOOS085a2
BLK = reagent blank with 2.0 mL of Genpure water
IS = internal standard prepared 2024-01-09 see RND Line 664
100 μL of 1 % sulfamic acid + 50 μL of Et3OBF4 aqueous solution
Estimation of the standard deviation on the single measurement
0.2492% | RSD of all 20 data
0.1212% | median of the RSD of repeted measurement
0.1583% | RSD from the factor way ANOVA
Unit | Storage, t | Days | rAB-NORM | rAB-NORM | rAB-NORM-2 | rAB-NORM-2
ID | °C | n. | value | uncertainty | value | uncertainty
MOOS009a | 21 | 12 | 0.096012 | 0.000152 | 0.99713 | 0.0019334
MOOS047a | 21 | 27 | 0.096078 | 0.00015211 | 0.99782 | 0.0019348
MOOS085a | 4 | # ref | 0.096341 | 0.00015253 | 1.0005 | 0.0019401
MOOS123a | 40 | 12 | 0.0965 | 0.00015278 | 1.0022 | 0.0019433
MOOS161a | 40 | 27 | 0.096322 | 0.0001525 | 1.0004 | 0.0019397
MOOS199a | 21 | 35 | 0.095908 | 0.00015184 | 0.99606 | 0.0019314
MOOS237a | 4 | # ref | 0.096235 | 0.00015236 | 0.99945 | 0.0019379
MOOS275a | 21 | 43 | 0.096074 | 0.0001521 | 0.99778 | 0.0019347
MOOS313a | 40 | 35 | 0.096338 | 0.00015252 | 1.0005 | 0.00194
MOOS351a | 40 | 43 | 0.096587 | 0.00015292 | 1.0031 | 0.001945
Isochronous study on nitrate at 21 °C |  |  |  |  | Isochronous study on nitrate at 40 °C
Day | rAB-NORM-2 | rAB-NORM-2 |  |  | Day | rAB-NORM-2 | rAB-NORM-2
number | value | uncertainty |  |  | number | value | uncertainty
0 | 1 | 0.001939 |  |  | 0 | 1 | 0.001939
12 | 0.99713 | 0.0019334 |  |  | 12 | 1.0022 | 0.0019433
27 | 0.99782 | 0.0019348 |  |  | 27 | 1.0004 | 0.0019397
35 | 0.99606 | 0.0019314 |  |  | 35 | 1.0005 | 0.00194
43 | 0.99778 | 0.0019347 |  |  | 43 | 1.0031 | 0.001945
Conclusion
Nitrate is stable under transportation conditions
No clear pattern in the isochronous study
Very tight uncertainty
Homogeneity of nitrate on the base of the stability data by ID HS GC-MS (verified on 2024-04-10)
Nitrate ANOVA experiment |  |  |  | Anova: Single Factor
Unit | rAB-NORM-I | rAB-NORM-II |  | SUMMARY
ID |  |  |  | Groups | Count | Sum | Average | Variance
MOOS009a | 0.096057 | 0.095967 |  | MOOS009a | 2 | 0.19202 | 0.096012 | 4.1105e-09
MOOS047a | 0.09617 | 0.095985 |  | MOOS047a | 2 | 0.19216 | 0.096078 | 1.7193e-08
MOOS085a | 0.096275 | 0.096407 |  | MOOS085a | 2 | 0.19268 | 0.096341 | 8.6925e-09
MOOS123a | 0.096381 | 0.096619 |  | MOOS123a | 2 | 0.193 | 0.0965 | 2.8275e-08
MOOS161a | 0.096299 | 0.096345 |  | MOOS161a | 2 | 0.19264 | 0.096322 | 1.0594e-09
MOOS199a | 0.095721 | 0.096096 |  | MOOS199a | 2 | 0.19182 | 0.095908 | 7.046e-08
MOOS237a | 0.096407 | 0.096063 |  | MOOS237a | 2 | 0.19247 | 0.096235 | 5.9138e-08
MOOS275a | 0.096089 | 0.09606 |  | MOOS275a | 2 | 0.19215 | 0.096074 | 4.2459e-10
MOOS313a | 0.096466 | 0.096211 |  | MOOS313a | 2 | 0.19268 | 0.096338 | 3.236e-08
MOOS351a | 0.096659 | 0.096515 |  | MOOS351a | 2 | 0.19317 | 0.096587 | 1.0437e-08
ANOVA
|  |  |  | Source of Variation | SS | df | MS | F | P-value | F crit
Between groups: Mb |  |  |  | Between Groups | 8.6072e-07 | 9 | 9.5636e-08 | 4.1195 | 0.018803 | 3.0204
Within groups: Mw |  |  |  | Within Groups | 2.3215e-07 | 10 | 2.3215e-08
|  |  |  | Total | 1.0929e-06 | 19
Ellison et al. textbook P65
F>Fcrit, there is a small homogeneiety component
Homogeneiety uncertainty component on nitrate
uR | 0.1977% | k = 1
Method uncertainty component on nitrate
uR | 0.1583% | k = 1
Nitrate DSL (R MetRology) experiment
Unit | rAB-NORM-I | rAB-NORM-II | Average | Uncertainty
ID
MOOS009a | 0.096057 | 0.095967 | 0.096012 | 0.000152
MOOS047a | 0.09617 | 0.095985 | 0.096078 | 0.00015211
MOOS085a | 0.096275 | 0.096407 | 0.096341 | 0.00015253
MOOS123a | 0.096381 | 0.096619 | 0.0965 | 0.00015278
MOOS161a | 0.096299 | 0.096345 | 0.096322 | 0.0001525
MOOS199a | 0.095721 | 0.096096 | 0.095908 | 0.00015184
MOOS237a | 0.096407 | 0.096063 | 0.096235 | 0.00015236
MOOS275a | 0.096089 | 0.09606 | 0.096074 | 0.0001521
MOOS313a | 0.096466 | 0.096211 | 0.096338 | 0.00015252
MOOS351a | 0.096659 | 0.096515 | 0.096587 | 0.00015292
Run on 2024-01-17
x | 9.60119751946321E-02, 9.60776350294193E-02, 9.63406331078862E-02, 9.65000587810263E-02, 0.096322319816342, 0.095908402452984, 9.62352249447595E-02, 9.60741454760132E-02, 9.63383174209506E-02, 9.65868661714798E-02
s | 1.52004846717576E-04, 1.52108798470491E-04, 1.52525173433336E-04, 1.52777573979824E-04, 1.52496180080492E-04, 1.51840871768759E-04, 1.52358292670323E-04, 1.52103273857238E-04, 1.52521507269433E-04, 1.52915006253799E-04
n | 2
$mu
0.096239
$s
6.9142e-05
$tau
0.00019026
Homogeneiety uncertainty component on nitrate
> dsl(x, s, n)$method.details$tau/dsl(x, s, n)$method.details$mu
0.1977%
Short term stability nitrite by UV-Vis (verified 2024-04-10)
Date | 2024-01-15
Sample | Vial | Sample | Chemistry | mS | mF | A | ANORM | Random | ANORM
ID | g | g | g | g | g |  |  | nitrite | SD
BLK1 | 6.1216 | 11.2235 | 11.4273 | 5.1019 | 5.3057 |  | 0 | BLK
BLK2 | 6.1492 | 11.2619 | 11.4627 | 5.1127 | 5.3135 |  | 0 | BLK-IS
MOOS009a1 | 6.1372 | 11.2325 | 11.4335 | 5.0953 | 5.2963 | 0.085949 | 0.08934 | MOOS351a1 | 0.1391%
MOOS009a2 | 6.1356 | 11.2302 | 11.431 | 5.0946 | 5.2954 | 0.086121 | 0.089515 | MOOS313a2
MOOS047a1 | 6.1213 | 11.2185 | 11.4202 | 5.0972 | 5.2989 | 0.085613 | 0.089001 | MOOS085a1 | 0.1122%
MOOS047a2 | 6.1227 | 11.2094 | 11.4092 | 5.0867 | 5.2865 | 0.085773 | 0.089142 | MOOS237a2
MOOS085a1 | 6.1416 | 11.2356 | 11.4362 | 5.094 | 5.2946 | 0.084193 | 0.087508 | MOOS351a2 | 0.9107%
MOOS085a2 | 6.15 | 11.2538 | 11.4545 | 5.1038 | 5.3045 | 0.085289 | 0.088643 | MOOS047a2
MOOS123a1 | 6.1194 | 11.2103 | 11.4114 | 5.0909 | 5.292 | 0.086719 | 0.090145 | MOOS275a1 | 0.6535%
MOOS123a2 | 6.1526 | 11.2505 | 11.4522 | 5.0979 | 5.2996 | 0.085916 | 0.089315 | MOOS009a2
MOOS161a1 | 6.1159 | 11.2166 | 11.4174 | 5.1007 | 5.3015 | 0.08647 | 0.089874 | MOOS161a2 | 0.2758%
MOOS161a2 | 6.0978 | 11.1862 | 11.3874 | 5.0884 | 5.2896 | 0.086119 | 0.089524 | MOOS313a1
MOOS199a1 | 6.1394 | 11.2445 | 11.4457 | 5.1051 | 5.3063 | 0.086315 | 0.089717 | MOOS123a2 | 0.6475%
MOOS199a2 | 6.1634 | 11.2654 | 11.4668 | 5.102 | 5.3034 | 0.085523 | 0.088899 | MOOS199a1
MOOS237a1 | 6.1512 | 11.2542 | 11.4557 | 5.103 | 5.3045 | 0.085775 | 0.089162 | MOOS275a2 | 0.8909%
MOOS237a2 | 6.12 | 11.207 | 11.4085 | 5.087 | 5.2885 | 0.084691 | 0.088046 | MOOS123a1
MOOS275a1 | 6.1324 | 11.2435 | 11.4448 | 5.1111 | 5.3124 | 0.086173 | 0.089567 | MOOS009a1 | 0.1385%
MOOS275a2 | 6.1383 | 11.2462 | 11.4486 | 5.1079 | 5.3103 | 0.086322 | 0.089743 | MOOS085a2
MOOS313a1 | 6.1219 | 11.2111 | 11.4113 | 5.0892 | 5.2894 | 0.086039 | 0.089424 | MOOS237a1 | 0.9823%
MOOS313a2 | 6.1232 | 11.2301 | 11.431 | 5.1069 | 5.3078 | 0.084852 | 0.08819 | MOOS161a1
MOOS351a1 | 6.1194 | 11.2185 | 11.4189 | 5.0991 | 5.2995 | 0.084748 | 0.088079 | MOOS199a2 | 0.4663%
MOOS351a2 | 6.1305 | 11.2447 | 11.4463 | 5.1142 | 5.3158 | 0.085299 | 0.088661 | MOOS047a1
BLK = reagent blank with LNSW
Zero the instrument against cuvette + GenPure water
Estimation of the standard deviation on the single measurement
0.7822% | RSD of all 20 data
0.5559% | median of the RSD of repeted measurement
0.6126% | RSD from the factor way ANOVA
Unit | Storage, t | Days | ANORM | ANORM | ANORM-2 | ANORM-2
ID | °C | n. | value | uncertainty | value | uncertainty
MOOS009a | 21 | 12 | 0.089427 | 0.00054783 | 1.0123 | 0.0075951
MOOS047a | 21 | 27 | 0.089071 | 0.00054565 | 1.0083 | 0.0075649
MOOS085a | 4 | # ref | 0.088076 | 0.00053955 | 0.99701 | 0.0074803
MOOS123a | 40 | 12 | 0.08973 | 0.00054968 | 1.0157 | 0.0076208
MOOS161a | 40 | 27 | 0.089699 | 0.00054949 | 1.0154 | 0.0076182
MOOS199a | 21 | 35 | 0.089308 | 0.00054709 | 1.011 | 0.0075849
MOOS237a | 4 | # ref | 0.088604 | 0.00054278 | 1.003 | 0.0075251
MOOS275a | 21 | 43 | 0.089655 | 0.00054922 | 1.0149 | 0.0076144
MOOS313a | 40 | 35 | 0.088807 | 0.00054402 | 1.0053 | 0.0075424
MOOS351a | 40 | 43 | 0.08837 | 0.00054135 | 1.0003 | 0.0075053
Isochronous study on nitrite at 21 °C |  |  |  |  | Isochronous study on nitrite at 40 °C
Day | ANORM-2 | ANORM-2 |  |  | Day | ANORM-2 | ANORM-2
number | value | uncertainty |  |  | number | value | uncertainty
0 | 1 | 0.0075027 |  |  | 0 | 1 | 0.0075027
12 | 1.0123 | 0.0075951 |  |  | 12 | 1.0157 | 0.0076208
27 | 1.0083 | 0.0075649 |  |  | 27 | 1.0154 | 0.0076182
35 | 1.011 | 0.0075849 |  |  | 35 | 1.0053 | 0.0075424
43 | 1.0149 | 0.0076144 |  |  | 43 | 1.0003 | 0.0075053
Conclusion
Nitrite is stable under transportation conditions
No clear pattern in the isochronous study
Homogeneity of nitrite on the base of the stability data by UV-vis (verified on 2024-04-10)
Nitrite ANOVA experiment |  |  |  | Anova: Single Factor
Unit | ANORM-I | ANORM-II |  | SUMMARY
ID |  |  |  | Groups | Count | Sum | Average | Variance
MOOS009a | 0.08934 | 0.089515 |  | MOOS009a | 2 | 0.17885 | 0.089427 | 1.5465e-08
MOOS047a | 0.089001 | 0.089142 |  | MOOS047a | 2 | 0.17814 | 0.089071 | 9.9827e-09
MOOS085a | 0.087508 | 0.088643 |  | MOOS085a | 2 | 0.17615 | 0.088076 | 6.4341e-07
MOOS123a | 0.090145 | 0.089315 |  | MOOS123a | 2 | 0.17946 | 0.08973 | 3.4384e-07
MOOS161a | 0.089874 | 0.089524 |  | MOOS161a | 2 | 0.1794 | 0.089699 | 6.1199e-08
MOOS199a | 0.089717 | 0.088899 |  | MOOS199a | 2 | 0.17862 | 0.089308 | 3.3441e-07
MOOS237a | 0.089162 | 0.088046 |  | MOOS237a | 2 | 0.17721 | 0.088604 | 6.2305e-07
MOOS275a | 0.089567 | 0.089743 |  | MOOS275a | 2 | 0.17931 | 0.089655 | 1.5416e-08
MOOS313a | 0.089424 | 0.08819 |  | MOOS313a | 2 | 0.17761 | 0.088807 | 7.6092e-07
MOOS351a | 0.088079 | 0.088661 |  | MOOS351a | 2 | 0.17674 | 0.08837 | 1.6981e-07
ANOVA
|  |  |  | Source of Variation | SS | df | MS | F | P-value | F crit
Between groups: Mb |  |  |  | Between Groups | 6.2451e-06 | 9 | 6.939e-07 | 2.3305 | 0.10185 | 3.0204
Within groups: Mw |  |  |  | Within Groups | 2.9775e-06 | 10 | 2.9775e-07
|  |  |  | Total | 9.2226e-06 | 19
Ellison et al. textbook P65
F<Fcrit, single factor anova does not find an uncertainty component
Homogeneiety uncertainty component on nitrite
uR | 0.4996% | k = 1
Method uncertainty component on nitrite
uR | 0.6126% | k = 1
Nitrite DSL (R MetRology) experiment
Unit | ANORM-I | ANORM-II | Average | Uncertainty
ID
MOOS009a | 0.08934 | 0.089515 | 0.089427 | 0.00054783
MOOS047a | 0.089001 | 0.089142 | 0.089071 | 0.00054565
MOOS085a | 0.087508 | 0.088643 | 0.088076 | 0.00053955
MOOS123a | 0.090145 | 0.089315 | 0.08973 | 0.00054968
MOOS161a | 0.089874 | 0.089524 | 0.089699 | 0.00054949
MOOS199a | 0.089717 | 0.088899 | 0.089308 | 0.00054709
MOOS237a | 0.089162 | 0.088046 | 0.088604 | 0.00054278
MOOS275a | 0.089567 | 0.089743 | 0.089655 | 0.00054922
MOOS313a | 0.089424 | 0.08819 | 0.088807 | 0.00054402
MOOS351a | 0.088079 | 0.088661 | 0.08837 | 0.00054135
Run on 2024-01-17
x | 8.94274618646746E-02, 8.90714197921282E-02, 8.80756829735431E-02, 8.97299273750277E-02, 8.96991506418348E-02, 8.93079028410775E-02, 8.86038185199901E-02, 8.96547062081091E-02, 8.88068035670311E-02, 8.83700714044544E-02
s | 5.47826519733884E-04, 5.45645430330081E-04, 5.39545614630047E-04, 5.49679402779217E-04, 5.4949086661502E-04, 5.47094109326227E-04, 5.42780936893756E-04, 5.49218603051437E-04, 5.44024409419532E-04, 5.41349018038724E-04
n | 2
$mu
0.089072
$s
0.00018673
$tau
0.00044699
Homogeneiety uncertainty component on nitrite
> dsl(x, s, n)$method.details$tau/dsl(x, s, n)$method.details$mu
0.5018%
Short term stability phosphate by UV-Vis (verified on 2024-04-10)
Date | 2024-01-15
Sample | Vial | Sample | Chemistry | mS | mF | A | ANORM | Random | ANORM
ID | g | g | g | g | g |  |  | phosphate | SD
BLK1 | 6.1173 | 12.2803 | 13.0634 | 6.163 | 6.9461 |  | 0 | BLK
BLK2 | 6.1121 | 12.281 | 13.0669 | 6.1689 | 6.9548 |  | 0 | BLK-IS
MOOS009a1 | 6.1238 | 12.2575 | 13.0417 | 6.1337 | 6.9179 | 0.26948 | 0.30393 | MOOS199a2 | 0.3367%
MOOS009a2 | 6.1337 | 12.2771 | 13.0616 | 6.1434 | 6.9279 | 0.2708 | 0.30538 | MOOS085a1
MOOS047a1 | 6.1306 | 12.2754 | 13.0601 | 6.1448 | 6.9295 | 0.27071 | 0.30528 | MOOS085a2 | 0.005636%
MOOS047a2 | 6.1397 | 12.2957 | 13.0812 | 6.156 | 6.9415 | 0.27071 | 0.30525 | MOOS313a1
MOOS085a1 | 6.1233 | 12.2804 | 13.0661 | 6.1571 | 6.9428 | 0.27006 | 0.30452 | MOOS047a2 | 0.1358%
MOOS085a2 | 6.1397 | 12.297 | 13.0824 | 6.1573 | 6.9427 | 0.27059 | 0.30511 | MOOS047a1
MOOS123a1 | 6.1282 | 12.2701 | 13.0558 | 6.1419 | 6.9276 | 0.27003 | 0.30457 | MOOS351a1 | 0.162%
MOOS123a2 | 6.1351 | 12.2805 | 13.0684 | 6.1454 | 6.9333 | 0.27058 | 0.30527 | MOOS237a2
MOOS161a1 | 6.1082 | 12.2618 | 13.0403 | 6.1536 | 6.9321 | 0.27256 | 0.30704 | MOOS313a2 | 0.156%
MOOS161a2 | 6.1277 | 12.2744 | 13.0595 | 6.1467 | 6.9318 | 0.27167 | 0.30637 | MOOS351a2
MOOS199a1 | 6.1593 | 12.3029 | 13.0876 | 6.1436 | 6.9283 | 0.27053 | 0.30508 | MOOS275a2 | 0.09017%
MOOS199a2 | 6.1483 | 12.2875 | 13.0708 | 6.1392 | 6.9225 | 0.27091 | 0.30547 | MOOS199a1
MOOS237a1 | 6.1473 | 12.302 | 13.0866 | 6.1547 | 6.9393 | 0.27094 | 0.30547 | MOOS237a1 | 0.03485%
MOOS237a2 | 6.1476 | 12.3054 | 13.0911 | 6.1578 | 6.9435 | 0.27077 | 0.30532 | MOOS161a1
MOOS275a1 | 6.1419 | 12.2659 | 13.0512 | 6.124 | 6.9093 | 0.27059 | 0.30528 | MOOS123a1 | 0.04899%
MOOS275a2 | 6.1493 | 12.2879 | 13.0735 | 6.1386 | 6.9242 | 0.27083 | 0.30549 | MOOS009a2
MOOS313a1 | 6.1309 | 12.275 | 13.0608 | 6.1441 | 6.9299 | 0.2727 | 0.30757 | MOOS009a1 | 0.1193%
MOOS313a2 | 6.1367 | 12.3061 | 13.0932 | 6.1694 | 6.9565 | 0.27231 | 0.30705 | MOOS123a2
MOOS351a1 | 6.1299 | 12.274 | 13.0574 | 6.1441 | 6.9275 | 0.27294 | 0.30774 | MOOS161a2 | 0.08146%
MOOS351a2 | 6.1081 | 12.2705 | 13.0563 | 6.1624 | 6.9482 | 0.27325 | 0.3081 | MOOS275a1
BLK = reagent blank with LNSW
Zero the instrument against cuvette + GenPure water
Estimation of the standard deviation on the single measurement
0.3757% | RSD of all 20 data
0.105% | median of the RSD of repeted measurement
0.1464% | RSD from the factor way ANOVA
Unit | Storage, t | Days | ANORM | ANORM | ANORM-2 | ANORM-2
ID | °C | n. | value | uncertainty | value | uncertainty
MOOS009a | 21 | 12 | 0.30466 | 0.00044603 | 0.99852 | 0.0017904
MOOS047a | 21 | 27 | 0.30526 | 0.00044692 | 1.0005 | 0.001794
MOOS085a | 4 | # ref | 0.30481 | 0.00044626 | 0.99904 | 0.0017914
MOOS123a | 40 | 12 | 0.30492 | 0.00044642 | 0.9994 | 0.001792
MOOS161a | 40 | 27 | 0.30671 | 0.00044903 | 1.0052 | 0.0018025
MOOS199a | 21 | 35 | 0.30528 | 0.00044694 | 1.0006 | 0.0017941
MOOS237a | 4 | # ref | 0.3054 | 0.00044712 | 1.001 | 0.0017948
MOOS275a | 21 | 43 | 0.30539 | 0.0004471 | 1.0009 | 0.0017947
MOOS313a | 40 | 35 | 0.30731 | 0.00044992 | 1.0072 | 0.0018061
MOOS351a | 40 | 43 | 0.30792 | 0.00045081 | 1.0092 | 0.0018096
Isochronous study on phosphate at 21 °C |  |  |  |  | Isochronous study on phosphate at 40 °C
Day | ANORM-2 | ANORM-2 |  |  | Day | ANORM-2 | ANORM-2
number | value | uncertainty |  |  | number | value | uncertainty
0 | 1 | 0.0017931 |  |  | 0 | 1 | 0.0017931
12 | 0.99852 | 0.0017904 |  |  | 12 | 0.9994 | 0.001792
27 | 1.0005 | 0.001794 |  |  | 27 | 1.0052 | 0.0018025
35 | 1.0006 | 0.0017941 |  |  | 35 | 1.0072 | 0.0018061
43 | 1.0009 | 0.0017947 |  |  | 43 | 1.0092 | 0.0018096
Conclusion
Phosphate shows a trend at 40 °C which will be considered to estimate the stability under storage conditions
On the base of Q10 calculated in cell:
[ESM_2.xlsm]LONG!$A$83
it was possible to calculate an uncertainty component (short-term, transportation) equivalent to 10 days at 30 °C
v (30 °C) | 6.085e-05
uR,SHORT | 0.06085% | Covering a 10 days period (k = 1)
Homogeneity of phosphate on the base of the stability data by UV-vis (verified on 2024-04-10)
Phosphate ANOVA experiment |  |  |  | Anova: Single Factor
Unit | ANORM-I | ANORM-II |  | SUMMARY
ID |  |  |  | Groups | Count | Sum | Average | Variance
MOOS009a | 0.30393 | 0.30538 |  | MOOS009a | 2 | 0.60931 | 0.30466 | 1.0523e-06
MOOS047a | 0.30528 | 0.30525 |  | MOOS047a | 2 | 0.61053 | 0.30526 | 2.9603e-10
MOOS085a | 0.30452 | 0.30511 |  | MOOS085a | 2 | 0.60963 | 0.30481 | 1.7145e-07
MOOS123a | 0.30457 | 0.30527 |  | MOOS123a | 2 | 0.60985 | 0.30492 | 2.4405e-07
MOOS161a | 0.30704 | 0.30637 |  | MOOS161a | 2 | 0.61341 | 0.30671 | 2.2904e-07
MOOS199a | 0.30508 | 0.30547 |  | MOOS199a | 2 | 0.61055 | 0.30528 | 7.5778e-08
MOOS237a | 0.30547 | 0.30532 |  | MOOS237a | 2 | 0.6108 | 0.3054 | 1.133e-08
MOOS275a | 0.30528 | 0.30549 |  | MOOS275a | 2 | 0.61078 | 0.30539 | 2.238e-08
MOOS313a | 0.30757 | 0.30705 |  | MOOS313a | 2 | 0.61463 | 0.30731 | 1.3445e-07
MOOS351a | 0.30774 | 0.3081 |  | MOOS351a | 2 | 0.61584 | 0.30792 | 6.2919e-08
ANOVA
|  |  |  | Source of Variation | SS | df | MS | F | P-value | F crit
Between groups: Mb |  |  |  | Between Groups | 2.3071e-05 | 9 | 2.5634e-06 | 12.7918 | 0.00021906 | 3.0204
Within groups: Mw |  |  |  | Within Groups | 2.004e-06 | 10 | 2.004e-07
|  |  |  | Total | 2.5075e-05 | 19
Ellison et al. textbook P65
F>Fcrit, single factor anova finds an uncertainty component
Homogeneiety uncertainty component on phosphate
uR | 0.3555% | k = 1
Method uncertainty component on phosphate
uR | 0.1464% | k = 1
Phosphate DLS (R MetRology) experiment
Unit | ANORM-I | ANORM-II | Average | Uncertainty
ID
MOOS009a | 0.30393 | 0.30538 | 0.30466 | 0.00044603
MOOS047a | 0.30528 | 0.30525 | 0.30526 | 0.00044692
MOOS085a | 0.30452 | 0.30511 | 0.30481 | 0.00044626
MOOS123a | 0.30457 | 0.30527 | 0.30492 | 0.00044642
MOOS161a | 0.30704 | 0.30637 | 0.30671 | 0.00044903
MOOS199a | 0.30508 | 0.30547 | 0.30528 | 0.00044694
MOOS237a | 0.30547 | 0.30532 | 0.3054 | 0.00044712
MOOS275a | 0.30528 | 0.30549 | 0.30539 | 0.0004471
MOOS313a | 0.30757 | 0.30705 | 0.30731 | 0.00044992
MOOS351a | 0.30774 | 0.3081 | 0.30792 | 0.00045081
Run on 2024-01-17
x | 0.304655270575334, 0.305263388296773, 0.304813689854298, 0.304922794485598, 0.306705767142286, 0.305276222348644, 0.305399604517702, 0.305388757594469, 0.307313192683291, 0.307919556839951
s | 4.46030953237752E-04, 4.46921268794945E-04, 4.46262887193379E-04, 4.46422621973642E-04, 4.49032985458108E-04, 4.46940058505621E-04, 4.47120696333989E-04, 4.47104815881632E-04, 4.49922287627618E-04, 4.50810035875968E-04
n | 2
$mu
0.30577
$s
0.00035686
$tau
0.0010832
Homogeneiety uncertainty component on phosphate
> dsl(x, s, n)$method.details$tau/dsl(x, s, n)$method.details$mu
0.3543%
2024-06-25 internal audit
Considering the trend obsered in the phosphate data, value in cell
$A$522
was not used as an estimate for u(hom). Instead the result from cell
[ESM_2.xlsm]CHAR_HOM!$A$1338
was used
Short term stability silicate by UV-Vis (verified on 2024-04-10)
Date | 2024-01-15
Sample | Vial | Sample | Chemistry | mS | mF | A | ANORM | Random | ANORM
ID | g | g | g | g | g |  |  | silicate | SD
BLK1 | 6.1527 | 9.2263 | 9.5499 | 3.0736 | 3.3972 |  | 0 | BLK
BLK2 | 6.1311 | 9.194 | 9.5163 | 3.0629 | 3.3852 |  | 0 | BLK-IS
MOOS009a1 | 6.1562 | 9.2044 | 9.5275 | 3.0482 | 3.3713 | 0.11918 | 0.13181 | MOOS313a1 | 1.595%
MOOS009a2 | 6.1494 | 9.2067 | 9.5284 | 3.0573 | 3.379 | 0.12198 | 0.13482 | MOOS009a2
MOOS047a1 | 6.1138 | 9.1729 | 9.4956 | 3.0591 | 3.3818 | 0.12144 | 0.13425 | MOOS199a2 | 0.3316%
MOOS047a2 | 6.1209 | 9.1833 | 9.5064 | 3.0624 | 3.3855 | 0.12087 | 0.13362 | MOOS047a2
MOOS085a1 | 6.1343 | 9.1855 | 9.508 | 3.0512 | 3.3737 | 0.12169 | 0.13456 | MOOS085a2 | 0.818%
MOOS085a2 | 6.1164 | 9.1772 | 9.5005 | 3.0608 | 3.3841 | 0.1203 | 0.13301 | MOOS085a1
MOOS123a1 | 6.163 | 9.227 | 9.5506 | 3.064 | 3.3876 | 0.12115 | 0.13395 | MOOS009a1 | 0.08507%
MOOS123a2 | 6.1231 | 9.1864 | 9.5089 | 3.0633 | 3.3858 | 0.12104 | 0.13379 | MOOS161a1
MOOS161a1 | 6.1422 | 9.2016 | 9.524 | 3.0594 | 3.3818 | 0.12055 | 0.13325 | MOOS123a2 | 1.023%
MOOS161a2 | 6.13 | 9.1969 | 9.5198 | 3.0669 | 3.3898 | 0.11883 | 0.13134 | MOOS275a2
MOOS199a1 | 6.1654 | 9.226 | 9.5487 | 3.0606 | 3.3833 | 0.12064 | 0.13336 | MOOS161a2 | 0.9862%
MOOS199a2 | 6.1455 | 9.2105 | 9.5325 | 3.065 | 3.387 | 0.11901 | 0.13152 | MOOS237a2
MOOS237a1 | 6.1324 | 9.2031 | 9.5256 | 3.0707 | 3.3932 | 0.1204 | 0.13304 | MOOS199a1 | 0.2012%
MOOS237a2 | 6.1331 | 9.1964 | 9.5197 | 3.0633 | 3.3866 | 0.12 | 0.13266 | MOOS351a2
MOOS275a1 | 6.1552 | 9.2174 | 9.5402 | 3.0622 | 3.385 | 0.11773 | 0.13014 | MOOS275a1 | 1.293%
MOOS275a2 | 6.1092 | 9.1768 | 9.4988 | 3.0676 | 3.3896 | 0.11995 | 0.13254 | MOOS237a1
MOOS313a1 | 6.141 | 9.1951 | 9.5174 | 3.0541 | 3.3764 | 0.12022 | 0.1329 | MOOS313a2 | 0.6406%
MOOS313a2 | 6.1368 | 9.2016 | 9.5238 | 3.0648 | 3.387 | 0.11918 | 0.1317 | MOOS351a1
MOOS351a1 | 6.1194 | 9.1767 | 9.4997 | 3.0573 | 3.3803 | 0.11916 | 0.13174 | MOOS047a1 | 0.2864%
MOOS351a2 | 6.1288 | 9.1935 | 9.5158 | 3.0647 | 3.387 | 0.11969 | 0.13228 | MOOS123a1
BLK = reagent blank with LNSW
Zero the instrument against cuvette + GenPure water
Estimation of the standard deviation on the single measurement
0.9036% | RSD of all 20 data
0.731% | median of the RSD of repeted measurement
0.8678% | RSD from the factor way ANOVA
Unit | Storage, t | Days | ANORM | ANORM | ANORM-2 | ANORM-2
ID | °C | n. | value | uncertainty | value | uncertainty
MOOS009a | 21 | 12 | 0.13332 | 0.0011569 | 1 | 0.010628
MOOS047a | 21 | 27 | 0.13393 | 0.0011622 | 1.0046 | 0.010677
MOOS085a | 4 | # ref | 0.13378 | 0.001161 | 1.0035 | 0.010665
MOOS123a | 40 | 12 | 0.13387 | 0.0011617 | 1.0041 | 0.010672
MOOS161a | 40 | 27 | 0.1323 | 0.0011481 | 0.99233 | 0.010547
MOOS199a | 21 | 35 | 0.13244 | 0.0011493 | 0.99341 | 0.010558
MOOS237a | 4 | # ref | 0.13285 | 0.0011529 | 0.99652 | 0.010591
MOOS275a | 21 | 43 | 0.13134 | 0.0011398 | 0.98516 | 0.010471
MOOS313a | 40 | 35 | 0.1323 | 0.0011481 | 0.99239 | 0.010547
MOOS351a | 40 | 43 | 0.13201 | 0.0011456 | 0.99021 | 0.010524
Isochronous study on silicate at 21 °C |  |  |  |  | Isochronous study on silicate at 40 °C
Day | ANORM-2 | ANORM-2 |  |  | Day | ANORM-2 | ANORM-2
number | value | uncertainty |  |  | number | value | uncertainty
0 | 1 | 0.010628 |  |  | 0 | 1 | 0.010628
12 | 1 | 0.010628 |  |  | 12 | 1.0041 | 0.010672
27 | 1.0046 | 0.010677 |  |  | 27 | 0.99233 | 0.010547
35 | 0.99341 | 0.010558 |  |  | 35 | 0.99239 | 0.010547
43 | 0.98516 | 0.010471 |  |  | 43 | 0.99021 | 0.010524
Conclusion
Silicate is stable under transportation conditions
No clear pattern in the isochronous study
Homogeneity of silicate on the base of the stability data by UV-vis (verified on 2024-04-10)
Silicate ANOVA experiment |  |  |  | Anova: Single Factor
Unit | ANORM-I | ANORM-II |  | SUMMARY
ID |  |  |  | Groups | Count | Sum | Average | Variance
MOOS009a | 0.13181 | 0.13482 |  | MOOS009a | 2 | 0.26663 | 0.13332 | 4.5206e-06
MOOS047a | 0.13425 | 0.13362 |  | MOOS047a | 2 | 0.26786 | 0.13393 | 1.9724e-07
MOOS085a | 0.13456 | 0.13301 |  | MOOS085a | 2 | 0.26756 | 0.13378 | 1.1975e-06
MOOS123a | 0.13395 | 0.13379 |  | MOOS123a | 2 | 0.26774 | 0.13387 | 1.297e-08
MOOS161a | 0.13325 | 0.13134 |  | MOOS161a | 2 | 0.26459 | 0.1323 | 1.8331e-06
MOOS199a | 0.13336 | 0.13152 |  | MOOS199a | 2 | 0.26488 | 0.13244 | 1.7058e-06
MOOS237a | 0.13304 | 0.13266 |  | MOOS237a | 2 | 0.26571 | 0.13285 | 7.1448e-08
MOOS275a | 0.13014 | 0.13254 |  | MOOS275a | 2 | 0.26268 | 0.13134 | 2.8838e-06
MOOS313a | 0.1329 | 0.1317 |  | MOOS313a | 2 | 0.26461 | 0.1323 | 7.1839e-07
MOOS351a | 0.13174 | 0.13228 |  | MOOS351a | 2 | 0.26402 | 0.13201 | 1.4298e-07
ANOVA
|  |  |  | Source of Variation | SS | df | MS | F | P-value | F crit
Between groups: Mb |  |  |  | Between Groups | 1.4079e-05 | 9 | 1.5643e-06 | 1.1776 | 0.39866 | 3.0204
Within groups: Mw |  |  |  | Within Groups | 1.3284e-05 | 10 | 1.3284e-06
|  |  |  | Total | 2.7362e-05 | 19
Ellison et al. textbook P65
F<Fcrit, single factor anova does not find an uncertainty component
Homogeneiety uncertainty component on silicate
uR | 0.2586% | k = 1
Method uncertainty component on silicate
uR | 0.8678% | k = 1
Silicate DLS (R MetRology) experiment
Unit | ANORM-I | ANORM-II | Average | Uncertainty
ID
MOOS009a | 0.13181 | 0.13482 | 0.13332 | 0.0011569
MOOS047a | 0.13425 | 0.13362 | 0.13393 | 0.0011622
MOOS085a | 0.13456 | 0.13301 | 0.13378 | 0.001161
MOOS123a | 0.13395 | 0.13379 | 0.13387 | 0.0011617
MOOS161a | 0.13325 | 0.13134 | 0.1323 | 0.0011481
MOOS199a | 0.13336 | 0.13152 | 0.13244 | 0.0011493
MOOS237a | 0.13304 | 0.13266 | 0.13285 | 0.0011529
MOOS275a | 0.13014 | 0.13254 | 0.13134 | 0.0011398
MOOS313a | 0.1329 | 0.1317 | 0.1323 | 0.0011481
MOOS351a | 0.13174 | 0.13228 | 0.13201 | 0.0011456
Run on 2024-01-17
x | 0.1333172613285, 0.133930960848318, 0.133781694456943, 0.133867873554469, 0.132295113360174, 0.132438583342709, 0.132852676090355, 0.131339020215813, 0.132303100262537, 0.132012063411144
s | 1.1569234362299E-03, 1.16224910336561E-03, 1.16095377382837E-03, 1.16170163360773E-03, 1.14805326497025E-03, 1.14929829343495E-03, 1.15289177862772E-03, 1.1397563156112E-03, 1.14812257508385E-03, 1.1455969654904E-03
n | 2
$mu
0.1328
$s
0.0002798
$tau
0.00034459
Homogeneiety uncertainty component on silicate
> dsl(x, s, n)$method.details$tau/dsl(x, s, n)$method.details$mu
0.2595%

## CHAR_HOM
Index
Line | Description
32 | Number of MonteCarlo simulations for uncertainty evaluation set at 2000
35 | Uncertainty on mass measurements (verified on 2024-04-10)
50 | Determination of density MOOS-4, LNWS, SALT-1, MOOS-3 (verified on 2024-04-10)
74 | Sequences for the determination of nutrients in MOOS-4 week-1
109 | Sequences for the determination of nutrients in MOOS-4 week-2
143 | Planning the quantitative analysis: preparation of the calibration standards
155 | Primary standards and atomic weights (verified on 2024-04-10)
216 | Preparation of the SALT-1 QC sample (verified 2024-04-10)
239 | Weight of the MOOS-4 units
257 | Matrix effect for nitrite, phosphate, silicate by UV-vis (MOOS-4 matrix vs LNSW, verified 2024-05-03)
270 | 2024-02-20 and 2024-02-21 Characterization MOOS-4 [WEEK 1; DAY 1 and 2]: sample preparation (verified 2024-04-11)
545 | 2024-02-20 and 2024-02-21 Characterization MOOS-4 [WEEK 1; DAY 1 and 2]: data analysis (verified 2024-04-19)
762 | 2024-03-05 and 2024-03-06 Characterization MOOS-4 [WEEK 2; DAY 1 and 2]: sample preparation (verified 2024-04-19)
927 | 2024-03-05 and 2024-03-06 Characterization MOOS-4 [WEEK 2; DAY 1 and 2]: data analysis (verified 2024-04-19)
1074 | EtONO and EtONO2 signals on MOOS-4 without derivatization
1082 | Difference between isotopic composition of sample (RA) and standard (RA*)
1120 | 2024-03-21 Characterization MOOS-4 [Silicate LC-ICPMS Zuzana]: sample preparation (verified 2024-04-24)
1160 | 2024-03-21 Characterization MOOS-4 [Silicate LC-ICPMS Zuzana]: data analysis (verified 2024-05-10)
1189 | CHAR summary nitrite (verified 2024-05-08)
1230 | CHAR summary nitrate (verified 2024-05-08)
1255 | CHAR summary phosphate (verified 2024-05-08)
1281 | CHAR summary silicate (verified 2024-05-08)
1312 | HOM data phosphate (verified 2024-05-09)
1340 | HOM data nitrite (verified 2024-05-10)
1368 | HOM data silicate (verified 2024-05-10)
Number of MonteCarlo simulations for uncertainty evaluation set at 2000
2000
Uncertainty on mass measurements (verified on 2024-04-10)
Balance: MS204S in B5
um | 0.0002 | g | from 0.1 g to 200 g
When measuring the mass of a cold solution the possible effect of condensation on the bottle should be accounted
See RND line number: 30
um | 0.002 | g | from 0.1 g to 200 g
uncertainty on the difference between two masses m1-m0
u(m1-m0) | 0.00028284 | g | from 0.1 g to 200 g
Determination of density MOOS-4, LNWS, SALT-1, MOOS-3 (verified on 2024-04-10)
Percentage variance components: First-order Taylor expansion
Sample | Date | Analyst | Balance | t | Rel Hum | m0 (load) | m0 (unload) | m1 (load) | m1 (unload) | m2 (load) | m2 (unload) | ρH2O | u[ρH2O] | ρair | u[ρair] | m0 | u[m0] | m1 | u[m1] | m2 | u[m2] | ρsample | u[ρsample] | ρH2O | ρair | m0 | m1 | m2
ID | yyyy-mm-dd | initials | model | °C | % | g | g | g | g | g | g | kg/m3 | kg/m3 | kg/m3 | kg/m3 | g | g | g | g | g | g | kg/m3 | kg/m3
MOOS085a + MOOS237a | 2024-01-23 | EP | MS204S | 21.8 | 18.9 | 23.2601 | -0.0001 | 48.2081 | -0.0001 | 48.795 | 0 | 997.8182 | 0.43189 | 1.2094 | 0.006 | 23.2602 | 0.0002 | 48.2082 | 0.0002 | 48.795 | 0.0002 | 1021.2614 | 0.44219 | 99.93% | 1.019e-05% | 1.806e-05% | 0.0342% | 0.03264%
MOOS199a + MOOS275a | 2024-01-23 | EP | MS204S | 21.8 | 18.9 | 23.2608 | 0.0002 | 48.2093 | -0.0002 | 48.7948 | 0 | 997.8182 | 0.43189 | 1.2094 | 0.006 | 23.2607 | 0.0002 | 48.2094 | 0.0002 | 48.7948 | 0.0002 | 1021.2028 | 0.44217 | 99.93% | 1.014e-05% | 1.797e-05% | 0.0342% | 0.03265%
MOOS009a + MOOS047a | 2024-01-23 | EP | MS204S | 21.8 | 18.9 | 23.2603 | 0 | 48.2079 | 0.0001 | 48.7928 | -0.0002 | 997.8182 | 0.43189 | 1.2094 | 0.006 | 23.2603 | 0.0002 | 48.2079 | 0.0002 | 48.7929 | 0.0002 | 1021.1899 | 0.44216 | 99.93% | 1.013e-05% | 1.796e-05% | 0.0342% | 0.03265%
LNSW | 2024-02-22 | EP | MS204S | 21.8 | 24.4 | 23.2605 | -0.0001 | 48.2083 | 0.0001 | 48.8705 | -0.0003 | 997.8182 | 0.43189 | 1.2088 | 0.006 | 23.2606 | 0.0002 | 48.2082 | 0.0002 | 48.8706 | 0.0002 | 1024.2797 | 0.4435 | 99.93% | 1.29e-05% | 2.288e-05% | 0.0342% | 0.03245%
LNSW | 2024-02-22 | EP | MS204S | 21.8 | 24.4 | 23.2613 | -0.0001 | 48.2092 | 0.0001 | 48.873 | 0 | 997.8182 | 0.43189 | 1.2088 | 0.006 | 23.2614 | 0.0002 | 48.2092 | 0.0002 | 48.873 | 0.0002 | 1024.3375 | 0.44353 | 99.93% | 1.296e-05% | 2.298e-05% | 0.0342% | 0.03245%
LNSW | 2024-02-22 | EP | MS204S | 21.8 | 24.4 | 23.2611 | -0.0001 | 48.2094 | 0 | 48.8751 | 0.0001 | 997.8182 | 0.43189 | 1.2088 | 0.006 | 23.2612 | 0.0002 | 48.2094 | 0.0002 | 48.8751 | 0.0002 | 1024.409 | 0.44356 | 99.93% | 1.303e-05% | 2.31e-05% | 0.0342% | 0.03244%
MOOS052a + MOOS130a + MOOS207a + MOOS284a + MOOS360a | 2024-03-06 | EP | MS204S | 22.7 | 31.1 | 23.2601 | 0 | 48.2053 | -0.0002 | 48.7861 | 0 | 997.6115 | 0.43351 | 1.2046 | 0.006 | 23.2601 | 0.0002 | 48.2054 | 0.0002 | 48.7861 | 0.0002 | 1020.8068 | 0.44375 | 99.93% | 9.907e-06% | 1.756e-05% | 0.03394% | 0.03241%
MOOS052a + MOOS130a + MOOS207a + MOOS284a + MOOS360a | 2024-03-06 | EP | MS204S | 22.7 | 31.1 | 23.2605 | -0.0001 | 48.204 | 0 | 48.785 | -0.0001 | 997.6115 | 0.43351 | 1.2046 | 0.006 | 23.2606 | 0.0002 | 48.204 | 0.0002 | 48.785 | 0.0002 | 1020.8225 | 0.44376 | 99.93% | 9.92e-06% | 1.759e-05% | 0.03394% | 0.03241%
MOOS052a + MOOS130a + MOOS207a + MOOS284a + MOOS360a | 2024-03-06 | EP | MS204S | 22.7 | 31.1 | 23.2612 | -0.0001 | 48.2036 | 0 | 48.7806 | 0 | 997.6115 | 0.43351 | 1.2046 | 0.006 | 23.2613 | 0.0002 | 48.2036 | 0.0002 | 48.7806 | 0.0002 | 1020.6617 | 0.44369 | 99.93% | 9.786e-06% | 1.735e-05% | 0.03394% | 0.03243%
LNSW | 2024-03-06 | EP | MS204S | 22.7 | 31.1 | 23.2614 | 0 | 48.202 | -0.0003 | 48.8616 | 0 | 997.6115 | 0.43351 | 1.2046 | 0.006 | 23.2614 | 0.0002 | 48.2021 | 0.0002 | 48.8616 | 0.0002 | 1023.9572 | 0.44512 | 99.93% | 1.27e-05% | 2.253e-05% | 0.03395% | 0.03222%
LNSW | 2024-03-06 | EP | MS204S | 22.7 | 31.1 | 23.2606 | 0 | 48.2012 | 0 | 48.8633 | 0 | 997.6115 | 0.43351 | 1.2046 | 0.006 | 23.2606 | 0.0002 | 48.2012 | 0.0002 | 48.8633 | 0.0002 | 1024.0632 | 0.44517 | 99.93% | 1.28e-05% | 2.27e-05% | 0.03395% | 0.03222%
Temperature and relative humidity measured with fisherbrand 15-077-963 11724196 S/N 230419093 (Expiry date 21 Jun 2025)
Sample | Density | STDEV | u | Unit | uR
MOOS-4 | 1.021 | 0.00025621 | 0.00050111 | g/mL | 0.04908% |  | Linear pooling 2024-05-09, https://consensus.nist.gov/app/nicob
LNSW | 1.0242 | 0.00019116 | 0.00044356 | g/mL | 0.04344%
MOOS-3 | 1.0307 |  |  | g/mL |  |  | Data obtained from the certification of SALT-1 CRM
SALT-1 | 1.0232 |  |  | g/mL |  |  | Data obtained from the certification of SALT-1 CRM
Sequences for the determination of nutrients in MOOS-4 week-1
Nitrate GC-MS |  |  | Nitrite UV-Vis |  |  | Phosphate UV-Vis |  |  | Nitrite GC-MS |  |  | Silicate UV-Vis
Week 1 Day 1 run x3 |  |  | Week 1 Day 1 |  |  | Week 1 Day 1 |  |  | Week 1 Day 2 run x3 |  |  | Week 1 Day 2
Reag BLK DIW |  |  | 1 cm cuvette |  |  | 5 cm cuvette |  |  | Reag BLK DIW |  |  | 5 cm cuvette
Reag BLK 3.5% NaCl |  |  | Zero baseline DIW |  |  | Zero baseline DIW |  |  | Reag BLK 3.5% NaCl |  |  | Zero baseline DIW
IntST BLK DIW |  |  | Baseline DIW |  |  | Baseline DIW |  |  | IntST BLK DIW |  |  | Baseline DIW
InST BLK 3.5% NaCl |  |  | Baseline LNSW |  |  | Baseline LNSW |  |  | InST BLK 3.5% NaCl |  |  | Baseline LNSW
A*B-1 |  |  | Baseline MOOS-4 |  |  | Baseline MOOS-4 |  |  | A*B-1 |  |  | Baseline MOOS-4
A*B-2A |  |  | Baseline SALT-1 |  |  | Baseline SALT-1 |  |  | A*B-2A |  |  | Baseline SALT-1
A*B-2B |  |  | Baseline MOOS-3 |  |  | Baseline MOOS-3 |  |  | A*B-2B |  |  | Baseline MOOS-3
A*B-3 |  |  | Reag BLK DIW |  |  | Reag BLK DIW |  |  | A*B-3 |  |  | Reag BLK DIW
MOOS001a_i |  |  | Reag BLK LNSW |  |  | Reag BLK LNSW |  |  | MOOS001a_i |  |  | Reag BLK LNSW
MOOS078a_i |  |  | CAL-1 |  |  | CAL-1 |  |  | MOOS078a_i |  |  | CAL-1
MOOS155a_i |  |  | CAL-2 |  |  | CAL-2 |  |  | MOOS155a_i |  |  | CAL-2
MOOS232a_i |  |  | CAL-3 |  |  | CAL-3 |  |  | MOOS232a_i |  |  | CAL-3
MOOS310a_i |  |  | CAL-4 |  |  | CAL-4 |  |  | MOOS310a_i |  |  | CAL-4
MOOS001a_ii |  |  | CAL-5 |  |  | CAL-5 |  |  | MOOS001a_ii |  |  | CAL-5
MOOS078a_ii |  |  | WASH CUVETTE |  |  | WASH CUVETTE |  |  | MOOS078a_ii |  |  | WASH CUVETTE
MOOS155a_ii |  |  | MOOS001a_i |  |  | MOOS001a_i |  |  | MOOS155a_ii |  |  | MOOS001a_i
MOOS232a_ii |  |  | MOOS078a_i |  |  | MOOS078a_i |  |  | MOOS232a_ii |  |  | MOOS078a_i
MOOS310a_ii |  |  | MOOS155a_i |  |  | MOOS155a_i |  |  | MOOS310a_ii |  |  | MOOS155a_i
SALT-1 |  |  | MOOS232a_i |  |  | MOOS232a_i |  |  | MOOS-3 |  |  | MOOS232a_i
MOOS-3 |  |  | MOOS310a_i |  |  | MOOS310a_i |  |  | A*B-2C |  |  | MOOS310a_i
A*B-2C |  |  | MOOS001a_ii |  |  | MOOS001a_ii |  |  | A*B-4 |  |  | MOOS001a_ii
|  |  | MOOS078a_ii |  |  | MOOS078a_ii | [5 empty] | MOOS078a_ii
|  |  | MOOS155a_ii |  |  | MOOS155a_ii | [5 empty] | MOOS155a_ii
|  |  | MOOS232a_ii |  |  | MOOS232a_ii | [5 empty] | MOOS232a_ii
|  |  | MOOS310a_ii |  |  | MOOS310a_ii | [5 empty] | MOOS310a_ii
|  |  | WASH CUVETTE |  |  | WASH CUVETTE | [5 empty] | WASH CUVETTE
|  |  | SALT-1 |  |  | SALT-1 | [5 empty] | SALT-1
|  |  | MOOS-3 |  |  | MOOS-3 | [5 empty] | MOOS-3
corr EP 2024-06-25 internal audit
Sequences for the determination of nutrients in MOOS-4 week-2
Nitrate GC-MS |  |  | Phosphate UV-Vis |  |  | Nitrite GC-MS |  |  | Silicate UV-Vis
Week 1 Day 1 run x3 |  |  | Week 1 Day 1 |  |  | Week 1 Day 2 run x3 |  |  | Week 1 Day 2
Non-derivatized sample |  |  | 5 cm cuvette |  |  | Non-derivatized sample |  |  | 5 cm cuvette
Reag BLK DIW |  |  | Zero baseline DIW |  |  | Reag BLK DIW |  |  | Zero baseline DIW
Reag BLK 3.5% NaCl |  |  | Baseline DIW |  |  | Reag BLK 3.5% NaCl |  |  | Baseline DIW
IntST BLK DIW |  |  | Baseline LNSW |  |  | IntST BLK DIW |  |  | Baseline LNSW
InST BLK 3.5% NaCl |  |  | Baseline MOOS-4 |  |  | InST BLK 3.5% NaCl |  |  | Baseline MOOS-4
A*B-1 |  |  | Baseline SALT-1 |  |  | A*B-1 |  |  | Baseline SALT-1
A*B-2A |  |  | Baseline MOOS-3 |  |  | A*B-2A |  |  | Baseline MOOS-3
A*B-2B |  |  | Reag BLK DIW |  |  | A*B-2B |  |  | Reag BLK DIW
A*B-3 |  |  | Reag BLK LNSW |  |  | A*B-3 |  |  | Reag BLK LNSW
MOOS052a_i |  |  | CAL-1 |  |  | MOOS052a_i |  |  | CAL-1
MOOS130a_i |  |  | CAL-2 |  |  | MOOS130a_i |  |  | CAL-2
MOOS207a_i |  |  | CAL-3 |  |  | MOOS207a_i |  |  | CAL-3
MOOS284a_i |  |  | CAL-4 |  |  | MOOS284a_i |  |  | CAL-4
MOOS360a_i |  |  | CAL-5 |  |  | MOOS360a_i |  |  | CAL-5
MOOS052a_ii |  |  | WASH CUVETTE |  |  | MOOS052a_ii |  |  | WASH CUVETTE
MOOS130a_ii |  |  | MOOS052a_i |  |  | MOOS130a_ii |  |  | MOOS052a_i
MOOS207a_ii |  |  | MOOS130a_i |  |  | MOOS207a_ii |  |  | MOOS130a_i
MOOS284a_ii |  |  | MOOS207a_i |  |  | MOOS284a_ii |  |  | MOOS207a_i
MOOS360a_ii |  |  | MOOS284a_i |  |  | MOOS360a_ii |  |  | MOOS284a_i
SALT-1 |  |  | MOOS360a_i |  |  | MOOS-3 |  |  | MOOS360a_i
MOOS-3 |  |  | MOOS052a_ii |  |  | A*B-2C |  |  | MOOS052a_ii
A*B-2C |  |  | MOOS130a_ii |  |  | A*B-4 |  |  | MOOS130a_ii
|  |  | MOOS207a_ii | [5 empty] | MOOS207a_ii
|  |  | MOOS284a_ii | [5 empty] | MOOS284a_ii
|  |  | MOOS360a_ii | [5 empty] | MOOS360a_ii
|  |  | WASH CUVETTE | [5 empty] | WASH CUVETTE
|  |  | SALT-1 | [5 empty] | SALT-1
|  |  | MOOS-3 | [5 empty] | MOOS-3
Planning the quantitative analysis: preparation of the calibration standards
Nitrate (GC-MS) |  |  |  | Nitrite (UV-vis) |  |  |  | Phosphate (UV-vis) |  |  |  | Nitrite (GC-MS) |  |  |  | Silicate (UV-vis)
0.5 | 0.53092 | µg/g |  | 0.5 | 0.037996 | µg/g |  | 0.5 | 0.13155 | µg/g |  | 0.5 | 0.038793 | µg/g |  | 0.5 | 0.21185 | µg/g
1 | 1.0618 | µg/g |  | 0.75 | 0.056994 | µg/g |  | 0.75 | 0.19733 | µg/g |  | 1 | 0.077587 | µg/g |  | 0.75 | 0.31778 | µg/g
2 | 2.1237 | µg/g |  | 1 | 0.075992 | µg/g |  | 1 | 0.2631 | µg/g |  | 2 | 0.15517 | µg/g |  | 1 | 0.4237 | µg/g
3 | 3.1855 | µg/g |  | 1.5 | 0.11399 | µg/g |  | 1.5 | 0.39465 | µg/g |  | 3 | 0.23276 | µg/g |  | 1.5 | 0.63555 | µg/g
|  |  |  | 2 | 0.18 | µg/g |  | 2 | 0.5262 | µg/g | [5 empty] | 2 | 0.8474 | µg/g
w(NO3) | 1.0618 | µg/g |  | w(NO2) | 0.075992 | µg/g |  | w(PO4) | 0.2631 | µg/g |  | w(NO2) | 0.077587 | µg/g |  | w(SiO2) | 0.4237 | µg/g
Primary standards and atomic weights (verified on 2024-04-10)
Nitrate
N=NIST | SRM3185 | Lot # | 170309
Certified value of nitrate
1.0006 | mg/g | 0.0018 | mg/g
Certified value as NO3 with k = 1
1000.6 | µg/g | 0.91603 | µg/g
Nitrite
S=SIGMA | 67276 | Lot # | BCCK5026
Certified value of nitrite IDMS
1004.0619 | mg/kg | 5.773 | mg/kg
Certified value as NO2 with k = 1
1004.0619 | µg/g | 2.8865 | µg/g
Phosphate
N=NIST | SRM3186 | Lot # | 170606
Certified value of phosphate:
1.0005 | mg/g | 0.0041 | mg/g
Certified value as PO4 with k = 1:
1000.5 | µg/g | 1.7951 | µg/g
Silicate
N=NIST | SRM3150 | Lot # | 211509
Certified value of silicon (Si):
10.018 | mg/g | 0.025 | mg/g
Certified value as SiO2 with k = 1:
21432.0637 | µg/g | 26.5429 | µg/g
Commission on Isotopic Abundances and Atomic Weights (ciaaw.org)
Access: 2024-02-16
Silicon
Ar(Si) = [28.084, 28.086] since 2009
28.085 | g/mol
Nitrogen
Ar(N) = [14.006 43, 14.007 28] since 2009
14.0069 | g/mol
Oxygen
Ar(O) = [15.999 03, 15.999 77] since 2009
15.9994 | g/mol
Phosphorus
Ar(P) = 30.973 761 998(5) since 2013
30.9738 | g/mol
Nitrate (NO3−)
62.0051 | g/mol
Nitrite (NO2−)
46.0057 | g/mol
Phosphate (PO43−)
94.9714 | g/mol
Silica (SiO2)
60.0838 | g/mol
Preparation of the SALT-1 QC sample (verified 2024-04-10)
|  | Full unit | Full unit | Sol:bottle | Sol:bottle | Sol:+SALT | Sol:+SALT | Sol:+H2O | Sol:+H2O | Empty unit | Empty unit |  |  |  |  | Full | Sol:bottle | Sol:+SALT | Sol:+H2O | Empty
Unit | Unit | m0 load | m0 unload | n0 load | n0 unload | n1 load | n1 unload | n2 load | n2 unload | m1 load | m1 unload | Date | Analyst | t | Rel Hum | m0 | n0 | n1 | n2 | m1 | Balance | u[m] | w
ID unique | number | g | g | g | g | g | g | g | g | g | g | yyyy-mm-dd | initials | °C | % | g | g | g | g | g | model | g | g/g
SALT0113 | 113 | 11.4885 | -0.0001 | 21.804 | 0 | 25.8768 | -0.0001 | 123.5979 | 0.0001 | 7.3994 | 0.0001 | 2024-02-19 | EP | 21.8 | 12.1 | 11.4886 | 21.804 | 25.8769 | 123.5978 | 7.3994 | MS204S | 0.0002 | 0.040011
SALT0309 | 309 | 11.4932 | 0 | 21.7899 | 0 | 25.8973 | 0 | 124.6211 | 0.0001 | 7.3763 | -0.0001 | 2024-03-04 | EP | 22.6 | 32 | 11.4932 | 21.7899 | 25.8973 | 124.621 | 7.3763 | MS204S | 0.0002 | 0.039943
SALT0197 | 197 | 11.4691 | -0.0001 | 22.0245 | 0.0001 | 26.0884 | -0.0001 | 123.6198 | 0 | 7.3891 | 0.0001 | 2024-05-07 | EP | 22.5 | 32.2 | 11.4691 | 22.0244 | 26.0885 | 123.6198 | 7.3891 | MS204S | 0.0002 | 0.040002
SALT-1 certified values (verified on 2024-04-10)
measurand | value | u (k = 1) | U (k = 2) | UR (%) | units | Reference
w(NO3) | 28.8159 | 0.28447 | 0.56895 | 1.974% | µg/g | At the time of bottling
w(PO4) | 3.7552 | 0.024486 | 0.048972 | 1.304% | µg/g | At the time of bottling
w(SiO2) | 13.0785 | 0.20903 | 0.41806 | 3.197% | µg/g | At the time of bottling
Calculation of nutrient concentration in the SALT-1 solution
Unit | Unit | SALT-1 mass | SALT-1 mass | w(NO3) | uw(NO3) | w(PO4) | uw(PO4) | w(SiO2) | uw(SiO2)
ID unique | number | 2018-05-07 | 2024-02-19 | µg/g | µg/g | µg/g | µg/g | µg/g | µg/g
SALT0113 | 113 | 4.0666 | 4.0892 | 1.1466 | 0.011319 | 0.14942 | 0.00097429 | 0.52039 | 0.0083172
SALT0309 | 309 | 4.0946 | 4.1169 | 1.1448 | 0.011301 | 0.14918 | 0.00097276 | 0.51957 | 0.0083042
Weight of the MOOS-4 units
|  |  |  | Mass loss | Mass loss | Mass loss
|  | weight /g | weight /g | g | g | %
Unit | Date | CHAR date | 2023-05-03 |  | sample
MOOS001a | 2024-02-20 | 65.5465 | 65.5573 | -0.0108 | 53.7463 | -0.02009%
MOOS078a | 2024-02-20 | 62.6827 | 62.6936 | -0.0109 | 50.8481 | -0.02144%
MOOS155a | 2024-02-20 | 65.5372 | 65.5485 | -0.0113 | 53.7512 | -0.02102%
MOOS232a | 2024-02-20 | 64.0366 | 64.0476 | -0.011 | 52.2179 | -0.02107%
MOOS310a | 2024-02-20 | 63.2937 | 63.305 | -0.0113 | 51.4625 | -0.02196%
MOOS052a | 2024-03-05 | 66.2553 | 66.2645 | -0.0092 | 54.4696 | -0.01689%
MOOS130a | 2024-03-05 | 65.5698 | 65.5789 | -0.0091 | 53.7432 | -0.01693%
MOOS207a | 2024-03-05 | 65.4042 | 65.4132 | -0.009 | 53.5378 | -0.01681%
MOOS284a | 2024-03-05 | 65.2307 | 65.2399 | -0.0092 | 53.3981 | -0.01723%
MOOS360a | 2024-03-05 | 66.3902 | 66.4001 | -0.0099 | 54.5577 | -0.01815%
MOOS104a | 2024-05-02 | 62.8265 | 62.8398 | -0.0133 | 51.0218 | -0.02607%
MOOS258a | 2024-05-03 | 64.5284 | 64.5417 | -0.0133 | 52.7379 | -0.02522%
Matrix effect for nitrite, phosphate, silicate by UV-vis (MOOS-4 matrix vs LNSW, verified 2024-05-03)
SC = corrected analytical signal that would have been observed if the sample matrix was matched with LNSW
SM = analytical signal measured in the sample matrix
SM = α * SC
α is the ratio between the slope of the calibration plot obtain from the sample (MOOS-4) and the slope of the calibration plot obtained from the standard (LNSW)
SC = SM / α
Analyte | α | u(α)
Nitrite | 1.0096 | 0.001395
Phosphate | 1 | 0
Silicate | 1 | 0
2024-02-20 and 2024-02-21 Characterization MOOS-4 [WEEK 1; DAY 1 and 2]: sample preparation (verified 2024-04-11)
Balance calibration
Weigh the unoped units
UV-vis performance verification
Preparation of standard solutions silicate in DIW
| Origin |  | m0 = vial |  | m1 = m0 + SiO2 standard |  | m2 = m1 + DIW | [7 empty] | Final
ID solution | w(SiO2) | u[w(SiO2)] | load | unload | load | unload | load | unload | m0 | m1 | m2 | u(m0) | u(m1) | u(m2) | w(SiO2) | u[w(SiO2)] | Target | Traciability | Prep. Date | Exp. Date | t | Rel Hum | Analyst | Balance
unique | µg/g | µg/g | g | g | g | g | g | g | g | g | g | g | g | g | µg/g | µg/g | % |  | yyyymmdd | yyyymmdd | °C | % | initials | model
Sw1-D | 21432.0637 | 0 | 11.7469 | 0 | 12.7981 | 0.0001 | 41.562 | 0.0001 | 11.7469 | 12.798 | 41.5619 | 0.0002 | 0.0002 | 0.0002 | 755.6021 | 0.20344 | N/A | NIST | 20240220 | 20240221 | 21.5 | 12.5 | EP | MS204S
In a 50 mL bottle
Preparation of standard solutions nitrite, phosphate and silicate in LNSW for UV-vis experiment
| Origin |  | Origin |  | Origin |  | m0 = vial |  | m1 = m0 + NO2 standard |  | m2 = m1 + PO4 standard |  | m3 = m2 + SiO2 standard |  | m4 = m3 + LNSW | [11 empty] | Final |  | Final |  | Final |  | Target | Target | Target
ID solution | w(NO2) | u[w(NO2)] | w(PO4) | u[w(PO4)] | w(SiO2) | u[w(SiO2)] | load | unload | load | unload | load | unload | load | unload | load | unload | m0 | m1 | m2 | m3 | m4 | u(m0) | u(m1) | u(m2) | u(m3) | u(m4) | w(NO2) | u[w(NO2)] | w(PO4) | u[w(PO4)] | w(SiO2) | u[w(SiO2)] | NO2 | PO4 | SiO2 | Traciability | Prep. Date | Exp. Date | t | Rel Hum | Analyst | Balance
unique | µg/g | µg/g | µg/g | µg/g | µg/g | µg/g | g | g | g | g | g | g | g | g | g | g | g | g | g | g | g | g | g | g | g | g | µg/g | µg/g | µg/g | µg/g | µg/g | µg/g | % | % | % |  | yyyymmdd | yyyymmdd | °C | % | initials | model
N2PSw1-ST | 1004.0619 | 0 | 1000.5 | 0 | 755.6021 | 0.20344 | 16.4315 | 0 | 16.9341 | 0 | 18.5648 | 0.0001 | 22.2452 | 0.0001 | 46.549 | 0 | 16.4315 | 16.9341 | 18.5648 | 22.2451 | 46.549 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 0.002 | 16.7558 | 0.0094955 | 54.17 | 0.010067 | 92.3356 | 0.026578 | N/A | N/A | N/A | NIST | 20240220 | 20240221 | 21.5 | 12.5 | EP | MS204S
In a 100 mL bottle
Preparation of standard solutions nitrite, phosphate and silicate in LNSW for UV-vis experiment
| Origin |  | Origin |  | Origin |  | m0 = vial |  | m1 = m0 + standard |  | m2 = m1 + LNSW | [7 empty] | Final |  | Final |  | Final |  | Target
ID solution | w(NO2) | u[w(NO2)] | w(PO4) | u[w(PO4)] | w(SiO2) | u[w(SiO2)] | load | unload | load | unload | load | unload | m0 | m1 | m2 | u(m0) | u(m1) | u(m2) | w(NO2) | u[w(NO2)] | w(PO4) | u[w(PO4)] | w(SiO2) | u[w(SiO2)] | NO2 | PO4 | SiO2 | Traciability | Prep. Date | Exp. Date | t | Rel Hum | Analyst | Balance
unique | µg/g | µg/g | µg/g | µg/g | µg/g | µg/g | g | g | g | g | g | g | g | g | g | g | g | g | µg/g | µg/g | µg/g | µg/g | µg/g | µg/g | % | % | % |  | yyyymmdd | yyyymmdd | °C | % | initials | model
N2PSw1-0 | 16.7558 | 0.0094955 | 54.17 | 0.010067 | 92.3356 | 0.026578 | 8.2118 | -0.0002 | 9.904 | 0 | 35.0969 | 0.0001 | 8.2119 | 9.904 | 35.0968 | 0.0002 | 0.0002 | 0.002 | 1.0546 | 0.00062806 | 3.4094 | 0.00088952 | 5.8115 | 0.0019826 | N/A | N/A | N/A | NIST | 20240220 | 20240221 | 21.5 | 12.5 | EP | MS204S
N2PSw1-1 | 1.0546 | 0.00062806 | 3.4094 | 0.00088952 | 5.8115 | 0.0019826 | 8.2263 | 0 | 9.2556 | 0 | 35.3679 | -0.0001 | 8.2263 | 9.2556 | 35.368 | 0.0002 | 0.0002 | 0.002 | 0.039993 | 2.6398e-05 | 0.12929 | 4.9919e-05 | 0.22039 | 9.7913e-05 | 105.3% | 98.29% | 104% | NIST | 20240220 | 20240221 | 21.5 | 12.5 | EP | MS204S
N2PSw1-2 | 1.0546 | 0.00062806 | 3.4094 | 0.00088952 | 5.8115 | 0.0019826 | 8.196 | 0 | 9.7403 | -0.0002 | 35.4846 | 0 | 8.196 | 9.7404 | 35.4846 | 0.0002 | 0.0002 | 0.002 | 0.059684 | 3.7447e-05 | 0.19295 | 6.3128e-05 | 0.3289 | 0.00012963 | 104.7% | 97.78% | 103.5% | NIST | 20240220 | 20240221 | 21.5 | 12.5 | EP | MS204S
N2PSw1-3 | 1.0546 | 0.00062806 | 3.4094 | 0.00088952 | 5.8115 | 0.0019826 | 8.1678 | -0.0001 | 10.2194 | -0.0001 | 35.6436 | -0.0001 | 8.1678 | 10.2195 | 35.6437 | 0.0002 | 0.0002 | 0.002 | 0.078745 | 4.848e-05 | 0.25458 | 7.7396e-05 | 0.43394 | 0.00016279 | 103.6% | 96.76% | 102.4% | NIST | 20240220 | 20240221 | 21.5 | 12.5 | EP | MS204S
N2PSw1-4 | 1.0546 | 0.00062806 | 3.4094 | 0.00088952 | 5.8115 | 0.0019826 | 8.2251 | 0 | 11.3054 | -0.0001 | 36.3157 | 0 | 8.2251 | 11.3055 | 36.3157 | 0.0002 | 0.0002 | 0.002 | 0.11564 | 7.0175e-05 | 0.37386 | 0.00010681 | 0.63727 | 0.00022971 | 101.5% | 94.73% | 100.3% | NIST | 20240220 | 20240221 | 21.5 | 12.5 | EP | MS204S
N2PSw1-5 | 1.0546 | 0.00062806 | 3.4094 | 0.00088952 | 5.8115 | 0.0019826 | 8.2536 | -0.0001 | 12.3651 | -0.0001 | 35.7585 | 0 | 8.2537 | 12.3651 | 35.7585 | 0.0002 | 0.0002 | 0.002 | 0.15764 | 9.5208e-05 | 0.50964 | 0.00014247 | 0.86871 | 0.00030892 | 87.58% | 96.85% | 102.5% | NIST | 20240220 | 20240221 | 21.5 | 12.5 | EP | MS204S
In 30 mL bottles
Preparation of standard solutions nitrite, nitrate in water for IDMS experiment
| Origin |  | m0 = vial |  | m1 = m0 + standard |  | m2 = m1 + DIW | [7 empty] | Final
ID solution | w(NO2) | u[w(NO2)] | load | unload | load | unload | load | unload | m0 | m1 | m2 | u(m0) | u(m1) | u(m2) | w(NO2) | u[w(NO2)] | Target | Traciability | Prep. Date | Exp. Date | t | Rel Hum | Analyst | Balance
unique | µg/g | µg/g | g | g | g | g | g | g | g | g | g | g | g | g | µg/g | µg/g | % |  | yyyymmdd | yyyymmdd | °C | % | initials | model
N2w1-D | 1004.0619 | 0 | 11.8027 | -0.0002 | 14.8086 | -0.0001 | 41.788 | 0.0001 | 11.8028 | 14.8087 | 41.7879 | 0.0002 | 0.0002 | 0.0002 | 100.6518 | 0.0095185 | N/A | N/A | 20240220 | 20240221 | 21.5 | 12.5 | EP | MS204S
In a 50 mL bottle
| Origin |  | Origin |  | m0 = vial |  | m1 = m0 + NO2 standard |  | m2 = m1 + NO3 standard |  | m3 = m2 + DIW | [9 empty] | Final |  | Final |  | Target
ID solution | w(NO2) | u[w(NO2)] | w(NO3) | u[w(NO3)] | load | unload | load | unload | load | unload | load | unload | m0 | m1 | m2 | m3 | u(m0) | u(m1) | u(m2) | u(m3) | w(NO2) | u[w(NO2)] | w(NO3) | u[w(NO3)] | NO2 | NO3 | Traciability | Prep. Date | Exp. Date | t | Rel Hum | Analyst | Balance
unique | µg/g | µg/g | µg/g | µg/g | g | g | g | g | g | g | g | g | g | g | g | g | g | g | g | g | µg/g | µg/g | µg/g | µg/g | % | % |  | yyyymmdd | yyyymmdd | °C | % | initials | model
N2N3w1-ST | 100.6518 | 0.0095185 | 1000.6 | 0 | 11.8131 | 0 | 12.5926 | 0 | 13.5966 | 0 | 41.8625 | 0.0001 | 11.8131 | 12.5926 | 13.5966 | 41.8624 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 2.611 | 0.00097935 | 33.4318 | 0.0094235 | N/A | N/A | NIST | 20240220 | 20240221 | 21.5 | 12.5 | EP | MS204S
In a 50 mL bottle
| Origin |  | Origin |  | m0 = vial |  | m1 = m0 + mass of standard |  | m2 = m1 + mass of LNSW | [7 empty] | Final |  | Final |  | Target
ID solution | w(NO2) | u[w(NO2)] | w(NO3) | u[w(NO3)] | load | unload | load | unload | load | unload | m0 | m1 | m2 | u(m0) | u(m1) | u(m2) | w(NO2) | u[w(NO2)] | w(NO3) | u[w(NO3)] | NO2 | NO3 | Traciability | Prep. Date | Exp. Date | t | Rel Hum | Analyst | Balance
unique | µg/g | µg/g | µg/g | µg/g | g | g | g | g | g | g | g | g | g | g | g | g | µg/g | µg/g | µg/g | µg/g | % | % |  | yyyymmdd | yyyymmdd | °C | % | initials | model
N2N3w1-1 | 2.611 | 0.00097935 | 33.4318 | 0.0094235 | 8.1698 | 0 | 8.5916 | 0.0001 | 35.7325 | 0.0001 | 8.1698 | 8.5915 | 35.7325 | 0.0002 | 0.0002 | 0.0002 | 0.039952 | 3.0702e-05 | 0.51156 | 0.00037218 | 103% | 96.35% | NIST | 20240220 | 20240221 | 21.5 | 12.5 | EP | MS204S
N2N3w1-2 | 2.611 | 0.00097935 | 33.4318 | 0.0094235 | 8.1972 | 0 | 9.0408 | 0 | 35.6275 | -0.0002 | 8.1972 | 9.0408 | 35.6276 | 0.0002 | 0.0002 | 0.0002 | 0.080298 | 4.0406e-05 | 1.0282 | 0.00045049 | 103.5% | 96.83% | NIST | 20240220 | 20240221 | 21.5 | 12.5 | EP | MS204S
N2N3w1-3 | 2.611 | 0.00097935 | 33.4318 | 0.0094235 | 8.2822 | -0.0002 | 9.9693 | -0.0001 | 36.1404 | 0.0001 | 8.2823 | 9.9694 | 36.1403 | 0.0002 | 0.0002 | 0.0002 | 0.15812 | 6.4983e-05 | 2.0246 | 0.00066431 | 101.9% | 95.33% | NIST | 20240220 | 20240221 | 21.5 | 12.5 | EP | MS204S
N2N3w1-4 | 2.611 | 0.00097935 | 33.4318 | 0.0094235 | 8.2226 | 0.0002 | 10.7586 | 0.0001 | 35.5673 | 0.0002 | 8.2225 | 10.7585 | 35.5672 | 0.0002 | 0.0002 | 0.0002 | 0.24215 | 9.4792e-05 | 3.1006 | 0.00094044 | 104% | 97.33% | NIST | 20240220 | 20240221 | 21.5 | 12.5 | EP | MS204S
In 30 mL bottles
Preparation of isotope dilution blends for NO2 and NO3 analysis by HS GC-MS
|  |  |  | m0 = mass of vial |  |  | m1 = m0 + mass of sample/ST |  |  | m2 = m1 + IntST
|  |  |  | load | unload | mass | load | unload | mass | load | unload | mass |  | mA | mB
Type | Description |  | ID | g | g | g | g | g | g | g | g | g | Date | g | g
BLK | Reag BLK DIW |  | BLK-DIW | [9 empty] | 2024-02-20
BLK | Reag BLK 3.5% NaCl |  | BLK-NaCl
BLK | IntST BLK DIW |  | BLK-IS-DIW | 6.152 | -0.0002 | 6.1521 | 11.1427 | -0.0001 | 11.1427 | 11.643 | 0 | 11.643 |  | 4.9906 | 0.50025
BLK | InST BLK 3.5% NaCl |  | BLK-IS-DIW | 6.1495 | 0.0003 | 6.1494 | 11.2466 | -0.0001 | 11.2467 | 11.7451 | -0.0002 | 11.7452 |  | 5.0973 | 0.49855
CAL | A*B-1 |  | AB1 | 6.1461 | 0 | 6.1461 | 11.1339 | 0.0001 | 11.1339 | 11.6308 | 0 | 11.6308 |  | 4.9878 | 0.49695
CAL | A*B-2A |  | AB2a | 6.1231 | -0.0001 | 6.1231 | 11.1036 | 0 | 11.1036 | 11.6009 | 0.0001 | 11.6008 |  | 4.9805 | 0.49725
CAL | A*B-2B |  | AB2b | 6.1215 | 0.0002 | 6.1214 | 11.1095 | 0 | 11.1095 | 11.6066 | 0 | 11.6066 |  | 4.9881 | 0.4971
CAL | A*B-3 |  | AB3 | 6.0868 | -0.0001 | 6.0869 | 11.0723 | -0.0001 | 11.0724 | 11.5704 | 0.0001 | 11.5703 |  | 4.9855 | 0.498
SAMPLE | MOOS001a_i |  | 001a_1 | 6.1317 | -0.0002 | 6.1318 | 11.2374 | 0.0001 | 11.2373 | 11.7353 | -0.0002 | 11.7354 |  | 5.1055 | 0.49805
SAMPLE | MOOS078a_i |  | 078a_1 | 6.1165 | 0.0001 | 6.1165 | 11.2091 | -0.0002 | 11.2092 | 11.7057 | -0.0001 | 11.7058 |  | 5.0927 | 0.49655
SAMPLE | MOOS155a_i |  | 155a_1 | 6.1327 | 0 | 6.1327 | 11.215 | -0.0001 | 11.215 | 11.7121 | -0.0001 | 11.7121 |  | 5.0823 | 0.4971
SAMPLE | MOOS232a_i |  | 232a_1 | 6.1237 | -0.0003 | 6.1239 | 11.212 | 0.0001 | 11.2119 | 11.7081 | -0.0001 | 11.7081 |  | 5.0881 | 0.4962
SAMPLE | MOOS310a_i |  | 310a_1 | 6.1464 | 0 | 6.1464 | 11.2257 | 0 | 11.2257 | 11.7225 | -0.0003 | 11.7226 |  | 5.0793 | 0.49695
SAMPLE | MOOS001a_ii |  | 001a_2 | 6.1299 | 0.0001 | 6.1299 | 11.2286 | 0 | 11.2286 | 11.7259 | 0.0001 | 11.7258 |  | 5.0987 | 0.49725
SAMPLE | MOOS078a_ii |  | 078a_2 | 6.1389 | -0.0001 | 6.1389 | 11.2304 | -0.0001 | 11.2304 | 11.7279 | 0 | 11.7279 |  | 5.0915 | 0.49745
SAMPLE | MOOS155a_ii |  | 155a_2 | 6.155 | -0.0001 | 6.1551 | 11.2334 | 0 | 11.2334 | 11.731 | -0.0002 | 11.7311 |  | 5.0783 | 0.4977
SAMPLE | MOOS232a_ii |  | 232a_2 | 6.1222 | -0.0001 | 6.1223 | 11.1992 | 0.0001 | 11.1991 | 11.6967 | 0.0001 | 11.6966 |  | 5.0769 | 0.4975
SAMPLE | MOOS310a_ii |  | 310a_2 | 6.1513 | 0.0001 | 6.1513 | 11.2281 | -0.0002 | 11.2282 | 11.7254 | 0.0002 | 11.7253 |  | 5.0769 | 0.4971
QC | SALT-1 |  | SALT1_1 | 6.1384 | 0 | 6.1384 | 11.2526 | -0.0001 | 11.2526 | 11.7507 | 0 | 11.7507 |  | 5.1142 | 0.49805
QC | MOOS-3 |  | MOOS3 | 6.1021 | 0 | 6.1021 | 11.2455 | 0.0001 | 11.2454 | 11.7426 | 0 | 11.7426 |  | 5.1433 | 0.49715
CAL | A*B-2C |  | AB2c | 6.1266 | -0.0001 | 6.1266 | 11.114 | 0 | 11.114 | 11.6104 | 0.0001 | 11.6104 |  | 4.9874 | 0.49635
CAL | A*B-4 |  | AB4 | 6.1146 | 0.0001 | 6.1146 | 11.0995 | 0 | 11.0995 | 11.5977 | -0.0003 | 11.5978 | 2024-02-20 | 4.985 | 0.49835
Day_1: determination of nitrate after nitrite removal by sulfamic acid
2 mL blend in 10 mL combiPAL vial + 0.1 mL sulfamic acid + 0.04 mL triethyloxonium tetrafluoroborate solution in water
Reag BLK = 1.82 mL of blank solution
Measure using SALT1 method
Preparation of external calibration for NO2 analysis by UV-vis
|  |  |  | m0 = mass of vial |  |  | m1 = m0 + mass of sample/ST |  |  | m2 = m1 + Reagents
|  |  |  | load | unload | mass | load | unload | mass | load | unload | mass |  | mS | mF
Type | Description |  | ID | g | g | g | g | g | g | g | g | g | Date | g | g
BLK | Baseline DIW |  | BSL-DIW | [9 empty] | 2024-02-20
BLK | Baseline LNSW |  | BSL-LNSW
BLK | Baseline MOOS-4 |  | BSL-M4
BLK | Baseline MOOS-3 |  | BSL-M3
BLK | Reag BLK DIW |  | BLK-DIW
BLK | Reag BLK LNSW |  | BLK-LNSW
STD | CAL-1 |  | CAL-1 | 6.119 | 0.0001 | 6.1189 | 11.233 | -0.0002 | 11.2331 | 11.4339 | 0 | 11.4339 |  | 5.1142 | 5.3149
STD | CAL-2 |  | CAL-2 | 6.1569 | 0 | 6.1569 | 11.2804 | -0.0002 | 11.2805 | 11.4817 | 0.0001 | 11.4817 |  | 5.1236 | 5.3247
STD | CAL-3 |  | CAL-3 | 6.11 | -0.0001 | 6.1101 | 11.2279 | 0 | 11.2279 | 11.4294 | 0 | 11.4294 |  | 5.1178 | 5.3193
STD | CAL-4 |  | CAL-4 | 6.1104 | 0 | 6.1104 | 11.2195 | -0.0001 | 11.2195 | 11.4202 | 0 | 11.4202 |  | 5.1091 | 5.3098
STD | CAL-5 |  | CAL-5 | 6.1312 | 0 | 6.1312 | 11.2752 | 0.0002 | 11.2751 | 11.4759 | 0.0002 | 11.4758 |  | 5.1439 | 5.3446
SAMPLE | MOOS001a_i |  | 001a_1 | 6.1006 | -0.0002 | 6.1007 | 11.2049 | -0.0001 | 11.205 | 11.406 | -0.0001 | 11.4061 |  | 5.1043 | 5.3054
SAMPLE | MOOS078a_i |  | 078a_1 | 6.1437 | 0 | 6.1437 | 11.2352 | -0.0001 | 11.2353 | 11.4363 | 0.0002 | 11.4362 |  | 5.0916 | 5.2925
SAMPLE | MOOS155a_i |  | 155a_1 | 6.1263 | 0.0001 | 6.1262 | 11.228 | 0.0002 | 11.2279 | 11.4286 | -0.0001 | 11.4286 |  | 5.1017 | 5.3024
SAMPLE | MOOS232a_i |  | 232a_1 | 6.115 | -0.0001 | 6.1151 | 11.2147 | 0 | 11.2147 | 11.4159 | 0.0001 | 11.4159 |  | 5.0997 | 5.3008
SAMPLE | MOOS310a_i |  | 310a_1 | 6.1026 | 0 | 6.1026 | 11.1918 | 0 | 11.1918 | 11.3923 | 0 | 11.3923 |  | 5.0892 | 5.2897
SAMPLE | MOOS001a_ii |  | 001a_2 | 6.1283 | 0.0001 | 6.1283 | 11.2381 | -0.0001 | 11.2381 | 11.4391 | 0 | 11.4391 |  | 5.1099 | 5.3108
SAMPLE | MOOS078a_ii |  | 078a_2 | 6.1556 | 0 | 6.1556 | 11.2493 | -0.0001 | 11.2493 | 11.4501 | 0 | 11.4501 |  | 5.0938 | 5.2945
SAMPLE | MOOS155a_ii |  | 155a_2 | 6.1339 | -0.0001 | 6.1339 | 11.2301 | 0 | 11.2301 | 11.4305 | 0 | 11.4305 |  | 5.0962 | 5.2966
SAMPLE | MOOS232a_ii |  | 232a_2 | 6.1146 | 0.0001 | 6.1146 | 11.1989 | -0.0002 | 11.199 | 11.3997 | 0 | 11.3997 |  | 5.0844 | 5.2851
SAMPLE | MOOS310a_ii |  | 310a_2 | 6.1488 | 0 | 6.1488 | 11.2456 | 0 | 11.2456 | 11.4455 | -0.0001 | 11.4455 |  | 5.0968 | 5.2967
QC | MOOS-3 |  | MOOS3 | 6.1156 | 0.0001 | 6.1155 | 11.2608 | 0 | 11.2608 | 11.4618 | 0 | 11.4618 | 2024-02-20 | 5.1452 | 5.3463
Use the 1 cm cuvette and zero the instrument against genpure water
Measure ABS at 541 nm with existing method
Preparation of external calibration for PO4 analysis by UV-vis
|  |  |  | m0 = mass of vial |  |  | m1 = m0 + mass of sample/ST |  |  | m2 = m1 + Reagents
|  |  |  | load | unload | mass | load | unload | mass | load | unload | mass |  | mS | mF
Type | Description |  | ID | g | g | g | g | g | g | g | g | g | Date | g | g
BLK | Baseline DIW |  | BSL-DIW | [9 empty] | 2024-02-20
BLK | Baseline LNSW |  | BSL-LNSW
BLK | Baseline MOOS-4 |  | BSL-M4
BLK | Baseline SALT-1 |  | BSL-SALT
BLK | Baseline MOOS-3 |  | BSL-M3
BLK | Reag BLK DIW |  | BLK-DIW
BLK | Reag BLK LNSW |  | BLK-LNSW
STD | CAL-1 |  | CAL-1 | 6.1279 | -0.0001 | 6.128 | 12.2898 | -0.0001 | 12.2898 | 13.0738 | 0 | 13.0738 |  | 6.1619 | 6.9459
STD | CAL-2 |  | CAL-2 | 6.1152 | 0.0001 | 6.1151 | 12.2578 | 0 | 12.2578 | 13.0423 | -0.0001 | 13.0423 |  | 6.1426 | 6.9272
STD | CAL-3 |  | CAL-3 | 6.1109 | -0.0001 | 6.1109 | 12.2486 | -0.0001 | 12.2486 | 13.0324 | -0.0001 | 13.0325 |  | 6.1377 | 6.9215
STD | CAL-4 |  | CAL-4 | 6.122 | 0.0001 | 6.122 | 12.2705 | 0 | 12.2705 | 13.0553 | 0.0004 | 13.0551 |  | 6.1486 | 6.9332
STD | CAL-5 |  | CAL-5 | 6.109 | 0.0001 | 6.109 | 12.2521 | 0 | 12.2521 | 13.0369 | -0.0004 | 13.0371 |  | 6.1432 | 6.9281
SAMPLE | MOOS001a_i |  | 001a_1 | 6.1437 | -0.0001 | 6.1437 | 12.2696 | 0 | 12.2696 | 13.0554 | 0.0001 | 13.0554 |  | 6.1259 | 6.9116
SAMPLE | MOOS078a_i |  | 078a_1 | 6.1503 | 0 | 6.1503 | 12.2797 | 0 | 12.2797 | 13.0654 | 0.0001 | 13.0654 |  | 6.1294 | 6.9151
SAMPLE | MOOS155a_i |  | 155a_1 | 6.1312 | 0 | 6.1312 | 12.2628 | -0.0003 | 12.263 | 13.0482 | -0.0001 | 13.0482 |  | 6.1318 | 6.917
SAMPLE | MOOS232a_i |  | 232a_1 | 6.1332 | 0.0001 | 6.1332 | 12.2776 | 0.0001 | 12.2775 | 13.0586 | 0 | 13.0586 |  | 6.1444 | 6.9254
SAMPLE | MOOS310a_i |  | 310a_1 | 6.1133 | -0.0002 | 6.1134 | 12.2391 | -0.0002 | 12.2392 | 13.0224 | 0.0001 | 13.0223 |  | 6.1258 | 6.9089
SAMPLE | MOOS001a_ii |  | 001a_2 | 6.1272 | 0 | 6.1272 | 12.2514 | 0.0002 | 12.2513 | 13.0353 | 0.0001 | 13.0352 |  | 6.1241 | 6.908
SAMPLE | MOOS078a_ii |  | 078a_2 | 6.1452 | -0.0001 | 6.1452 | 12.2869 | 0 | 12.2869 | 13.0711 | 0 | 13.0711 |  | 6.1416 | 6.9258
SAMPLE | MOOS155a_ii |  | 155a_2 | 6.1125 | -0.0001 | 6.1125 | 12.2303 | -0.0001 | 12.2303 | 13.0138 | 0.0001 | 13.0137 |  | 6.1178 | 6.9012
SAMPLE | MOOS232a_ii |  | 232a_2 | 6.1496 | 0 | 6.1496 | 12.2887 | 0 | 12.2887 | 13.0728 | 0.0002 | 13.0727 |  | 6.1391 | 6.9231
SAMPLE | MOOS310a_ii |  | 310a_2 | 6.1204 | 0 | 6.1204 | 12.2386 | -0.0002 | 12.2387 | 13.0229 | 0 | 13.0229 |  | 6.1183 | 6.9025
QC | SALT-1 |  | SALT1_1 | 6.1394 | 0.0001 | 6.1394 | 12.2784 | -0.0002 | 12.2785 | 13.0634 | 0 | 13.0634 |  | 6.1391 | 6.924
QC | MOOS-3 |  | MOOS3 | 6.1525 | 0 | 6.1525 | 12.3063 | 0.0002 | 12.3062 | 13.0888 | -0.0001 | 13.0889 | 2024-02-20 | 6.1537 | 6.9364
Use the 5 cm cuvette and zero the instrument against genpure water
Measure ABS at 890 nm with existing method
=== Day break carry on next day ===
Day_2: determination of nitrite
2 mL blend prepared previous day in 10 mL combiPAL vial + 0.04 mL triethyloxonium tetrafluoroborate solution in water
Reag BLK = 1.82 mL of blank solution
Measure using MOOS4 method
Preparation of standard solutions silicate in DIW
| Origin |  | m0 = vial |  | m1 = m0 + standard |  | m2 = m1 + DIW | [7 empty] | Final
ID solution | w(SiO2) | u[w(SiO2)] | load | unload | load | unload | load | unload | m0 | m1 | m2 | u(m0) | u(m1) | u(m2) | w(SiO2) | u[w(SiO2)] | Target | Traciability | Prep. Date | Exp. Date | t | Rel Hum | Analyst | Balance
unique | µg/g | µg/g | g | g | g | g | g | g | g | g | g | g | g | g | µg/g | µg/g | % |  | yyyymmdd | yyyymmdd | °C | % | initials | model
Sw1d2-D | 21432.0637 | 0 | 8.164 | 0 | 8.8962 | -0.0003 | 35.0048 | 0 | 8.164 | 8.8963 | 35.0048 | 0.0002 | 0.0002 | 0.0002 | 584.7729 | 0.22593 | N/A | NIST | 20240221 | 20240222 | 21.5 | 13.9 | EP | MS204S
In a 30 mL bottle
Preparation of standard solutions silicate in LNSW
| Origin |  | m0 = vial |  | m1 = m0 + standard |  | m2 = m1 + LNSW | [7 empty] | Final
ID solution | w(SiO2) | u[w(SiO2)] | load | unload | load | unload | load | unload | m0 | m1 | m2 | u(m0) | u(m1) | u(m2) | w(SiO2) | u[w(SiO2)] | Target | Traciability | Prep. Date | Exp. Date | t | Rel Hum | Analyst | Balance
unique | µg/g | µg/g | g | g | g | g | g | g | g | g | g | g | g | g | µg/g | µg/g | % |  | yyyymmdd | yyyymmdd | °C | % | initials | model
Sw1d2-ST | 584.7729 | 0.22593 | 8.2661 | 0 | 8.8163 | -0.0001 | 35.4263 | 0 | 8.2661 | 8.8163 | 35.4263 | 0.0002 | 0.0002 | 0.002 | 11.8472 | 0.0076684 | N/A | NIST | 20240221 | 20240222 | 21.5 | 13.9 | EP | MS204S
In a 30 mL bottle
| Origin |  | m0 = vial |  | m1 = m0 + standard |  | m2 = m1 + LNSW | [7 empty] | Final
ID solution | w(SiO2) | u[w(SiO2)] | load | unload | load | unload | load | unload | m0 | m1 | m2 | u(m0) | u(m1) | u(m2) | w(SiO2) | u[w(SiO2)] | Target | Traciability | Prep. Date | Exp. Date | t | Rel Hum | Analyst | Balance
unique | µg/g | µg/g | g | g | g | g | g | g | g | g | g | g | g | g | µg/g | µg/g | % |  | yyyymmdd | yyyymmdd | °C | % | initials | model
Sw1d2-1 | 11.8472 | 0.0076684 | 8.2068 | -0.0001 | 8.7209 | -0.0001 | 36.7566 | 0.0001 | 8.2068 | 8.721 | 36.7565 | 0.0002 | 0.0002 | 0.002 | 0.21333 | 0.00018185 | 100.7% | NIST | 20240221 | 20240222 | 21.5 | 13.9 | EP | MS204S
Sw1d2-2 | 11.8472 | 0.0076684 | 8.1531 | 0 | 8.9246 | 0 | 37.2917 | 0.0001 | 8.1531 | 8.9246 | 37.2916 | 0.0002 | 0.0002 | 0.002 | 0.31368 | 0.00023434 | 98.71% | NIST | 20240221 | 20240222 | 21.5 | 13.9 | EP | MS204S
Sw1d2-3 | 11.8472 | 0.0076684 | 8.2684 | 0 | 9.2921 | 0 | 37.6176 | 0.0003 | 8.2684 | 9.2921 | 37.6175 | 0.0002 | 0.0002 | 0.002 | 0.41323 | 0.0002922 | 97.53% | NIST | 20240221 | 20240222 | 21.5 | 13.9 | EP | MS204S
Sw1d2-4 | 11.8472 | 0.0076684 | 8.2666 | 0 | 9.8104 | -0.0001 | 37.5717 | 0.0002 | 8.2666 | 9.8104 | 37.5716 | 0.0002 | 0.0002 | 0.002 | 0.62413 | 0.00042204 | 98.2% | NIST | 20240221 | 20240222 | 21.5 | 13.9 | EP | MS204S
Sw1d2-5 | 11.8472 | 0.0076684 | 8.2036 | 0.0001 | 10.255 | -0.0002 | 37.7905 | 0 | 8.2035 | 10.2551 | 37.7905 | 0.0002 | 0.0002 | 0.002 | 0.82148 | 0.00054651 | 96.94% | NIST | 20240221 | 20240222 | 21.5 | 13.9 | EP | MS204S
Sw1d2-6 | 11.8472 | 0.0076684 | 8.231 | 0 | 14.3393 | -0.0001 | 39.6892 | 0.0003 | 8.231 | 14.3393 | 39.6891 | 0.0002 | 0.0002 | 0.002 | 2.3004 | 0.0015 | 121.1% | NIST | 20240221 | 20240222 | 21.5 | 13.9 | EP | MS204S
Preparation of external calibration for SiO2 analysis by UV-vis
|  |  |  | m0 = mass of vial |  |  | m1 = m0 + mass of sample/ST |  |  | m2 = m1 + Reagents
|  |  |  | load | unload | mass | load | unload | mass | load | unload | mass |  | mS | mF
Type | Description |  | ID | g | g | g | g | g | g | g | g | g | Date | g | g
BLK | Baseline DIW |  | BSL-DIW | [9 empty] | 2024-02-21
BLK | Baseline LNSW |  | BSL-LNSW
BLK | Baseline MOOS-4 |  | BSL-M4
BLK | Baseline SALT-1 |  | BSL-SALT
BLK | Baseline MOOS-3 |  | BSL-M3
BLK | Reag BLK DIW |  | BLK-DIW
BLK | Reag BLK LNSW |  | BLK-LNSW
STD | CAL-1 |  | CAL-1 | 6.1488 | -0.0001 | 6.1488 | 9.213 | 0 | 9.213 | 9.5355 | 0.0001 | 9.5355 |  | 3.0641 | 3.3866
STD | CAL-2 |  | CAL-2 | 6.1515 | -0.0001 | 6.1516 | 9.2227 | -0.0001 | 9.2227 | 9.5437 | 0 | 9.5437 |  | 3.0712 | 3.3921
STD | CAL-3 |  | CAL-3 | 6.1329 | -0.0002 | 6.133 | 9.199 | 0 | 9.199 | 9.5209 | -0.0002 | 9.521 |  | 3.066 | 3.388
STD | CAL-4 |  | CAL-4 | 6.1517 | 0.0002 | 6.1516 | 9.2228 | 0 | 9.2228 | 9.5444 | 0.0001 | 9.5443 |  | 3.0712 | 3.3927
STD | CAL-5 |  | CAL-5 | 6.0977 | 0 | 6.0977 | 9.1664 | 0.0001 | 9.1663 | 9.4884 | 0.0001 | 9.4884 |  | 3.0686 | 3.3907
SAMPLE | MOOS001a_i |  | 001a_1 | 6.131 | 0.0001 | 6.131 | 9.1761 | 0.0001 | 9.1761 | 9.4979 | 0.0001 | 9.4978 |  | 3.0451 | 3.3669
SAMPLE | MOOS078a_i |  | 078a_1 | 6.1186 | -0.0001 | 6.1186 | 9.1574 | 0.0002 | 9.1573 | 9.4794 | 0 | 9.4794 |  | 3.0387 | 3.3608
SAMPLE | MOOS155a_i |  | 155a_1 | 6.1476 | 0 | 6.1476 | 9.1806 | -0.0002 | 9.1807 | 9.502 | -0.0002 | 9.5021 |  | 3.0331 | 3.3545
SAMPLE | MOOS232a_i |  | 232a_1 | 6.1326 | 0.0001 | 6.1326 | 9.167 | 0 | 9.167 | 9.4885 | 0 | 9.4885 |  | 3.0344 | 3.3559
SAMPLE | MOOS310a_i |  | 310a_1 | 6.1266 | 0 | 6.1266 | 9.1714 | 0.0001 | 9.1714 | 9.4919 | 0 | 9.4919 |  | 3.0448 | 3.3653
SAMPLE | MOOS001a_ii |  | 001a_2 | 6.135 | -0.0001 | 6.135 | 9.1635 | 0 | 9.1635 | 9.4852 | -0.0001 | 9.4853 |  | 3.0285 | 3.3502
SAMPLE | MOOS078a_ii |  | 078a_2 | 6.128 | -0.0001 | 6.128 | 9.1566 | 0 | 9.1566 | 9.4778 | 0.0001 | 9.4778 |  | 3.0285 | 3.3497
SAMPLE | MOOS155a_ii |  | 155a_2 | 6.1151 | 0 | 6.1151 | 9.1455 | -0.0002 | 9.1456 | 9.4671 | 0 | 9.4671 |  | 3.0305 | 3.352
SAMPLE | MOOS232a_ii |  | 232a_2 | 6.1332 | 0 | 6.1332 | 9.1729 | -0.0001 | 9.173 | 9.4946 | 0 | 9.4946 |  | 3.0397 | 3.3614
SAMPLE | MOOS310a_ii |  | 310a_2 | 6.1341 | 0.0001 | 6.1341 | 9.1744 | -0.0002 | 9.1745 | 9.4959 | -0.0001 | 9.496 |  | 3.0404 | 3.3619
QC | SALT-1 |  | SALT1_1 | 6.1252 | -0.0001 | 6.1253 | 9.1619 | 0.0001 | 9.1618 | 9.4825 | 0.0001 | 9.4825 |  | 3.0366 | 3.3572
QC | MOOS-3 |  | MOOS3 | 6.1052 | 0 | 6.1052 | 9.1673 | 0.0002 | 9.1672 | 9.4893 | 0 | 9.4893 |  | 3.062 | 3.3841
STD | CAL-6 |  | CAL-6 | 6.1456 | 0 | 6.1456 | 9.207 | -0.0001 | 9.2071 | 9.5291 | 0 | 9.5291 | 2024-02-21 | 3.0615 | 3.3835
Use the 1 cm cuvette and zero the instrument against genpure water
Measure ABS at 810 nm with existing method
Repeat nitrite UV
Preparation of standard solutions nitrite in DIW
| Origin |  | m0 = vial |  | m1 = m0 + standard |  | m2 = m1 + DIW | [7 empty] | Final
ID solution | w(NO2) | u[w(NO2)] | load | unload | load | unload | load | unload | m0 | m1 | m2 | u(m0) | u(m1) | u(m2) | w(NO2) | u[w(NO2)] | Target | Traciability | Prep. Date | Exp. Date | t | Rel Hum | Analyst | Balance
unique | µg/g | µg/g | g | g | g | g | g | g | g | g | g | g | g | g | µg/g | µg/g | % |  | yyyymmdd | yyyymmdd | °C | % | initials | model
N2w1d2-D | 1004.0619 | 0 | 8.1894 | 0.0001 | 9.1935 | 0.0002 | 35.1691 | 0 | 8.1893 | 9.1934 | 35.1691 | 0.0002 | 0.0002 | 0.0002 | 37.3661 | 0.010533 | N/A | NIST | 20240221 | 20240222 | 21.5 | 13.9 | EP | MS204S
In a 30 mL bottle
Preparation of standard solutions nitrite in LNSW
| Origin |  | m0 = vial |  | m1 = m0 + standard |  | m2 = m1 + LNSW | [7 empty] | Final
ID solution | w(NO2) | u[w(NO2)] | load | unload | load | unload | load | unload | m0 | m1 | m2 | u(m0) | u(m1) | u(m2) | w(NO2) | u[w(NO2)] | Target | Traciability | Prep. Date | Exp. Date | t | Rel Hum | Analyst | Balance
unique | µg/g | µg/g | g | g | g | g | g | g | g | g | g | g | g | g | µg/g | µg/g | % |  | yyyymmdd | yyyymmdd | °C | % | initials | model
N2w1d2-ST | 37.3661 | 0.010533 | 8.1235 | 0.0001 | 9.667 | -0.0001 | 35.1403 | 0 | 8.1235 | 9.667 | 35.1403 | 0.0002 | 0.0002 | 0.002 | 2.1349 | 0.00073515 | N/A | NIST | 20240221 | 20240222 | 21.5 | 13.9 | EP | MS204S
In a 30 mL bottle
| Origin |  | m0 = vial |  | m1 = m0 + standard |  | m2 = m1 + LNSW | [7 empty] | Final
ID solution | w(NO2) | u[w(NO2)] | load | unload | load | unload | load | unload | m0 | m1 | m2 | u(m0) | u(m1) | u(m2) | w(NO2) | u[w(NO2)] | Target | Traciability | Prep. Date | Exp. Date | t | Rel Hum | Analyst | Balance
unique | µg/g | µg/g | g | g | g | g | g | g | g | g | g | g | g | g | µg/g | µg/g | % |  | yyyymmdd | yyyymmdd | °C | % | initials | model
N2w1d2-1 | 2.1349 | 0.00073515 | 8.1888 | -0.0001 | 8.7018 | -0.0003 | 35.7922 | -0.0002 | 8.1889 | 8.702 | 35.7923 | 0.0002 | 0.0002 | 0.002 | 0.039684 | 2.5954e-05 | 104.4% | NIST | 20240221 | 20240222 | 21.5 | 13.9 | EP | MS204S
N2w1d2-2 | 2.1349 | 0.00073515 | 8.2674 | -0.0001 | 9.0373 | -0.0001 | 36.1531 | 0 | 8.2675 | 9.0373 | 36.1531 | 0.0002 | 0.0002 | 0.002 | 0.058943 | 2.9982e-05 | 103.4% | NIST | 20240221 | 20240222 | 21.5 | 13.9 | EP | MS204S
N2w1d2-3 | 2.1349 | 0.00073515 | 8.1986 | -0.0002 | 9.2237 | 0 | 36.0701 | 0 | 8.1987 | 9.2237 | 36.0701 | 0.0002 | 0.0002 | 0.002 | 0.078513 | 3.5105e-05 | 103.3% | NIST | 20240221 | 20240222 | 21.5 | 13.9 | EP | MS204S
N2w1d2-4 | 2.1349 | 0.00073515 | 8.2099 | 0 | 9.7536 | 0 | 36.0253 | 0 | 8.2099 | 9.7536 | 36.0253 | 0.0002 | 0.0002 | 0.002 | 0.11848 | 4.7002e-05 | 103.9% | NIST | 20240221 | 20240222 | 21.5 | 13.9 | EP | MS204S
N2w1d2-5 | 2.1349 | 0.00073515 | 8.2542 | -0.0002 | 10.6173 | 0 | 35.6878 | -0.0001 | 8.2543 | 10.6173 | 35.6879 | 0.0002 | 0.0002 | 0.002 | 0.18389 | 6.8379e-05 | 102.2% | NIST | 20240221 | 20240222 | 21.5 | 13.9 | EP | MS204S
Preparation of external calibration for NO2 analysis by UV-vis
|  |  |  | m0 = mass of vial |  |  | m1 = m0 + mass of sample/ST |  |  | m2 = m1 + Reagents
|  |  |  | load | unload | mass | load | unload | mass | load | unload | mass |  | mS | mF
Type | Description |  | ID | g | g | g | g | g | g | g | g | g | Date | g | g
BLK | Baseline DIW |  | BSL-DIW | [9 empty] | 2024-02-21
BLK | Baseline LNSW |  | BSL-LNSW
BLK | Baseline MOOS-4 |  | BSL-M4
BLK | Baseline MOOS-3 |  | BSL-M3
BLK | Reag BLK DIW |  | BLK-DIW
BLK | Reag BLK LNSW |  | BLK-LNSW
STD | CAL-1 |  | CAL-1 | 7.337 | 0 | 7.337 | 12.452 | -0.0001 | 12.452 | 12.6531 | 0 | 12.6531 |  | 5.1151 | 5.3161
STD | CAL-2 |  | CAL-2 | 7.3019 | 0.0002 | 7.3018 | 12.4175 | -0.0003 | 12.4177 | 12.6183 | 0.0001 | 12.6182 |  | 5.1159 | 5.3164
STD | CAL-3 |  | CAL-3 | 7.3347 | 0 | 7.3347 | 12.4648 | 0.0003 | 12.4647 | 12.6659 | -0.0001 | 12.666 |  | 5.13 | 5.3313
STD | CAL-4 |  | CAL-4 | 7.3373 | -0.0001 | 7.3373 | 12.4712 | 0 | 12.4712 | 12.6728 | 0 | 12.6728 |  | 5.1338 | 5.3355
STD | CAL-5 |  | CAL-5 | 7.3178 | 0.0001 | 7.3178 | 12.456 | 0 | 12.456 | 12.657 | 0.0001 | 12.657 |  | 5.1382 | 5.3392
SAMPLE | MOOS001a_i |  | 001a_1 | 7.3215 | 0.0001 | 7.3215 | 12.4002 | -0.0001 | 12.4002 | 12.6019 | -0.0001 | 12.602 |  | 5.0788 | 5.2805
SAMPLE | MOOS078a_i |  | 078a_1 | 7.3086 | -0.0002 | 7.3087 | 12.395 | 0.0001 | 12.3949 | 12.5968 | -0.0001 | 12.5968 |  | 5.0862 | 5.2881
SAMPLE | MOOS155a_i |  | 155a_1 | 7.3305 | 0.0001 | 7.3304 | 12.4243 | 0 | 12.4243 | 12.6257 | 0.0001 | 12.6257 |  | 5.0939 | 5.2952
SAMPLE | MOOS232a_i |  | 232a_1 | 7.3357 | 0 | 7.3357 | 12.4315 | -0.0002 | 12.4316 | 12.6326 | -0.0002 | 12.6327 |  | 5.0959 | 5.297
SAMPLE | MOOS310a_i |  | 310a_1 | 7.3327 | -0.0001 | 7.3327 | 12.4239 | 0.0002 | 12.4238 | 12.6252 | 0.0002 | 12.6251 |  | 5.0911 | 5.2923
QC | MOOS-3 |  | MOOS3 | 7.3119 | 0.0001 | 7.3118 | 12.2307 | 0 | 12.2307 | 12.4318 | 0 | 12.4318 | 2024-02-21 | 4.9189 | 5.12
2024-02-20 and 2024-02-21 Characterization MOOS-4 [WEEK 1; DAY 1 and 2]: data analysis (verified 2024-04-19)
1. Determination of nitrate by IDMS w1d1
Sample ID | rA*B | u(rA*B) | Date | wA* | u(wA*) | mA* | mB | wA*(mA*/mB) | u(wA*(mA*/mB))
AB1 | 0.49291 | 0.010385 | 2024-02-20 13:48 | 0.51156 | 0.00037218 | 4.9878 | 0.49695 | 5.1343 | 0.0047516
AB2A | 0.98737 | 0.020803 | 2024-02-20 13:58 | 1.0282 | 0.00045049 | 4.9805 | 0.49725 | 10.2981 | 0.0074171
AB2B | 0.9571 | 0.020165 | 2024-02-20 14:08 | 1.0282 | 0.00045049 | 4.9881 | 0.4971 | 10.317 | 0.0074321
AB3 | 1.8695 | 0.039388 | 2024-02-20 14:18 | 2.0246 | 0.00066431 | 4.9855 | 0.498 | 20.2682 | 0.013344
AB2C | 1.0012 | 0.021095 | 2024-02-20 16:28 | 1.0282 | 0.00045049 | 4.9874 | 0.49635 | 10.3311 | 0.0074492
[8 empty] | Gauss's method (GUM's Linear Approximation) |  |  |  | Monte Carlo method |  |  | Ratio wA | Ratio u(wA) | Z-score | Degree of
Sample ID | rAB | u(rAB) | Date | mA | mB | m = mB/mA | u(m) | wA | u(wA) | uR(wA) | R^2 | wA | u(wA) | uR(wA) | GS/MT | GS/MT | value | matching
001_A1 | 1.0578 | 0.0095499 | 2024-02-20 14:28 | 5.1055 | 0.49805 | 0.097551 | 5.5662e-05 | 1.0875 | 0.017869 | 1.643% | 0.99893 | 1.0883 | 0.017163 | 1.577% | 99.92% | 104.1% | 2.5199 | 107.7%
078_A1 | 1.0146 | 0.0091909 | 2024-02-20 14:38 | 5.0927 | 0.49655 | 0.097501 | 5.5802e-05 | 1.0412 | 0.017134 | 1.646% | 0.99893 | 1.042 | 0.015816 | 1.518% | 99.92% | 108.3% | -0.20903 | 103.3%
155_A1 | 1.0015 | 0.0090629 | 2024-02-20 14:48 | 5.0823 | 0.4971 | 0.097809 | 5.5918e-05 | 1.0305 | 0.016954 | 1.645% | 0.99893 | 1.0316 | 0.016108 | 1.561% | 99.89% | 105.3% | -0.83661 | 102%
232_A1 | 1.0016 | 0.0090679 | 2024-02-20 14:58 | 5.0881 | 0.4962 | 0.097522 | 5.5853e-05 | 1.0276 | 0.016909 | 1.645% | 0.99893 | 1.0291 | 0.016293 | 1.583% | 99.85% | 103.8% | -1.0078 | 102%
310_A1 | 1.0003 | 0.0090689 | 2024-02-20 15:08 | 5.0793 | 0.49695 | 0.097838 | 5.5951e-05 | 1.0296 | 0.016949 | 1.646% | 0.99893 | 1.0315 | 0.015548 | 1.507% | 99.81% | 109% | -0.89124 | 101.9%
001_A2 | 1.0054 | 0.0091219 | 2024-02-20 15:18 | 5.0987 | 0.49725 | 0.097524 | 5.5736e-05 | 1.0316 | 0.016987 | 1.647% | 0.99893 | 1.0321 | 0.016159 | 1.566% | 99.95% | 105.1% | -0.76959 | 102.4%
078_A2 | 0.99937 | 0.0090416 | 2024-02-20 15:28 | 5.0915 | 0.49745 | 0.097702 | 5.5816e-05 | 1.0271 | 0.016897 | 1.645% | 0.99893 | 1.0276 | 0.016076 | 1.564% | 99.95% | 105.1% | -1.0349 | 101.8%
155_A2 | 1.0058 | 0.0091044 | 2024-02-20 15:38 | 5.0783 | 0.4977 | 0.098004 | 5.5963e-05 | 1.0371 | 0.017064 | 1.645% | 0.99893 | 1.038 | 0.015757 | 1.518% | 99.92% | 108.3% | -0.44601 | 102.4%
232_A2 | 1.0005 | 0.0090531 | 2024-02-20 15:48 | 5.0769 | 0.4975 | 0.097993 | 5.5979e-05 | 1.0314 | 0.016968 | 1.645% | 0.99893 | 1.0321 | 0.016341 | 1.583% | 99.93% | 103.8% | -0.78358 | 101.9%
310_A2 | 1.0689 | 0.0096427 | 2024-02-20 15:58 | 5.0769 | 0.4971 | 0.097913 | 5.5978e-05 | 1.1034 | 0.018121 | 1.642% | 0.99893 | 1.1041 | 0.017189 | 1.557% | 99.94% | 105.4% | 3.4589 | 108.9%
SALT1_113 | 1.0973 | 0.009934 | 2024-02-20 16:08 | 5.1142 | 0.49805 | 0.097385 | 5.5566e-05 | 1.1276 | 0.018514 | 1.642% | 0.99893 | 1.1294 | 0.017831 | 1.579% | 99.84% | 103.8%
MOOS3 | 1.3124 | 0.011853 | 2024-02-20 16:18 | 5.1433 | 0.49715 | 0.096659 | 5.5248e-05 | 1.3482 | 0.021133 | 1.568% | 0.99893 | 1.3505 | 0.021467 | 1.59% | 99.82% | 98.45%
Value with Z-score > 2.5: random contamination of nitrate was the likely cause for these higher results which were excluded for technical reasons
| CRM |  | Measurement |  |  | Comparability |  |  |  |  | Ratio
Sample | μ | unc | x | s | n | IRMM | ISO | Wald | t-test | Mood, B=2 | x/μ
SALT1_113 | 1.1466 | 0.011319 | 1.1276 | 0.018514 | 3 | TRUE | TRUE | TRUE | TRUE | TRUE | 0.98343
MOOS3 | 1.384 | 0.0075 | 1.3482 | 0.021133 | 3 | FALSE | TRUE | TRUE | TRUE | FALSE | 0.9741
Note ICM-21-02 for MOOS-3 stability: nitrate (n = 21) w = 1.3921 μg/g, uw = 0.020 μg/g
n = 3 as the isotope ratio was measured in triplicate
Comparability tests are for information only
MOOS-3 is expired and not properly informative as a QC sample at this point
2. Determination of nitrite by IDMS w1d2
Sample ID | rA*B | u(rA*B) | Date | wA* | u(wA*) | mA* | mB | wA*(mA*/mB) | u(wA*(mA*/mB))
AB1 | 0.491 | 0.001376 | 2024-02-21 12:18 | 0.039952 | 3.0702e-05 | 4.9878 | 0.49695 | 0.40099 | 0.00038413
AB2A | 0.9635 | 0.0027001 | 2024-02-21 12:32 | 0.080298 | 4.0406e-05 | 4.9805 | 0.49725 | 0.80427 | 0.00061251
AB2B | 0.96822 | 0.0027133 | 2024-02-21 12:46 | 0.080298 | 4.0406e-05 | 4.9881 | 0.4971 | 0.80575 | 0.00061373
AB3 | 1.8787 | 0.0052648 | 2024-02-21 13:00 | 0.15812 | 6.4983e-05 | 4.9855 | 0.498 | 1.5829 | 0.0011133
AB2C | 0.9674 | 0.002711 | 2024-02-21 15:33 | 0.080298 | 4.0406e-05 | 4.9874 | 0.49635 | 0.80684 | 0.00061508
AB4 | 2.8467 | 0.0079775 | 2024-02-21 16:01 | 0.24215 | 9.4792e-05 | 4.985 | 0.49835 | 2.4222 | 0.0016757
[8 empty] | Gauss's method (GUM's Linear Approximation) |  |  |  | Monte Carlo method |  |  | Ratio wA | Ratio u(wA) | Z-score | Degree of
Sample ID | rAB | u(rAB) | Date | mA | mB | m = mB/mA | u(m) | wA | u(wA) | uR(wA) | R^2 | wA | u(wA) | uR(wA) | GS/MT | GS/MT | value | matching
001_A1 | 0.93544 | 0.0026215 | 2024-02-21 13:14 | 5.1055 | 0.49805 | 0.097551 | 5.5662e-05 | 0.07595 | 0.00023596 | 0.3107% | 1 | 0.07596 | 0.00025823 | 0.34% | 99.99% | 91.38% | 1.4724 | 96.8%
078_A1 | 0.93369 | 0.0026166 | 2024-02-21 13:28 | 5.0927 | 0.49655 | 0.097501 | 5.5802e-05 | 0.075767 | 0.00023547 | 0.3108% | 1 | 0.075764 | 0.00025375 | 0.3349% | 100% | 92.79% | 0.692 | 96.62%
155_A1 | 0.92497 | 0.0025921 | 2024-02-21 13:42 | 5.0823 | 0.4971 | 0.097809 | 5.5918e-05 | 0.075282 | 0.00023421 | 0.3111% | 1 | 0.075289 | 0.00025015 | 0.3323% | 99.99% | 93.63% | -1.3707 | 95.72%
232_A1 | 0.93336 | 0.0026156 | 2024-02-21 13:56 | 5.0881 | 0.4962 | 0.097522 | 5.5853e-05 | 0.075755 | 0.00023544 | 0.3108% | 1 | 0.07576 | 0.00026098 | 0.3445% | 99.99% | 90.22% | 0.6421 | 96.58%
310_A1 | 0.92931 | 0.0026043 | 2024-02-21 14:09 | 5.0793 | 0.49695 | 0.097838 | 5.5951e-05 | 0.075665 | 0.00023527 | 0.3109% | 1 | 0.075673 | 0.00025661 | 0.3391% | 99.99% | 91.68% | 0.25796 | 96.16%
001_A2 | 0.92897 | 0.0026033 | 2024-02-21 14:23 | 5.0987 | 0.49725 | 0.097524 | 5.5736e-05 | 0.075394 | 0.00023443 | 0.3109% | 1 | 0.075391 | 0.00025215 | 0.3345% | 100% | 92.97% | -0.89404 | 96.13%
078_A2 | 0.9285 | 0.002602 | 2024-02-21 14:37 | 5.0915 | 0.49745 | 0.097702 | 5.5816e-05 | 0.075492 | 0.00023475 | 0.311% | 1 | 0.075501 | 0.00025622 | 0.3394% | 99.99% | 91.62% | -0.47719 | 96.08%
155_A2 | 0.92999 | 0.0026062 | 2024-02-21 14:51 | 5.0783 | 0.4977 | 0.098004 | 5.5963e-05 | 0.07585 | 0.00023581 | 0.3109% | 1 | 0.075847 | 0.00025374 | 0.3345% | 100% | 92.93% | 1.0449 | 96.24%
232_A2 | 0.92417 | 0.0025899 | 2024-02-21 15:05 | 5.0769 | 0.4975 | 0.097993 | 5.5979e-05 | 0.075357 | 0.00023446 | 0.3111% | 1 | 0.075353 | 0.00025067 | 0.3327% | 100% | 93.53% | -1.0526 | 95.63%
310_A2 | 0.92699 | 0.0025978 | 2024-02-21 15:19 | 5.0769 | 0.4971 | 0.097913 | 5.5978e-05 | 0.07553 | 0.00023492 | 0.311% | 1 | 0.075539 | 0.00025625 | 0.3392% | 99.99% | 91.67% | -0.3148 | 95.93%
MOOS3 | 1.8423 | 0.0051628 | 2024-02-21 15:47 | 5.1433 | 0.49715 | 0.096659 | 5.5248e-05 | 0.15013 | 0.0004566 | 0.3041% | 1 | 0.15015 | 0.00052224 | 0.3478% | 99.99% | 87.43%
| CRM |  | Measurement |  |  | Comparability |  |  |  |  | Ratio
Sample | μ | unc | x | s | n | IRMM | ISO | Wald | t-test | Mood, B=2 | x/μ
MOOS3 | 0.1581 | 0.00115 | 0.15013 | 0.0004566 | 3 | FALSE | FALSE | FALSE | FALSE | FALSE | 0.9496
Note ICM-21-02 for MOOS-3 stability: nitrate (n = 14) w = 0.1535 μg/g, uw = 0.020 μg/g
At this point the MOOS-3 was not a reliable CRM for validation of nitrite. Comparison of the data with the independent UV-vis method is a better alternative for validation
n = 3 as the isotope ratio was measured in triplicate
Comparability tests are for information only
MOOS-3 is expired and not properly informative as a QC sample at this point
3. Determination of nitrite by UV-vis w1d1
[5 empty] | Baseline | Blank | Grav | Density | Matrix | uncertainty | Linear model: Gauss's method (GUM's Linear Approximation) |  |  |  | Monte Carlo method |  |  | Ratio Gauss vs Monte Carlo |  | Quadratic | Quadratic | Ratio linear
Nitrite signals |  |  |  |  | signal | signal | signal | signal | signal | signal | w(NO2) | u[w(NO2)] | uR[w(NO2)] | R^2 | w(NO2) | u[w(NO2)] | uR[w(NO2)] | w(NO2) | u[w(NO2)] | w1(NO2) | w2(NO2) | vs quadratic | Z-score | Degree of
Sample ID | Date | 541nm (Abs) | mS | mF | A(541) | A(541) | A(541) | A(541) | A(541) | A(541) | µg/g | µg/g | % |  | µg/g | µg/g | % | % | % | µg/g | µg/g | % | value | matching
BLS-DIW | 2/20/2024 3:31:31 PM | -0.000252
BSL-LNWS | 2/20/2024 3:32:18 PM | -0.000417
BSL-M4 | 2/20/2024 3:33:13 PM | -0.000417
BSL-M3 | 2/20/2024 3:33:58 PM | -0.000377
BLK-DIW | 2/20/2024 3:34:54 PM | -0.000218 |  |  | 3.4e-05
BLK-LNSW | 2/20/2024 3:35:48 PM | 2.3e-05 |  |  | 0.00044
CAL1 | 2/20/2024 3:37:24 PM | 0.045523 | 5.1142 | 5.3149 | 0.04594 | 0.045534 | 0.047322 | 0.046203 | 0.046203 | 0.00020774 | 0.039993 | 2.6398e-05
CAL2 | 2/20/2024 3:38:23 PM | 0.064783 | 5.1236 | 5.3247 | 0.0652 | 0.064794 | 0.067338 | 0.065746 | 0.065746 | 0.00029561 | 0.059684 | 3.7447e-05
CAL3 | 2/20/2024 3:39:20 PM | 0.086985 | 5.1178 | 5.3193 | 0.087402 | 0.086996 | 0.090421 | 0.088284 | 0.088284 | 0.00039694 | 0.078745 | 4.848e-05
CAL4 | 2/20/2024 3:40:15 PM | 0.1271 | 5.1091 | 5.3098 | 0.12751 | 0.12711 | 0.1321 | 0.12898 | 0.12898 | 0.00057991 | 0.11564 | 7.0175e-05
CAL5 | 2/20/2024 3:41:06 PM | 0.17165 | 5.1439 | 5.3446 | 0.17207 | 0.17166 | 0.17836 | 0.17415 | 0.17415 | 0.00078299 | 0.15764 | 9.5208e-05
001a_1 | 2/20/2024 3:42:46 PM | 0.083905 | 5.1043 | 5.3054 | 0.084322 | 0.084322 | 0.087644 | 0.085842 | 0.085028 | 0.00039995 | 0.076063 | 0.00055289 | 0.7269% | 0.99974 | 0.076074 | 0.0003999 | 0.5257% | 99.98% | 138.3% | 0.076032 | 42.4163 | 100% | 0.75089 | 96.93%
078a_1 | 2/20/2024 3:43:32 PM | 0.083475 | 5.0916 | 5.2925 | 0.083892 | 0.083892 | 0.087203 | 0.08541 | 0.0846 | 0.00039793 | 0.075672 | 0.00055258 | 0.7302% | 0.99974 | 0.075659 | 0.00039046 | 0.5161% | 100% | 141.5% | 0.075642 | 42.4167 | 100% | 0.043768 | 96.44%
155a_1 | 2/20/2024 3:44:12 PM | 0.083304 | 5.1017 | 5.3024 | 0.083721 | 0.083721 | 0.087015 | 0.085226 | 0.084418 | 0.00039708 | 0.075506 | 0.00055246 | 0.7317% | 0.99974 | 0.075507 | 0.00038957 | 0.5159% | 100% | 141.8% | 0.075476 | 42.4168 | 100% | -0.25688 | 96.23%
232a_1 | 2/20/2024 3:44:55 PM | 0.083229 | 5.0997 | 5.3008 | 0.083646 | 0.083646 | 0.086945 | 0.085158 | 0.08435 | 0.00039676 | 0.075444 | 0.00055241 | 0.7322% | 0.99974 | 0.075441 | 0.00038826 | 0.5147% | 100% | 142.3% | 0.075414 | 42.4169 | 100% | -0.36924 | 96.16%
310a_1 | 2/20/2024 3:45:37 PM | 0.084108 | 5.0892 | 5.2897 | 0.084525 | 0.084525 | 0.087855 | 0.086049 | 0.085232 | 0.00040091 | 0.076249 | 0.00055305 | 0.7253% | 0.99974 | 0.07625 | 0.00041856 | 0.5489% | 100% | 132.1% | 0.076219 | 42.4161 | 100% | 1.0889 | 97.16%
001a_2 | 2/20/2024 3:46:18 PM | 0.083172 | 5.1099 | 5.3108 | 0.083589 | 0.083589 | 0.086876 | 0.08509 | 0.084283 | 0.00039644 | 0.075382 | 0.00055237 | 0.7328% | 0.99974 | 0.07539 | 0.00040078 | 0.5316% | 99.99% | 137.8% | 0.075353 | 42.4169 | 100% | -0.48005 | 96.08%
078a_2 | 2/20/2024 3:46:59 PM | 0.083875 | 5.0938 | 5.2945 | 0.084292 | 0.084292 | 0.087614 | 0.085813 | 0.084999 | 0.00039981 | 0.076036 | 0.00055287 | 0.7271% | 0.99974 | 0.076033 | 0.00039133 | 0.5147% | 100% | 141.3% | 0.076006 | 42.4163 | 100% | 0.7026 | 96.9%
155a_2 | 2/20/2024 3:47:41 PM | 0.083345 | 5.0962 | 5.2966 | 0.083762 | 0.083762 | 0.087056 | 0.085266 | 0.084457 | 0.00039726 | 0.075542 | 0.00055248 | 0.7314% | 0.99974 | 0.075553 | 0.00039727 | 0.5258% | 99.98% | 139.1% | 0.075512 | 42.4168 | 100% | -0.1921 | 96.28%
232a_2 | 2/20/2024 3:48:24 PM | 0.0831 | 5.0844 | 5.2851 | 0.083517 | 0.083517 | 0.086814 | 0.085029 | 0.084222 | 0.00039616 | 0.075327 | 0.00055233 | 0.7332% | 0.99974 | 0.075326 | 0.00039135 | 0.5195% | 100% | 141.1% | 0.075298 | 42.417 | 100% | -0.58022 | 96.01%
310a_2 | 2/20/2024 3:49:07 PM | 0.083043 | 5.0968 | 5.2967 | 0.08346 | 0.08346 | 0.086734 | 0.084951 | 0.084145 | 0.00039579 | 0.075257 | 0.00055228 | 0.7339% | 0.99974 | 0.075279 | 0.00039051 | 0.5188% | 99.97% | 141.4% | 0.075227 | 42.4171 | 100% | -0.70767 | 95.92%
MOOS3 | 2/20/2024 3:50:04 PM | 0.1634 | 5.1452 | 5.3463 | 0.16378 | 0.16378 | 0.17017 | 0.16511 | 0.16511 | 0.00074234 | 0.14915 | 0.00095809 | 0.6423% | 0.99974 | 0.14912 | 0.00084883 | 0.5692% | 100% | 112.9% | 0.14917 | 42.3431 | 99.99%
| CRM |  | Measurement |  |  | Comparability |  |  |  |  | Ratio
Sample | μ | unc | x | s | n | IRMM | ISO | Wald | t-test | Mood, B=2 | x/μ
MOOS3 | 0.1581 | 0.00115 | 0.14915 | 0.00095809 | 3 | FALSE | FALSE | FALSE | FALSE | FALSE | 0.94342
Note ICM-21-02 for MOOS-3 stability: nitrate (n = 14) w = 0.1535 μg/g, uw = 0.020 μg/g
At this point the MOOS-3 was not a reliable CRM for validation of nitrite. Comparison of the data with the independent UV-vis method is a better alternative for validation
n = 3 as the isotope ratio was measured in triplicate
Comparability tests are for information only
MOOS-3 is expired and not properly informative as a QC sample at this point
4. Determination of phosphate by UV-vis w1d1
[5 empty] | Baseline | Blank | Grav | Density | Matrix | uncertainty | Linear model: Gauss's method (GUM's Linear Approximation) |  |  |  | Monte Carlo method |  |  | Ratio Gauss vs Monte Carlo |  | Quadratic | Quadratic | Ratio linear
Phosphate signals |  |  |  |  | signal | signal | signal | signal | signal | signal | w(PO4) | u[w(PO4)] | uR[w(PO4)] | R^2 | w(PO4) | u[w(PO4)] | uR[w(PO4)] | w(PO4) | u[w(PO4)] | w1(PO4) | w2(PO4) | vs quadratic | Z-score | Degree of
Sample ID | Date | 890nm (Abs) | mS | mF | A(890) | A(890) | A(890) | A(890) | A(890) | A(890) | µg/g | µg/g | % |  | µg/g | µg/g | % | % | % | µg/g | µg/g | % | value | matching
BLS-DIW | 2/20/2024 5:17:55 PM | -0.000594
BSL-LNWS | 2/20/2024 5:18:37 PM | -0.00262
BSL-M4 | 2/20/2024 5:19:59 PM | -0.001955
BSL-SALT | 2/20/2024 5:20:43 PM | 0.000197
BSL-M3 | 2/20/2024 5:21:31 PM | -0.002503
BLK-DIW | 2/20/2024 5:22:30 PM | -0.000605 |  |  | -1.1e-05
BLK-LNSW | 2/20/2024 5:23:42 PM | -0.002192 |  |  | 0.000428
CAL1 | 2/20/2024 5:24:54 PM | 0.1337 | 6.1619 | 6.9459 | 0.13632 | 0.13588 | 0.15317 | 0.14955 | 0.14955 | 0.00058225 | 0.12929 | 4.9919e-05
CAL2 | 2/20/2024 5:25:55 PM | 0.20198 | 6.1426 | 6.9272 | 0.2046 | 0.20416 | 0.23023 | 0.22479 | 0.22479 | 0.0008752 | 0.19295 | 6.3128e-05
CAL3 | 2/20/2024 5:26:51 PM | 0.26679 | 6.1377 | 6.9215 | 0.26941 | 0.26897 | 0.30332 | 0.29615 | 0.29615 | 0.001153 | 0.25458 | 7.7396e-05
CAL4 | 2/20/2024 5:28:05 PM | 0.3908 | 6.1486 | 6.9332 | 0.39342 | 0.39298 | 0.44313 | 0.43266 | 0.43266 | 0.0016845 | 0.37386 | 0.00010681
CAL5 | 2/20/2024 5:29:08 PM | 0.53541 | 6.1432 | 6.9281 | 0.53803 | 0.53759 | 0.60628 | 0.59195 | 0.59195 | 0.0023047 | 0.50964 | 0.00014247
001a_1 | 2/20/2024 5:30:41 PM | 0.27373 | 6.1259 | 6.9116 | 0.27569 | 0.27569 | 0.31105 | 0.30466 | 0.30466 | 0.0011862 | 0.26245 | 0.0011008 | 0.4194% | 0.99997 | 0.26249 | 0.0011298 | 0.4304% | 99.99% | 97.43% | 0.26253 | -187.5869 | 99.97% | -0.65256 | 102.5%
078a_1 | 2/20/2024 5:31:27 PM | 0.27549 | 6.1294 | 6.9151 | 0.27744 | 0.27744 | 0.313 | 0.30657 | 0.30657 | 0.0011936 | 0.2641 | 0.0011064 | 0.4189% | 0.99997 | 0.2641 | 0.0011851 | 0.4487% | 100% | 93.36% | 0.26418 | -187.5886 | 99.97% | 0.84204 | 103.2%
155a_1 | 2/20/2024 5:32:15 PM | 0.27478 | 6.1318 | 6.917 | 0.27673 | 0.27673 | 0.31217 | 0.30575 | 0.30575 | 0.0011904 | 0.2634 | 0.001104 | 0.4191% | 0.99997 | 0.26342 | 0.0011261 | 0.4275% | 99.99% | 98.03% | 0.26348 | -187.5879 | 99.97% | 0.2051 | 102.9%
232a_1 | 2/20/2024 5:33:00 PM | 0.27465 | 6.1444 | 6.9254 | 0.2766 | 0.2766 | 0.31176 | 0.30536 | 0.30536 | 0.0011889 | 0.26305 | 0.0011028 | 0.4192% | 0.99997 | 0.26307 | 0.0011244 | 0.4274% | 99.99% | 98.08% | 0.26314 | -187.5875 | 99.97% | -0.10663 | 102.8%
310a_1 | 2/20/2024 5:33:44 PM | 0.27584 | 6.1258 | 6.9089 | 0.2778 | 0.2778 | 0.31331 | 0.30687 | 0.30687 | 0.0011948 | 0.26436 | 0.0011072 | 0.4188% | 0.99997 | 0.26442 | 0.001123 | 0.4247% | 99.98% | 98.6% | 0.26444 | -187.5888 | 99.97% | 1.0768 | 103.3%
001a_2 | 2/20/2024 5:34:23 PM | 0.27459 | 6.1241 | 6.908 | 0.27655 | 0.27655 | 0.31195 | 0.30554 | 0.30554 | 0.0011896 | 0.26321 | 0.0011033 | 0.4192% | 0.99997 | 0.26318 | 0.0011088 | 0.4213% | 100% | 99.5% | 0.26329 | -187.5877 | 99.97% | 0.036121 | 102.8%
078a_2 | 2/20/2024 5:35:06 PM | 0.27548 | 6.1416 | 6.9258 | 0.27743 | 0.27743 | 0.31285 | 0.30642 | 0.30642 | 0.001193 | 0.26397 | 0.0011059 | 0.419% | 0.99997 | 0.26399 | 0.0011343 | 0.4297% | 99.99% | 97.5% | 0.26406 | -187.5884 | 99.97% | 0.72732 | 103.1%
155a_2 | 2/20/2024 5:36:02 PM | 0.2747 | 6.1178 | 6.9012 | 0.27665 | 0.27665 | 0.31208 | 0.30566 | 0.30566 | 0.0011901 | 0.26332 | 0.0011037 | 0.4192% | 0.99997 | 0.26332 | 0.0011353 | 0.4312% | 100% | 97.22% | 0.2634 | -187.5878 | 99.97% | 0.13387 | 102.9%
232a_2 | 2/20/2024 5:36:55 PM | 0.27456 | 6.1391 | 6.9231 | 0.27652 | 0.27652 | 0.31183 | 0.30542 | 0.30542 | 0.0011891 | 0.26311 | 0.001103 | 0.4192% | 0.99997 | 0.26306 | 0.0011282 | 0.4289% | 100% | 97.77% | 0.26319 | -187.5876 | 99.97% | -0.055927 | 102.8%
310a_2 | 2/20/2024 5:37:45 PM | 0.27196 | 6.1183 | 6.9025 | 0.27391 | 0.27391 | 0.30902 | 0.30267 | 0.30267 | 0.0011784 | 0.26074 | 0.001095 | 0.42% | 0.99997 | 0.26076 | 0.0011349 | 0.4352% | 99.99% | 96.49% | 0.26082 | -187.5852 | 99.97% | -2.2061 | 101.9%
SALT1_113 | 2/20/2024 5:39:17 PM | 0.15677 | 6.1391 | 6.924 | 0.15657 | 0.15657 | 0.17659 | 0.17259 | 0.17259 | 0.00067195 | 0.14863 | 0.00082015 | 0.5518% | 0.99997 | 0.14864 | 0.00076821 | 0.5168% | 99.99% | 106.8% | 0.14857 | -187.4729 | 100%
MOOS3 | 2/20/2024 5:40:31 PM | 0.15339 | 6.1537 | 6.9364 | 0.15589 | 0.15589 | 0.17572 | 0.17049 | 0.17049 | 0.00066377 | 0.14682 | 0.00081796 | 0.5571% | 0.99997 | 0.1468 | 0.00076843 | 0.5234% | 100% | 106.4% | 0.14675 | -187.4711 | 100%
| CRM |  | Measurement |  |  | Comparability |  |  |  |  | Ratio
Sample | μ | unc | x | s | n | IRMM | ISO | Wald | t-test | Mood, B=2 | x/μ
SALT1_113 | 0.14942 | 0.00097429 | 0.14863 | 0.00082015 | 3 | TRUE | TRUE | TRUE | TRUE | TRUE | 0.99471
MOOS3 | 0.147 | 0.0065 | 0.14682 | 0.00081796 | 3 | TRUE | TRUE | TRUE | TRUE | TRUE | 0.99875
Note ICM-21-02 for MOOS-3 stability: phosphate (n = 21) w = 0.1569 μg/g, uw = 0.0095 μg/g
n = 3 as the isotope ratio was measured in triplicate
Comparability tests are for information only
MOOS-3 is expired and not properly informative as a QC sample at this point
5. Determination of silicate by UV-vis w1d2
[5 empty] | Baseline | Blank | Grav | Density | Matrix | uncertainty | Linear model: Gauss's method (GUM's Linear Approximation) |  |  |  | Monte Carlo method |  |  | Ratio Gauss vs Monte Carlo |  | Quadratic | Quadratic | Ratio linear
Silicate signals |  |  |  |  | signal | signal | signal | signal | signal | signal | w(SiO2) | u[w(SiO2)] | uR[w(SiO2)] | R^2 | w(SiO2) | u[w(SiO2)] | uR[w(SiO2)] | w(SiO2) | u[w(SiO2)] | w1(SiO2) | w2(SiO2) | vs quadratic | Z-score | Degree of
Sample ID | Date | 810nm (Abs) | mS | mF | A(810) | A(810) | A(810) | A(810) | A(810) | A(810) | µg/g | µg/g | % |  | µg/g | µg/g | % | % | % | µg/g | µg/g | % | value | matching
BLS-DIW | 2/21/2024 11:43:54 AM | 0.000828
BSL-LNWS | 2/21/2024 11:44:34 AM | -0.000282
BSL-M4 | 2/21/2024 11:45:09 AM | -0.000234
BSL-SALT | 2/21/2024 11:45:45 AM | 6e-06
BSL-M3 | 2/21/2024 11:46:19 AM | 3.8e-05
BLK-DIW | 2/21/2024 11:47:01 AM | 0.000217 |  |  | -0.000611
BLK-LNSW | 2/21/2024 11:47:41 AM | 0.011604 |  |  | 0.011886
CAL1 | 2/21/2024 11:48:24 AM | 0.071961 | 3.0641 | 3.3866 | 0.072243 | 0.059746 | 0.066033 | 0.064472 | 0.064472 | 0.00041986 | 0.21333 | 0.00018185
CAL2 | 2/21/2024 11:49:04 AM | 0.10009 | 3.0712 | 3.3921 | 0.10037 | 0.087873 | 0.097056 | 0.094762 | 0.094762 | 0.00061712 | 0.31368 | 0.00023434
CAL3 | 2/21/2024 11:49:47 AM | 0.12859 | 3.066 | 3.388 | 0.12887 | 0.11637 | 0.1286 | 0.12556 | 0.12556 | 0.00081766 | 0.41323 | 0.0002922
CAL4 | 2/21/2024 11:50:34 AM | 0.18786 | 3.0712 | 3.3927 | 0.18814 | 0.17564 | 0.19403 | 0.18945 | 0.18945 | 0.0012337 | 0.62413 | 0.00042204
CAL5 | 2/21/2024 11:51:30 AM | 0.24244 | 3.0686 | 3.3907 | 0.24272 | 0.23022 | 0.25438 | 0.24836 | 0.24836 | 0.0016174 | 0.82148 | 0.00054651
CAL6 | 2/21/2024 12:00:27 PM | 0.66193 | 3.0615 | 3.3835 | 0.66221 | 0.64971 | 0.71806 | 0.70109 | 0.70109 | 0.0045657 | 2.3004 | 0.0015
001a_1 | 2/21/2024 11:52:29 AM | 0.11764 | 3.0451 | 3.3669 | 0.11787 | 0.11787 | 0.13033 | 0.12765 | 0.12765 | 0.0008313 | 0.42144 | 0.002829 | 0.6713% | 0.99997 | 0.42157 | 0.0029818 | 0.7073% | 99.97% | 94.87% | 0.42038 | 38.1739 | 100.3% | -1.5831 | 101.3%
078a_1 | 2/21/2024 11:53:08 AM | 0.11819 | 3.0387 | 3.3608 | 0.11843 | 0.11843 | 0.13098 | 0.12829 | 0.12829 | 0.00083543 | 0.42354 | 0.0028419 | 0.671% | 0.99997 | 0.42355 | 0.003074 | 0.7258% | 100% | 92.45% | 0.42246 | 38.1718 | 100.3% | -0.84989 | 101.9%
155a_1 | 2/21/2024 11:53:53 AM | 0.11924 | 3.0331 | 3.3545 | 0.11947 | 0.11947 | 0.13213 | 0.12942 | 0.12942 | 0.0008428 | 0.42728 | 0.0028649 | 0.6705% | 0.99997 | 0.42734 | 0.0031124 | 0.7283% | 99.99% | 92.05% | 0.42618 | 38.1681 | 100.3% | 0.45887 | 102.8%
232a_1 | 2/21/2024 11:54:31 AM | 0.11856 | 3.0344 | 3.3559 | 0.11879 | 0.11879 | 0.13138 | 0.12868 | 0.12868 | 0.00083799 | 0.42484 | 0.0028499 | 0.6708% | 0.99997 | 0.42491 | 0.0031073 | 0.7313% | 99.98% | 91.72% | 0.42376 | 38.1705 | 100.3% | -0.39463 | 102.2%
310a_1 | 2/21/2024 11:55:10 AM | 0.11979 | 3.0448 | 3.3653 | 0.12002 | 0.12002 | 0.13265 | 0.12993 | 0.12993 | 0.00084612 | 0.42896 | 0.0028753 | 0.6703% | 0.99997 | 0.42897 | 0.0030937 | 0.7212% | 100% | 92.94% | 0.42786 | 38.1664 | 100.3% | 1.0484 | 103.2%
001a_2 | 2/21/2024 11:55:49 AM | 0.11821 | 3.0285 | 3.3502 | 0.11844 | 0.11844 | 0.13103 | 0.12833 | 0.12833 | 0.00083575 | 0.4237 | 0.0028429 | 0.671% | 0.99997 | 0.4238 | 0.0030368 | 0.7166% | 99.98% | 93.61% | 0.42262 | 38.1716 | 100.3% | -0.79372 | 101.9%
078a_2 | 2/21/2024 11:56:26 AM | 0.11983 | 3.0285 | 3.3497 | 0.12006 | 0.12006 | 0.13279 | 0.13006 | 0.13006 | 0.000847 | 0.42941 | 0.0028781 | 0.6702% | 0.99997 | 0.42958 | 0.0031066 | 0.7232% | 99.96% | 92.64% | 0.4283 | 38.166 | 100.3% | 1.2044 | 103.3%
155a_2 | 2/21/2024 11:57:02 AM | 0.11837 | 3.0305 | 3.352 | 0.1186 | 0.1186 | 0.13119 | 0.12849 | 0.12849 | 0.00083675 | 0.42421 | 0.002846 | 0.6709% | 0.99997 | 0.42433 | 0.0030372 | 0.7158% | 99.97% | 93.7% | 0.42313 | 38.1711 | 100.3% | -0.61528 | 102%
232a_2 | 2/21/2024 11:57:35 AM | 0.11974 | 3.0397 | 3.3614 | 0.11997 | 0.11997 | 0.13266 | 0.12994 | 0.12994 | 0.00084619 | 0.42899 | 0.0028755 | 0.6703% | 0.99997 | 0.429 | 0.003221 | 0.7508% | 100% | 89.27% | 0.42789 | 38.1664 | 100.3% | 1.0597 | 103.2%
310a_2 | 2/21/2024 11:58:10 AM | 0.11927 | 3.0404 | 3.3619 | 0.11951 | 0.11951 | 0.13214 | 0.12942 | 0.12942 | 0.00084284 | 0.4273 | 0.002865 | 0.6705% | 0.99997 | 0.42741 | 0.0030877 | 0.7224% | 99.97% | 92.79% | 0.4262 | 38.1681 | 100.3% | 0.46525 | 102.8%
SALT1_113 | 2/21/2024 11:59:02 AM | 0.1498 | 3.0366 | 3.3572 | 0.14979 | 0.14979 | 0.16561 | 0.16185 | 0.16185 | 0.001054 | 0.53442 | 0.0035478 | 0.6639% | 0.99997 | 0.53453 | 0.0039071 | 0.7309% | 99.98% | 90.81% | 0.53309 | 38.0612 | 100.2%
MOOS3 | 2/21/2024 11:59:42 AM | 0.49205 | 3.062 | 3.3841 | 0.49201 | 0.49201 | 0.54377 | 0.52757 | 0.52757 | 0.0034357 | 1.7326 | 0.011375 | 0.6565% | 0.99999 | 1.7329 | 0.014978 | 0.8643% | 99.98% | 75.95% | 1.7354 | -217.9594 | 99.84%
| CRM |  | Measurement |  |  | Comparability |  |  |  |  | Ratio
Sample | μ | unc | x | s | n | IRMM | ISO | Wald | t-test | Mood, B=2 | x/μ
SALT1_113 | 0.52039 | 0.0083172 | 0.53442 | 0.0035478 | 3 | TRUE | TRUE | TRUE | FALSE | TRUE | 1.027
MOOS3 | 1.77 | 0.02 | 1.7326 | 0.011375 | 3 | TRUE | TRUE | TRUE | FALSE | TRUE | 0.97889
Note ICM-21-02 for MOOS-3 stability: silicate (n = 7) w = 1.747 μg/g, uw = 0.010 μg/g
n = 3 as the isotope ratio was measured in triplicate
Comparability tests are for information only
MOOS-3 is expired and not properly informative as a QC sample at this point
6. Determination of nitrite by UV-vis w1d2
[5 empty] | Baseline | Blank | Grav | Density | Matrix | uncertainty | Linear model: Gauss's method (GUM's Linear Approximation) |  |  |  | Monte Carlo method |  |  | Ratio Gauss vs Monte Carlo |  | Quadratic | Quadratic | Ratio linear
Nitrite signals |  |  |  |  | signal | signal | signal | signal | signal | signal | w(NO2) | u[w(NO2)] | uR[w(NO2)] | R^2 | w(NO2) | u[w(NO2)] | uR[w(NO2)] | w(NO2) | u[w(NO2)] | w1(NO2) | w2(NO2) | vs quadratic | Z-score | Degree of
Sample ID | Date | 541nm (Abs) | mS | mF | A(541) | A(541) | A(541) | A(541) | A(541) | A(541) | µg/g | µg/g | % |  | µg/g | µg/g | % | % | % | µg/g | µg/g | % | value | matching
BLS-DIW | 2/21/2024 1:35:24 PM | 1.2e-05
BSL-LNWS | 2/21/2024 1:36:29 PM | -8.7e-05
BSL-M4 | 2/21/2024 1:37:26 PM | -0.000158
BSL-M3 | 2/21/2024 1:38:07 PM | 4.3e-05
BLK-DIW | 2/21/2024 1:38:43 PM | 9.1e-05 |  |  | 7.9e-05
BLK-LNSW | 2/21/2024 1:39:25 PM | 5e-06 |  |  | 9.2e-05
CAL1 | 2/21/2024 1:40:11 PM | 0.044309 | 5.1151 | 5.3161 | 0.044396 | 0.044383 | 0.046127 | 0.045037 | 0.045037 | 0.0002025 | 0.039684 | 2.5954e-05
CAL2 | 2/21/2024 1:40:59 PM | 0.065183 | 5.1159 | 5.3164 | 0.06527 | 0.065257 | 0.067816 | 0.066213 | 0.066213 | 0.00029771 | 0.058943 | 2.9982e-05
CAL3 | 2/21/2024 1:41:49 PM | 0.087086 | 5.13 | 5.3313 | 0.087173 | 0.08716 | 0.09058 | 0.088439 | 0.088439 | 0.00039764 | 0.078513 | 3.5105e-05
CAL4 | 2/21/2024 1:42:41 PM | 0.13109 | 5.1338 | 5.3355 | 0.13118 | 0.13117 | 0.13632 | 0.1331 | 0.1331 | 0.00059843 | 0.11848 | 4.7002e-05
CAL5 | 2/21/2024 1:43:29 PM | 0.2008 | 5.1382 | 5.3392 | 0.20088 | 0.20087 | 0.20873 | 0.20379 | 0.20379 | 0.00091629 | 0.18389 | 6.8379e-05
001a_1 | 2/21/2024 1:44:41 PM | 0.084117 | 5.0788 | 5.2805 | 0.084275 | 0.084275 | 0.087622 | 0.08582 | 0.085006 | 0.0003822 | 0.075665 | 0.00045504 | 0.6014% | 0.99991 | 0.075686 | 0.00037816 | 0.4996% | 99.97% | 120.3% | 0.075363 | 5.0359 | 100.4% | 0.1082 | 96.73%
078a_1 | 2/21/2024 1:45:23 PM | 0.08412 | 5.0862 | 5.2881 | 0.084278 | 0.084278 | 0.087623 | 0.085822 | 0.085008 | 0.00038221 | 0.075666 | 0.00045504 | 0.6014% | 0.99991 | 0.075678 | 0.0003829 | 0.506% | 99.98% | 118.8% | 0.075364 | 5.0359 | 100.4% | 0.11116 | 96.74%
155a_1 | 2/21/2024 1:46:03 PM | 0.084168 | 5.0939 | 5.2952 | 0.084326 | 0.084326 | 0.087659 | 0.085857 | 0.085043 | 0.00038237 | 0.075698 | 0.00045511 | 0.6012% | 0.99991 | 0.075704 | 0.00038156 | 0.504% | 99.99% | 119.3% | 0.075395 | 5.0358 | 100.4% | 0.18041 | 96.78%
232a_1 | 2/21/2024 1:46:44 PM | 0.084109 | 5.0959 | 5.297 | 0.084267 | 0.084267 | 0.087592 | 0.085792 | 0.084978 | 0.00038208 | 0.075639 | 0.00045498 | 0.6015% | 0.99991 | 0.07564 | 0.00040043 | 0.5294% | 100% | 113.6% | 0.075338 | 5.0359 | 100.4% | 0.051202 | 96.7%
310a_1 | 2/21/2024 1:47:23 PM | 0.083853 | 5.0911 | 5.2923 | 0.084011 | 0.084011 | 0.087333 | 0.085537 | 0.084726 | 0.00038094 | 0.07541 | 0.00045449 | 0.6027% | 0.99991 | 0.075432 | 0.00037698 | 0.4998% | 99.97% | 120.6% | 0.075113 | 5.0361 | 100.4% | -0.45097 | 96.41%
MOOS3 | 2/21/2024 1:48:24 PM | 0.16705 | 4.9189 | 5.12 | 0.16701 | 0.16701 | 0.17383 | 0.16866 | 0.16866 | 0.00075831 | 0.15155 | 0.00079823 | 0.5267% | 0.99991 | 0.15157 | 0.00085432 | 0.5636% | 99.98% | 93.43% | 0.15124 | 4.96 | 100.2%
| CRM |  | Measurement |  |  | Comparability |  |  |  |  | Ratio
Sample | μ | unc | x | s | n | IRMM | ISO | Wald | t-test | Mood, B=2 | x/μ
MOOS3 | 0.1581 | 0.00115 | 0.15155 | 0.00079823 | 3 | FALSE | FALSE | FALSE | FALSE | FALSE | 0.95856
Note ICM-21-02 for MOOS-3 stability: nitrate (n = 14) w = 0.1535 μg/g, uw = 0.020 μg/g
At this point the MOOS-3 was not a reliable CRM for validation of nitrite. Comparison of the data with the independent UV-vis method is a better alternative for validation
n = 3 as the isotope ratio was measured in triplicate
Comparability tests are for information only
MOOS-3 is expired and not properly informative as a QC sample at this point
2024-03-05 and 2024-03-06 Characterization MOOS-4 [WEEK 2; DAY 1 and 2]: sample preparation (verified 2024-04-19)
Balance calibration
Weigh the unoped units
UV-vis performance verification
Preparation of standard solutions silicate in DIW
| Origin |  | m0 = vial |  | m1 = m0 + standard |  | m2 = m1 + DIW | [7 empty] | Final
ID solution | w(SiO2) | u[w(SiO2)] | load | unload | load | unload | load | unload | m0 | m1 | m2 | u(m0) | u(m1) | u(m2) | w(SiO2) | u[w(SiO2)] | Target | Traciability | Prep. Date | Exp. Date | t | Rel Hum | Analyst | Balance
unique | µg/g | µg/g | g | g | g | g | g | g | g | g | g | g | g | g | µg/g | µg/g | % |  | yyyymmdd | yyyymmdd | °C | % | initials | model
Sw2d1-D | 21432.0637 | 0 | 11.7875 | 0 | 12.8359 | 0 | 41.9695 | 0 | 11.7875 | 12.8359 | 41.9695 | 0.0002 | 0.0002 | 0.0002 | 744.4628 | 0.20097 | N/A | NIST | 20240305 | 20240306 | 22.7 | 31.3 | EP | MS204S
In a 50 mL bottle
Preparation of standard solutions phosphate and silicate in LNSW
| Origin |  | Origin |  | m0 = vial |  | m1 = m0 + PO4 standard |  | m2 = m1 + SiO2 standard |  | m3 = m2 + LNSW | [9 empty] | Final |  | Final |  | Target
ID solution | w(PO4) | u[w(PO4)] | w(SiO2) | u[w(SiO2)] | load | unload | load | unload | load | unload | load | unload | m0 | m1 | m2 | m3 | u(m0) | u(m1) | u(m2) | u(m3) | w(PO4) | u[w(PO4)] | w(SiO2) | u[w(SiO2)] | PO4 | SiO2 | Traciability | Prep. Date | Exp. Date | t | Rel Hum | Analyst | Balance
unique | µg/g | µg/g | µg/g | µg/g | g | g | g | g | g | g | g | g | g | g | g | g | g | g | g | g | µg/g | µg/g | µg/g | µg/g | % | % |  | yyyymmdd | yyyymmdd | °C | % | initials | model
PSw2d1-ST | 1000.5 | 0 | 744.4628 | 0.20097 | 11.8017 | 0 | 12.575 | 0 | 14.312 | -0.0001 | 42.239 | 0.0002 | 11.8017 | 12.575 | 14.312 | 42.2389 | 0.0002 | 0.0002 | 0.0002 | 0.002 | 25.4191 | 0.0094476 | 42.4865 | 0.013685 | N/A | N/A | NIST | 20240305 | 20240306 | 22.7 | 31.3 | EP | MS204S
In a 50 mL bottle
| Origin |  | Origin |  | m0 = vial |  | m1 = m0 + standard |  | m2 = m1 + LNSW | [7 empty] | Final |  | Final |  | Target
ID solution | w(PO4) | u[w(PO4)] | w(SiO2) | u[w(SiO2)] | load | unload | load | unload | load | unload | m0 | m1 | m2 | u(m0) | u(m1) | u(m2) | w(PO4) | u[w(PO4)] | w(SiO2) | u[w(SiO2)] | PO4 | SiO2 | Traciability | Prep. Date | Exp. Date | t | Rel Hum | Analyst | Balance
unique | µg/g | µg/g | µg/g | µg/g | g | g | g | g | g | g | g | g | g | g | g | g | µg/g | µg/g | µg/g | µg/g | % | % |  | yyyymmdd | yyyymmdd | °C | % | initials | model
PSw2d1-0 | 25.4191 | 0.0094476 | 42.4865 | 0.013685 | 8.1929 | -0.0001 | 15.8493 | 0 | 36.2341 | -0.0002 | 8.1929 | 15.8493 | 36.2342 | 0.0002 | 0.0002 | 0.002 | 6.9404 | 0.0026396 | 11.6005 | 0.0038518 | N/A | N/A | NIST | 20240305 | 20240306 | 22.7 | 31.3 | EP | MS204S
PSw2d1-1 | 6.9404 | 0.0026396 | 11.6005 | 0.0038518 | 8.182 | -0.0001 | 8.6977 | 0 | 35.8851 | 0.0001 | 8.1821 | 8.6977 | 35.885 | 0.0002 | 0.0002 | 0.002 | 0.12919 | 8.6735e-05 | 0.21593 | 0.00013933 | 98.2% | 101.9% | NIST | 20240305 | 20240306 | 22.7 | 31.3 | EP | MS204S
PSw2d1-2 | 6.9404 | 0.0026396 | 11.6005 | 0.0038518 | 8.1876 | 0 | 8.9603 | -0.0002 | 35.5318 | 0.0002 | 8.1876 | 8.9604 | 35.5317 | 0.0002 | 0.0002 | 0.002 | 0.19615 | 0.00010453 | 0.32785 | 0.0001638 | 99.4% | 103.2% | NIST | 20240305 | 20240306 | 22.7 | 31.3 | EP | MS204S
PSw2d1-3 | 6.9404 | 0.0026396 | 11.6005 | 0.0038518 | 8.1214 | -0.0001 | 9.1493 | -0.0001 | 35.66 | 0.0001 | 8.1214 | 9.1494 | 35.6599 | 0.0002 | 0.0002 | 0.002 | 0.25906 | 0.00012307 | 0.433 | 0.00018938 | 98.46% | 102.2% | NIST | 20240305 | 20240306 | 22.7 | 31.3 | EP | MS204S
PSw2d1-4 | 6.9404 | 0.0026396 | 11.6005 | 0.0038518 | 8.2168 | 0 | 9.7594 | 0 | 35.7377 | 0.0001 | 8.2168 | 9.7594 | 35.7376 | 0.0002 | 0.0002 | 0.002 | 0.38902 | 0.00016669 | 0.65023 | 0.00025116 | 98.57% | 102.3% | NIST | 20240305 | 20240306 | 22.7 | 31.3 | EP | MS204S
PSw2d1-5 | 6.9404 | 0.0026396 | 11.6005 | 0.0038518 | 8.1149 | -0.0002 | 10.1706 | -0.0001 | 35.9109 | 0.0001 | 8.115 | 10.1707 | 35.9108 | 0.0002 | 0.0002 | 0.002 | 0.51328 | 0.00021089 | 0.85791 | 0.00031453 | 97.54% | 101.2% | NIST | 20240305 | 20240306 | 22.7 | 31.3 | EP | MS204S
PSw2d1-6 | 6.9404 | 0.0026396 | 11.6005 | 0.0038518 | 8.2155 | 0 | 12.8051 | 0.0002 | 37.3468 | 0.0001 | 8.2155 | 12.805 | 37.3468 | 0.0002 | 0.0002 | 0.002 | 1.0934 | 0.00042798 | 1.8276 | 0.00062995 | N/A | 101.5% | NIST | 20240305 | 20240306 | 22.7 | 31.3 | EP | MS204S
Preparation of standard solutions nitrite, nitrate in water for IDMS experiment
| Origin |  | m0 = vial |  | m1 = m0 + standard |  | m2 = m1 + DIW | [7 empty] | Final
ID solution | w(NO2) | u[w(NO2)] | load | unload | load | unload | load | unload | m0 | m1 | m2 | u(m0) | u(m1) | u(m2) | w(NO2) | u[w(NO2)] | Target | Traciability | Prep. Date | Exp. Date | t | Rel Hum | Analyst | Balance
unique | µg/g | µg/g | g | g | g | g | g | g | g | g | g | g | g | g | µg/g | µg/g | % |  | yyyymmdd | yyyymmdd | °C | % | initials | model
N2w2-D | 1004.0619 | 0 | 11.7917 | 0.0002 | 14.7713 | 0 | 42.2424 | 0.0001 | 11.7916 | 14.7713 | 42.2424 | 0.0002 | 0.0002 | 0.0002 | 98.2506 | 0.0093708 | N/A | NIST | 20240305 | 20240306 | 22.7 | 31.3 | EP | MS204S
In a 50 mL bottle
| Origin |  | Origin |  | m0 = vial |  | m1 = m0 + NO2 standard |  | m2 = m1 + NO3 standard |  | m3 = m2 + DIW | [9 empty] | Final |  | Final |  | Target
ID solution | w(NO2) | u[w(NO2)] | w(NO3) | u[w(NO3)] | load | unload | load | unload | load | unload | load | unload | m0 | m1 | m2 | m3 | u(m0) | u(m1) | u(m2) | u(m3) | w(NO2) | u[w(NO2)] | w(NO3) | u[w(NO3)] | NO2 | NO3 | Traciability | Prep. Date | Exp. Date | t | Rel Hum | Analyst | Balance
unique | µg/g | µg/g | µg/g | µg/g | g | g | g | g | g | g | g | g | g | g | g | g | g | g | g | g | µg/g | µg/g | µg/g | µg/g | % | % |  | yyyymmdd | yyyymmdd | °C | % | initials | model
N2N3w2-ST | 98.2506 | 0.0093708 | 1000.6 | 0 | 11.797 | 0 | 12.5489 | 0.0001 | 13.553 | 0.0001 | 39.9599 | 0.0001 | 11.797 | 12.5488 | 13.553 | 39.9598 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 2.6229 | 0.0010183 | 35.6747 | 0.010056 | N/A | N/A | NIST | 20240305 | 20240306 | 22.7 | 31.3 | EP | MS204S
In a 50 mL bottle
| Origin |  | Origin |  | m0 = mass of vial |  | m1 = m0 + mass of standard |  | m2 = m1 + mass of DIW | [7 empty] | Final |  | Final |  | Target
ID solution | w(NO2) | u[w(NO2)] | w(NO3) | u[w(NO3)] | load | unload | load | unload | load | unload | m0 | m1 | m2 | u(m0) | u(m1) | u(m2) | w(NO2) | u[w(NO2)] | w(NO3) | u[w(NO3)] | NO2 | NO3 | Traciability | Prep. Date | Exp. Date | t | Rel Hum | Analyst | Balance
unique | µg/g | µg/g | µg/g | µg/g | g | g | g | g | g | g | g | g | g | g | g | g | µg/g | µg/g | µg/g | µg/g | % | % |  | yyyymmdd | yyyymmdd | °C | % | initials | model
N2N3w2-1 | 2.6229 | 0.0010183 | 35.6747 | 0.010056 | 8.203 | 0.0002 | 8.604 | 0.0001 | 35.208 | 0.0001 | 8.2029 | 8.6039 | 35.2079 | 0.0002 | 0.0002 | 0.0002 | 0.038953 | 3.1362e-05 | 0.5298 | 0.00040242 | 100.4% | 99.79% | NIST | 20240305 | 20240306 | 22.7 | 31.3 | EP | MS204S
N2N3w2-2 | 2.6229 | 0.0010183 | 35.6747 | 0.010056 | 8.2763 | -0.0001 | 9.0776 | -0.0001 | 35.2139 | -0.0001 | 8.2764 | 9.0777 | 35.214 | 0.0002 | 0.0002 | 0.0002 | 0.078024 | 4.0948e-05 | 1.0612 | 0.00047949 | 100.6% | 99.94% | NIST | 20240305 | 20240306 | 22.7 | 31.3 | EP | MS204S
N2N3w2-3 | 2.6229 | 0.0010183 | 35.6747 | 0.010056 | 8.2601 | 0.0002 | 9.8675 | 0 | 35.3754 | 0.0001 | 8.26 | 9.8675 | 35.3753 | 0.0002 | 0.0002 | 0.0002 | 0.1555 | 6.6299e-05 | 2.1149 | 0.00070309 | 100.2% | 99.59% | NIST | 20240305 | 20240306 | 22.7 | 31.3 | EP | MS204S
N2N3w2-4 | 2.6229 | 0.0010183 | 35.6747 | 0.010056 | 8.2216 | -0.0001 | 10.6394 | 0 | 35.3063 | -0.0001 | 8.2217 | 10.6394 | 35.3064 | 0.0002 | 0.0002 | 0.0002 | 0.23414 | 9.4969e-05 | 3.1846 | 0.00097243 | 100.6% | 99.97% | NIST | 20240305 | 20240306 | 22.7 | 31.3 | EP | MS204S
In 30 mL bottles
Preparation of isotope dilution blends for NO2 and NO3 analysis by HS GC-MS
|  |  |  | m0 = mass of vial |  |  | m1 = m0 + mass of sample/ST |  |  | m2 = m1 + IntST
|  |  |  | load | unload | mass | load | unload | mass | load | unload | mass |  | mA | mB
Type | Description |  | ID | g | g | g | g | g | g | g | g | g | Date | g | g
BLK | Reag BLK DIW |  | BLK-DIW | [9 empty] | 2024-03-05
BLK | Reag BLK 3.5% NaCl |  | BLK-NaCl
BLK | IntST BLK DIW |  | BLK-IS-DIW | 7.3247 | -0.0002 | 7.3248 | 15.3167 | -0.0001 | 15.3168 | 16.1132 | 0.0001 | 16.1131 |  | 7.992 | 0.7964
BLK | InST BLK 3.5% NaCl |  | BLK-IS-NaCl | 7.3203 | 0 | 7.3203 | 15.5114 | 0.0001 | 15.5114 | 16.3071 | 0 | 16.3071 |  | 8.1911 | 0.79575
CAL | A*B-1 |  | AB1 | 7.3451 | 0.0001 | 7.3451 | 15.3301 | 0.0001 | 15.33 | 16.1255 | 0 | 16.1255 |  | 7.985 | 0.79545
CAL | A*B-2A |  | AB2a | 7.3262 | 0.0001 | 7.3262 | 15.3146 | 0.0002 | 15.3145 | 16.1108 | 0 | 16.1108 |  | 7.9884 | 0.7963
CAL | A*B-2B |  | AB2b | 7.3069 | -0.0002 | 7.307 | 15.336 | 0 | 15.336 | 16.1321 | 0 | 16.1321 |  | 8.029 | 0.7961
CAL | A*B-3 |  | AB3 | 7.2979 | 0 | 7.2979 | 15.2764 | 0 | 15.2764 | 16.0721 | 0.0001 | 16.072 |  | 7.9785 | 0.79565
SAMPLE | MOOS052a_i |  | 052a_1 | 7.3385 | -0.0001 | 7.3385 | 15.5317 | -0.0001 | 15.5318 | 16.3272 | 0 | 16.3272 |  | 8.1932 | 0.79545
SAMPLE | MOOS130a_i |  | 130a_1 | 7.2987 | 0 | 7.2987 | 15.4391 | 0 | 15.4391 | 16.2338 | 0 | 16.2338 |  | 8.1404 | 0.7947
SAMPLE | MOOS207a_i |  | 207a_1 | 7.3196 | -0.0001 | 7.3197 | 15.4413 | 0 | 15.4413 | 16.2377 | 0 | 16.2377 |  | 8.1216 | 0.7964
SAMPLE | MOOS284a_i |  | 284a_1 | 7.3131 | 0.0002 | 7.313 | 15.4336 | 0 | 15.4336 | 16.2296 | 0.0001 | 16.2295 |  | 8.1206 | 0.79595
SAMPLE | MOOS360a_i |  | 360a_1 | 7.3442 | -0.0001 | 7.3442 | 15.4595 | 0 | 15.4595 | 16.2525 | 0.0001 | 16.2524 |  | 8.1152 | 0.79295
SAMPLE | MOOS052a_ii |  | 052a_2 | 7.3095 | 0.0002 | 7.3094 | 15.4286 | -0.0002 | 15.4287 | 16.2229 | -0.0002 | 16.223 |  | 8.1193 | 0.7943
SAMPLE | MOOS130a_ii |  | 130a_2 | 7.3357 | -0.0002 | 7.3358 | 15.4546 | -0.0001 | 15.4546 | 16.2519 | 0 | 16.2519 |  | 8.1188 | 0.79725
SAMPLE | MOOS207a_ii |  | 207a_2 | 7.3204 | 0 | 7.3204 | 15.4232 | 0.0001 | 15.4231 | 16.217 | 0 | 16.217 |  | 8.1028 | 0.79385
SAMPLE | MOOS284a_ii |  | 284a_2 | 7.312 | 0 | 7.312 | 15.416 | 0 | 15.416 | 16.2134 | 0 | 16.2134 |  | 8.104 | 0.7974
SAMPLE | MOOS360a_ii |  | 360a_2 | 7.3153 | 0.0001 | 7.3152 | 15.427 | -0.0001 | 15.427 | 16.2211 | -0.0002 | 16.2212 |  | 8.1118 | 0.79415
QC | SALT-1 |  | SALT1_309 | 7.3137 | 0.0001 | 7.3136 | 15.5048 | 0 | 15.5048 | 16.2976 | -0.0003 | 16.2978 |  | 8.1912 | 0.79295
QC | MOOS-3 |  | MOOS3 | 7.3386 | 0 | 7.3386 | 15.57 | 0 | 15.57 | 16.3663 | 0 | 16.3663 |  | 8.2314 | 0.7963
CAL | A*B-2C |  | AB2c | 7.3103 | 0 | 7.3103 | 15.3302 | 0 | 15.3302 | 16.1249 | -0.0001 | 16.125 |  | 8.0199 | 0.79475
CAL | A*B-4 |  | AB4 | 7.3461 | 0 | 7.3461 | 15.3428 | -0.0001 | 15.3429 | 16.1393 | 0.0001 | 16.1392 | 2024-03-05 | 7.9968 | 0.7964
Sample without derivatization: 2 mL sample without further treatment
RA and RA* : 2 mL of sample (or AB2 standard ) following derivatization method
Day_1: determination of nitrate after nitrite removal by sulfamic acid
2 mL blend in 10 mL combiPAL vial + 0.1 mL sulfamic acid + 0.04 mL triethyloxonium tetrafluoroborate solution in water
Reag BLK = 1.82 mL of blank solution
Measure using SALT1 method
Vial washed "old method", rinsed 3 times with water, incubated with water for one week and blow dried with pressurized nitrogen
Preparation of external calibration for PO4 analysis by UV-vis
|  |  |  | m0 = mass of vial |  |  | m1 = m0 + mass of sample/ST |  |  | m2 = m1 + Reagents
|  |  |  | load | unload | mass | load | unload | mass | load | unload | mass |  | mS | mF
Type | Description |  | ID | g | g | g | g | g | g | g | g | g | Date | g | g
BLK | Baseline DIW |  | BSL-DIW | [9 empty] | 2024-03-05
BLK | Baseline LNSW |  | BSL-LNSW
BLK | Baseline MOOS-4 |  | BSL-M4
BLK | Baseline SALT-1 |  | BSL-SALT
BLK | Baseline MOOS-3 |  | BSL-M3
BLK | Reag BLK DIW |  | BLK-DIW
BLK | Reag BLK LNSW |  | BLK-LNSW
STD | CAL-1 |  | CAL-1 | 7.3141 | 0 | 7.3141 | 13.4858 | 0 | 13.4858 | 14.3227 | -0.0001 | 14.3227 |  | 6.1717 | 7.0086
STD | CAL-2 |  | CAL-2 | 7.3134 | 0 | 7.3134 | 13.4729 | 0 | 13.4729 | 14.3091 | 0 | 14.3091 |  | 6.1595 | 6.9957
STD | CAL-3 |  | CAL-3 | 7.3144 | 0 | 7.3144 | 13.4452 | 0 | 13.4452 | 14.2821 | 0.0001 | 14.282 |  | 6.1308 | 6.9676
STD | CAL-4 |  | CAL-4 | 7.3048 | 0 | 7.3048 | 13.4504 | -0.0001 | 13.4505 | 14.2884 | 0 | 14.2884 |  | 6.1456 | 6.9836
STD | CAL-5 |  | CAL-5 | 7.2849 | 0 | 7.2849 | 13.4371 | -0.0002 | 13.4372 | 14.2742 | -0.0001 | 14.2743 |  | 6.1523 | 6.9893
SAMPLE | MOOS052a_i |  | 052a_1 | 7.3135 | 0.0001 | 7.3135 | 13.4543 | 0 | 13.4543 | 14.2922 | 0.0001 | 14.2921 |  | 6.1408 | 6.9787
SAMPLE | MOOS130a_i |  | 130a_1 | 7.2738 | 0 | 7.2738 | 13.4335 | 0 | 13.4335 | 14.2717 | 0 | 14.2717 |  | 6.1597 | 6.9979
SAMPLE | MOOS207a_i |  | 207a_1 | 7.3392 | 0.0001 | 7.3392 | 13.4878 | 0 | 13.4878 | 14.326 | -0.0001 | 14.3261 |  | 6.1486 | 6.9869
SAMPLE | MOOS284a_i |  | 284a_1 | 7.3268 | 0 | 7.3268 | 13.467 | 0.0001 | 13.467 | 14.3028 | 0 | 14.3028 |  | 6.1402 | 6.976
SAMPLE | MOOS360a_i |  | 360a_1 | 7.2905 | -0.0001 | 7.2905 | 13.4471 | -0.0001 | 13.4472 | 14.2835 | 0 | 14.2835 |  | 6.1566 | 6.993
SAMPLE | MOOS052a_ii |  | 052a_2 | 7.3269 | 0 | 7.3269 | 13.4527 | 0 | 13.4527 | 14.2879 | 0 | 14.2879 |  | 6.1258 | 6.961
SAMPLE | MOOS130a_ii |  | 130a_2 | 7.3459 | 0 | 7.3459 | 13.49 | 0 | 13.49 | 14.3073 | 0 | 14.3073 |  | 6.1441 | 6.9614
SAMPLE | MOOS207a_ii |  | 207a_2 | 7.3024 | 0.0001 | 7.3023 | 13.4349 | 0.0001 | 13.4349 | 14.2709 | 0.0001 | 14.2708 |  | 6.1325 | 6.9685
SAMPLE | MOOS284a_ii |  | 284_a2 | 7.3347 | -0.0001 | 7.3347 | 13.481 | 0 | 13.481 | 14.3175 | 0 | 14.3175 |  | 6.1463 | 6.9828
SAMPLE | MOOS360a_ii |  | 360a_2 | 7.3126 | -0.0001 | 7.3126 | 13.4603 | -0.0001 | 13.4604 | 14.2987 | -0.0001 | 14.2988 |  | 6.1477 | 6.9861
QC | SALT-1 |  | SALT1_309 | 7.346 | 0.0001 | 7.346 | 13.4907 | 0 | 13.4907 | 14.328 | 0.0001 | 14.3279 |  | 6.1448 | 6.982
QC | MOOS-3 |  | MOOS3 | 7.3318 | -0.0001 | 7.3319 | 13.4951 | -0.0001 | 13.4952 | 14.324 | -0.0001 | 14.324 | 2024-03-05 | 6.1633 | 6.9922
Use the 5 cm cuvette and zero the instrument against genpure water
Measure ABS at 890 nm with existing method
=== Day break carry on next day ===
Day_2: determination of nitrite
2 mL blend prepared previous day in 10 mL combiPAL vial + 0.04 mL triethyloxonium tetrafluoroborate solution in water
Reag BLK = 1.82 mL of blank solution
Measure using MOOS4 method
Preparation of external calibration for SiO2 analysis by UV-vis
|  |  |  | m0 = mass of vial |  |  | m1 = m0 + mass of sample/ST |  |  | m2 = m1 + Reagents
|  |  |  | load | unload | mass | load | unload | mass | load | unload | mass |  | mS | mF
Type | Description |  | ID | g | g | g | g | g | g | g | g | g | Date | g | g
BLK | Baseline DIW |  | BSL-DIW | [9 empty] | 2024-03-06
BLK | Baseline LNSW |  | BSL-LNSW
BLK | Baseline MOOS-4 |  | BSL-M4
BLK | Baseline SALT-1 |  | BSL-SALT
BLK | Baseline MOOS-3 |  | BSL-M3
BLK | Reag BLK DIW |  | BLK-DIW
BLK | Reag BLK LNSW |  | BLK-LNSW
STD | CAL-1 |  | CAL-1 | 6.1097 | -0.0002 | 6.1098 | 9.1814 | -0.0002 | 9.1815 | 9.5035 | 0.0001 | 9.5035 |  | 3.0717 | 3.3937
STD | CAL-2 |  | CAL-2 | 6.1178 | 0.0002 | 6.1177 | 9.1826 | -0.0001 | 9.1827 | 9.5056 | -0.0001 | 9.5056 |  | 3.065 | 3.3879
STD | CAL-3 |  | CAL-3 | 6.1163 | 0 | 6.1163 | 9.1884 | -0.0001 | 9.1884 | 9.5115 | 0 | 9.5115 |  | 3.0721 | 3.3952
STD | CAL-4 |  | CAL-4 | 6.1222 | -0.0001 | 6.1223 | 9.1918 | 0.0001 | 9.1918 | 9.5143 | 0.0001 | 9.5143 |  | 3.0695 | 3.392
STD | CAL-5 |  | CAL-5 | 6.1077 | 0 | 6.1077 | 9.1814 | 0.0001 | 9.1814 | 9.5038 | -0.0001 | 9.5038 |  | 3.0736 | 3.3961
SAMPLE | MOOS052a_i |  | 052a_1 | 6.1329 | 0.0001 | 6.1329 | 9.1905 | -0.0001 | 9.1905 | 9.5135 | 0.0001 | 9.5135 |  | 3.0577 | 3.3806
SAMPLE | MOOS130a_i |  | 130a_1 | 6.1407 | 0.0001 | 6.1406 | 9.2043 | 0 | 9.2043 | 9.5257 | -0.0002 | 9.5258 |  | 3.0636 | 3.3852
SAMPLE | MOOS207a_i |  | 207a_1 | 6.1027 | 0.0001 | 6.1026 | 9.1644 | 0 | 9.1644 | 9.4873 | 0 | 9.4873 |  | 3.0618 | 3.3846
SAMPLE | MOOS284a_i |  | 284a_1 | 6.0976 | 0.0003 | 6.0975 | 9.1644 | 0 | 9.1644 | 9.4878 | -0.0001 | 9.4878 |  | 3.067 | 3.3904
SAMPLE | MOOS360a_i |  | 360a_1 | 6.1181 | -0.0001 | 6.1181 | 9.1862 | 0 | 9.1862 | 9.5089 | 0 | 9.5089 |  | 3.068 | 3.3908
SAMPLE | MOOS052a_ii |  | 052a_2 | 6.1271 | -0.0003 | 6.1273 | 9.185 | 0 | 9.185 | 9.5082 | 0.0001 | 9.5082 |  | 3.0578 | 3.3809
SAMPLE | MOOS130a_ii |  | 130a_2 | 6.1392 | 0.0002 | 6.1391 | 9.2107 | -0.0001 | 9.2107 | 9.5332 | 0.0002 | 9.5331 |  | 3.0716 | 3.394
SAMPLE | MOOS207a_ii |  | 207a_2 | 6.147 | 0 | 6.147 | 9.2148 | 0 | 9.2148 | 9.5361 | 0 | 9.5361 |  | 3.0678 | 3.3891
SAMPLE | MOOS284a_ii |  | 284_a2 | 6.1261 | 0 | 6.1261 | 9.1923 | 0 | 9.1923 | 9.5149 | 0 | 9.5149 |  | 3.0662 | 3.3888
SAMPLE | MOOS360a_ii |  | 360a_2 | 6.1402 | -0.0001 | 6.1402 | 9.2141 | 0.0001 | 9.2141 | 9.5366 | 0 | 9.5366 |  | 3.0738 | 3.3963
QC | SALT-1 |  | SALT1_309 | 6.1483 | 0 | 6.1483 | 9.2192 | -0.0001 | 9.2193 | 9.542 | 0.0002 | 9.5419 |  | 3.071 | 3.3936
QC | MOOS-3 |  | MOOS3 | 6.1438 | 0 | 6.1438 | 9.2272 | 0 | 9.2272 | 9.5499 | -0.0002 | 9.55 |  | 3.0834 | 3.4062
STD | CAL-6 |  | CAL-6 | 6.1472 | 0 | 6.1472 | 9.2105 | 0.0001 | 9.2104 | 9.5332 | 0.0001 | 9.5332 | 2024-03-06 | 3.0633 | 3.386
2024-03-05 and 2024-03-06 Characterization MOOS-4 [WEEK 2; DAY 1 and 2]: data analysis (verified 2024-04-19)
1. Determination of nitrate by IDMS w2d1
Sample ID | rA*B | u(rA*B) | Date | wA* | u(wA*) | mA* | mB | wA*(mA*/mB) | u(wA*(mA*/mB))
AB1 | 0.5277 | 0.011118 | 2024-03-05 12:58 | 0.5298 | 0.00040242 | 7.985 | 0.79545 | 5.3183 | 0.0044643
AB2A | 0.98614 | 0.020777 | 2024-03-05 13:08 | 1.0612 | 0.00047949 | 7.9884 | 0.7963 | 10.6458 | 0.0061301
AB2B | 0.99424 | 0.020947 | 2024-03-05 13:18 | 1.0612 | 0.00047949 | 8.029 | 0.7961 | 10.7026 | 0.0061633
AB3 | 1.9432 | 0.04094 | 2024-03-05 13:28 | 2.1149 | 0.00070309 | 7.9785 | 0.79565 | 21.2078 | 0.010349
AB2C | 1.0272 | 0.021641 | 2024-03-05 15:39 | 1.0612 | 0.00047949 | 8.0199 | 0.79475 | 10.7087 | 0.0061708
[8 empty] | Gauss's method (GUM's Linear Approximation) |  |  |  | Monte Carlo method |  |  | Ratio wA | Ratio u(wA) | Z-score | Degree of
Sample ID | rAB | u(rAB) | Date | mA | mB | m = mB/mA | u(m) | wA | u(wA) | uR(wA) | R^2 | wA | u(wA) | uR(wA) | GS/MT | GS/MT | value | matching
052_A1 | 1.0189 | 0.0090444 | 2024-03-05 13:38 | 8.1932 | 0.79545 | 0.097087 | 3.4684e-05 | 1.056 | 0.016089 | 1.524% | 0.99929 | 1.0564 | 0.016823 | 1.592% | 99.96% | 95.64% | 0.24615 | 101.6%
130_A1 | 1.0023 | 0.0088959 | 2024-03-05 13:48 | 8.1404 | 0.7947 | 0.097624 | 3.4911e-05 | 1.0435 | 0.015908 | 1.524% | 0.99929 | 1.0446 | 0.016623 | 1.591% | 99.89% | 95.7% | -0.53631 | 99.97%
207_A1 | 0.99549 | 0.0088377 | 2024-03-05 13:58 | 8.1216 | 0.7964 | 0.098059 | 3.4993e-05 | 1.0406 | 0.015871 | 1.525% | 0.99929 | 1.0417 | 0.015938 | 1.53% | 99.9% | 99.58% | -0.71435 | 99.3%
284_A1 | 1.0113 | 0.0089772 | 2024-03-05 14:08 | 8.1206 | 0.79595 | 0.098016 | 3.4997e-05 | 1.0576 | 0.016119 | 1.524% | 0.99929 | 1.0587 | 0.016193 | 1.529% | 99.9% | 99.54% | 0.35055 | 100.9%
360_A1 | 1.0015 | 0.0088916 | 2024-03-05 14:18 | 8.1152 | 0.79295 | 0.097711 | 3.5019e-05 | 1.0435 | 0.015911 | 1.525% | 0.99929 | 1.0442 | 0.016737 | 1.603% | 99.93% | 95.06% | -0.53196 | 99.9%
052_A2 | 1.0212 | 0.0090647 | 2024-03-05 14:28 | 8.1193 | 0.7943 | 0.097829 | 3.5002e-05 | 1.0665 | 0.016249 | 1.524% | 0.99929 | 1.0678 | 0.016975 | 1.59% | 99.89% | 95.72% | 0.90605 | 101.9%
130_A2 | 1.0254 | 0.0091022 | 2024-03-05 14:38 | 8.1188 | 0.79725 | 0.098197 | 3.5005e-05 | 1.0752 | 0.016379 | 1.523% | 0.99929 | 1.076 | 0.017367 | 1.614% | 99.93% | 94.31% | 1.4503 | 102.3%
207_A2 | 1.009 | 0.0089556 | 2024-03-05 14:49 | 8.1028 | 0.79385 | 0.097973 | 3.5074e-05 | 1.0546 | 0.016073 | 1.524% | 0.99929 | 1.0557 | 0.016147 | 1.53% | 99.9% | 99.54% | 0.16031 | 100.6%
284_A2 | 0.9937 | 0.0088201 | 2024-03-05 14:59 | 8.104 | 0.7974 | 0.098396 | 3.507e-05 | 1.0422 | 0.015895 | 1.525% | 0.99929 | 1.0436 | 0.017442 | 1.671% | 99.87% | 91.13% | -0.61486 | 99.12%
360_A2 | 0.99699 | 0.0088513 | 2024-03-05 15:09 | 8.1118 | 0.79415 | 0.097901 | 3.5035e-05 | 1.0406 | 0.015869 | 1.525% | 0.99929 | 1.0412 | 0.016832 | 1.617% | 99.94% | 94.28% | -0.71587 | 99.45%
SALT1_309 | 1.1249 | 0.0099878 | 2024-03-05 15:19 | 8.1912 | 0.79295 | 0.096806 | 3.4692e-05 | 1.1685 | 0.017729 | 1.517% | 0.99929 | 1.1705 | 0.018737 | 1.601% | 99.83% | 94.62%
MOOS3 | 1.3327 | 0.01182 | 2024-03-05 15:29 | 8.2314 | 0.7963 | 0.096739 | 3.4522e-05 | 1.3935 | 0.020207 | 1.45% | 0.99929 | 1.3946 | 0.021747 | 1.559% | 99.92% | 92.92%
| CRM |  | Measurement |  |  | Comparability |  |  |  |  | Ratio
Sample | μ | unc | x | s | n | IRMM | ISO | Wald | t-test | Mood, B=2 | x/μ
SALT1_309 | 1.1448 | 0.011301 | 1.1685 | 0.017729 | 3 | TRUE | TRUE | TRUE | TRUE | TRUE | 1.0208
MOOS3 | 1.384 | 0.0075 | 1.3935 | 0.020207 | 3 | TRUE | TRUE | TRUE | TRUE | TRUE | 1.0069
Note ICM-21-02 for MOOS-3 stability: nitrate (n = 21) w = 1.3921 μg/g, uw = 0.020 μg/g
n = 3 as the isotope ratio was measured in triplicate
Comparability tests are for information only
MOOS-3 is expired and not properly informative as a QC sample at this point
2. Determination of phosphate by UV-vis w2d1
[5 empty] | Baseline | Blank | Grav | Density | Matrix | uncertainty | Linear model: Gauss's method (GUM's Linear Approximation) |  |  |  | Monte Carlo method |  |  | Ratio Gauss vs Monte Carlo |  | Quadratic | Quadratic | Ratio linear
Phosphate signals |  |  |  |  | signal | signal | signal | signal | signal | signal | w(PO4) | u[w(PO4)] | uR[w(PO4)] | R^2 | w(PO4) | u[w(PO4)] | uR[w(PO4)] | w(PO4) | u[w(PO4)] | w1(PO4) | w2(PO4) | vs quadratic | Z-score | Degree of
Sample ID | Date | 890nm (Abs) | mS | mF | A(890) | A(890) | A(890) | A(890) | A(890) | A(890) | µg/g | µg/g | % |  | µg/g | µg/g | % | % | % | µg/g | µg/g | % | value | matching
BLS-DIW | 3/5/2024 2:15:27 PM | -0.000859
BSL-LNWS | 3/5/2024 2:16:17 PM | -0.001958
BSL-M4 | 3/5/2024 2:16:59 PM | -0.001641
BSL-SALT | 3/5/2024 2:17:47 PM | 0.000148
BSL-M3 | 3/5/2024 2:18:34 PM | -0.002501
BLK-DIW | 3/5/2024 2:19:35 PM | -0.000337 |  |  | 0.000522
BLK-LNSW | 3/5/2024 2:20:26 PM | -0.002227 |  |  | -0.000269
CAL1 | 3/5/2024 2:21:31 PM | 0.13287 | 6.1717 | 7.0086 | 0.13483 | 0.13483 | 0.15312 | 0.1495 | 0.1495 | 0.00043396 | 0.12919 | 8.6735e-05
CAL2 | 3/5/2024 2:22:33 PM | 0.20334 | 6.1595 | 6.9957 | 0.2053 | 0.2053 | 0.23317 | 0.22766 | 0.22766 | 0.00066085 | 0.19615 | 0.00010453
CAL3 | 3/5/2024 2:23:35 PM | 0.26915 | 6.1308 | 6.9676 | 0.27111 | 0.27111 | 0.30812 | 0.30083 | 0.30083 | 0.00087325 | 0.25906 | 0.00012307
CAL4 | 3/5/2024 2:24:35 PM | 0.40547 | 6.1456 | 6.9836 | 0.40742 | 0.40742 | 0.46297 | 0.45203 | 0.45203 | 0.0013121 | 0.38902 | 0.00016669
CAL5 | 3/5/2024 2:25:25 PM | 0.53436 | 6.1523 | 6.9893 | 0.53632 | 0.53632 | 0.60929 | 0.59489 | 0.59489 | 0.0017268 | 0.51328 | 0.00021089
052a_1 | 3/5/2024 2:26:44 PM | 0.27287 | 6.1408 | 6.9787 | 0.27451 | 0.27451 | 0.31197 | 0.30555 | 0.30555 | 0.00088695 | 0.26335 | 0.00080082 | 0.3041% | 0.99999 | 0.26335 | 0.00084412 | 0.3205% | 100% | 94.87% | 0.26294 | 39.4329 | 100.2% | -0.76392 | 101.2%
130a_1 | 3/5/2024 2:27:38 PM | 0.27442 | 6.1597 | 6.9979 | 0.27606 | 0.27606 | 0.31363 | 0.30718 | 0.30718 | 0.00089167 | 0.26475 | 0.00080454 | 0.3039% | 0.99999 | 0.26478 | 0.00083618 | 0.3158% | 99.99% | 96.22% | 0.26433 | 39.4315 | 100.2% | 0.98309 | 101.8%
207a_1 | 3/5/2024 2:28:21 PM | 0.27327 | 6.1486 | 6.9869 | 0.27491 | 0.27491 | 0.31239 | 0.30596 | 0.30596 | 0.00088814 | 0.2637 | 0.00080175 | 0.304% | 0.99999 | 0.26372 | 0.00082777 | 0.3139% | 99.99% | 96.86% | 0.26329 | 39.4325 | 100.2% | -0.32385 | 101.4%
284a_1 | 3/5/2024 2:29:01 PM | 0.27351 | 6.1402 | 6.976 | 0.27515 | 0.27515 | 0.3126 | 0.30618 | 0.30618 | 0.00088876 | 0.26388 | 0.00080224 | 0.304% | 0.99999 | 0.26391 | 0.00085946 | 0.3257% | 99.99% | 93.34% | 0.26347 | 39.4323 | 100.2% | -0.095319 | 101.5%
360a_1 | 3/5/2024 2:29:44 PM | 0.27404 | 6.1566 | 6.993 | 0.27568 | 0.27568 | 0.31313 | 0.3067 | 0.3067 | 0.00089027 | 0.26433 | 0.00080343 | 0.3039% | 0.99999 | 0.26437 | 0.00084402 | 0.3193% | 99.99% | 95.19% | 0.26392 | 39.4319 | 100.2% | 0.46372 | 101.6%
052a_2 | 3/5/2024 2:30:29 PM | 0.27306 | 6.1258 | 6.961 | 0.2747 | 0.2747 | 0.31215 | 0.30573 | 0.30573 | 0.00088747 | 0.2635 | 0.00080123 | 0.3041% | 0.99999 | 0.26351 | 0.0008263 | 0.3136% | 100% | 96.97% | 0.26309 | 39.4327 | 100.2% | -0.57077 | 101.3%
130a_2 | 3/5/2024 2:31:12 PM | 0.27565 | 6.1441 | 6.9614 | 0.27729 | 0.27729 | 0.31418 | 0.30772 | 0.30772 | 0.00089324 | 0.26521 | 0.00080578 | 0.3038% | 0.99999 | 0.26521 | 0.00084252 | 0.3177% | 100% | 95.64% | 0.2648 | 39.431 | 100.2% | 1.5644 | 102%
207a_2 | 3/5/2024 2:32:00 PM | 0.27207 | 6.1325 | 6.9685 | 0.27371 | 0.27371 | 0.31102 | 0.30463 | 0.30463 | 0.00088427 | 0.26255 | 0.0007987 | 0.3042% | 0.99999 | 0.26256 | 0.00083128 | 0.3166% | 100% | 96.08% | 0.26214 | 39.4337 | 100.2% | -1.7579 | 100.9%
284a_2 | 3/5/2024 2:32:40 PM | 0.27345 | 6.1463 | 6.9828 | 0.27509 | 0.27509 | 0.31254 | 0.30611 | 0.30611 | 0.00088856 | 0.26383 | 0.00080209 | 0.304% | 0.99999 | 0.26386 | 0.00084732 | 0.3211% | 99.99% | 94.66% | 0.26341 | 39.4324 | 100.2% | -0.1672 | 101.4%
360a_2 | 3/5/2024 2:33:28 PM | 0.27409 | 6.1477 | 6.9861 | 0.27573 | 0.27573 | 0.31333 | 0.30689 | 0.30689 | 0.00089082 | 0.2645 | 0.00080387 | 0.3039% | 0.99999 | 0.26448 | 0.00085039 | 0.3215% | 100% | 94.53% | 0.26408 | 39.4317 | 100.2% | 0.66781 | 101.7%
SALT1_309 | 3/5/2024 2:34:57 PM | 0.1569 | 6.1448 | 6.982 | 0.15676 | 0.15676 | 0.17811 | 0.17407 | 0.17407 | 0.0005053 | 0.14999 | 0.00055144 | 0.3676% | 0.99999 | 0.14999 | 0.00059312 | 0.3954% | 100% | 92.97% | 0.15026 | 39.5455 | 99.82%
MOOS3 | 3/5/2024 2:35:55 PM | 0.16153 | 6.1633 | 6.9922 | 0.16404 | 0.16404 | 0.1861 | 0.18055 | 0.18055 | 0.0005241 | 0.15558 | 0.0005603 | 0.3601% | 0.99999 | 0.15558 | 0.00058432 | 0.3756% | 100% | 95.89% | 0.1558 | 39.54 | 99.86%
| CRM |  | Measurement |  |  | Comparability |  |  |  |  | Ratio
Sample | μ | unc | x | s | n | IRMM | ISO | Wald | t-test | Mood, B=2 | x/μ
SALT1_309 | 0.14918 | 0.00097276 | 0.14999 | 0.00055144 | 3 | TRUE | TRUE | TRUE | TRUE | TRUE | 1.0054
MOOS3 | 0.147 | 0.0065 | 0.15558 | 0.0005603 | 3 | TRUE | TRUE | TRUE | FALSE | TRUE | 1.0584
Note ICM-21-02 for MOOS-3 stability: phosphate (n = 21) w = 0.1569 μg/g, uw = 0.0095 μg/g
n = 3 as the isotope ratio was measured in triplicate
Comparability tests are for information only
MOOS-3 is expired and not properly informative as a QC sample at this point
3. Determination of nitrite by IDMS w2d2
Sample ID | rA*B | u(rA*B) | Date | wA* | u(wA*) | mA* | mB | wA*(mA*/mB) | u(wA*(mA*/mB))
AB1 | 0.48329 | 0.0014433 | 2024-03-06 11:57 | 0.038953 | 3.1362e-05 | 7.985 | 0.79545 | 0.39103 | 0.00034444
AB2A | 0.94403 | 0.0028191 | 2024-03-06 12:12 | 0.078024 | 4.0948e-05 | 7.9884 | 0.7963 | 0.78272 | 0.00049679
AB2B | 0.94992 | 0.0028367 | 2024-03-06 12:26 | 0.078024 | 4.0948e-05 | 8.029 | 0.7961 | 0.7869 | 0.00049948
AB3 | 1.8523 | 0.0055315 | 2024-03-06 12:40 | 0.1555 | 6.6299e-05 | 7.9785 | 0.79565 | 1.5593 | 0.00086735
AB2C | 0.94856 | 0.0028327 | 2024-03-06 15:16 | 0.078024 | 4.0948e-05 | 8.0199 | 0.79475 | 0.78734 | 0.00050003
AB4 | 2.7626 | 0.0082498 | 2024-03-06 15:45 | 0.23414 | 9.4969e-05 | 7.9968 | 0.7964 | 2.351 | 0.0012702
[8 empty] | Gauss's method (GUM's Linear Approximation) |  |  |  | Monte Carlo method |  |  | Ratio wA | Ratio u(wA) | Z-score | Degree of
Sample ID | rAB | u(rAB) | Date | mA | mB | m = mB/mA | u(m) | wA | u(wA) | uR(wA) | R^2 | wA | u(wA) | uR(wA) | GS/MT | GS/MT | value | matching
052_A1 | 0.93791 | 0.0028009 | 2024-03-06 12:54 | 8.1932 | 0.79545 | 0.097087 | 3.4684e-05 | 0.075441 | 0.00023785 | 0.3153% | 1 | 0.075445 | 0.00027178 | 0.3602% | 99.99% | 87.51% | 0.10159 | 98.99%
130_A1 | 0.93257 | 0.0027849 | 2024-03-06 13:08 | 8.1404 | 0.7947 | 0.097624 | 3.4911e-05 | 0.075416 | 0.00023783 | 0.3154% | 1 | 0.075426 | 0.00027221 | 0.3609% | 99.99% | 87.37% | -0.004776 | 98.42%
207_A1 | 0.92738 | 0.0027694 | 2024-03-06 13:23 | 8.1216 | 0.7964 | 0.098059 | 3.4993e-05 | 0.075319 | 0.00023757 | 0.3154% | 1 | 0.075315 | 0.00026754 | 0.3552% | 100% | 88.8% | -0.41316 | 97.88%
284_A1 | 0.93208 | 0.0027834 | 2024-03-06 13:37 | 8.1206 | 0.79595 | 0.098016 | 3.4997e-05 | 0.075677 | 0.00023865 | 0.3154% | 1 | 0.075675 | 0.00027004 | 0.3568% | 100% | 88.38% | 1.0938 | 98.37%
360_A1 | 0.93057 | 0.0027789 | 2024-03-06 13:51 | 8.1152 | 0.79295 | 0.097711 | 3.5019e-05 | 0.075316 | 0.00023754 | 0.3154% | 1 | 0.075312 | 0.00026777 | 0.3555% | 100% | 88.71% | -0.42306 | 98.21%
052_A2 | 0.93151 | 0.0027818 | 2024-03-06 14:05 | 8.1193 | 0.7943 | 0.097829 | 3.5002e-05 | 0.075486 | 0.00023806 | 0.3154% | 1 | 0.075494 | 0.00028213 | 0.3737% | 99.99% | 84.38% | 0.28841 | 98.31%
130_A2 | 0.92812 | 0.0027716 | 2024-03-06 14:19 | 8.1188 | 0.79725 | 0.098197 | 3.5005e-05 | 0.075486 | 0.00023809 | 0.3154% | 1 | 0.075494 | 0.00026875 | 0.356% | 99.99% | 88.59% | 0.29133 | 97.95%
207_A2 | 0.92816 | 0.0027718 | 2024-03-06 14:34 | 8.1028 | 0.79385 | 0.097973 | 3.5074e-05 | 0.075318 | 0.00023757 | 0.3154% | 1 | 0.075323 | 0.00027615 | 0.3666% | 99.99% | 86.03% | -0.41787 | 97.96%
284_A2 | 0.92371 | 0.0027584 | 2024-03-06 14:48 | 8.104 | 0.7974 | 0.098396 | 3.507e-05 | 0.07527 | 0.00023746 | 0.3155% | 1 | 0.075274 | 0.00026896 | 0.3573% | 99.99% | 88.29% | -0.61845 | 97.49%
360_A2 | 0.93032 | 0.0027782 | 2024-03-06 15:02 | 8.1118 | 0.79415 | 0.097901 | 3.5035e-05 | 0.075441 | 0.00023793 | 0.3154% | 1 | 0.07545 | 0.00028184 | 0.3736% | 99.99% | 84.42% | 0.10215 | 98.19%
MOOS3 | 1.8899 | 0.0056438 | 2024-03-06 15:31 | 8.2314 | 0.7963 | 0.096739 | 3.4522e-05 | 0.15407 | 0.00048067 | 0.312% | 1 | 0.1541 | 0.00056225 | 0.3649% | 99.98% | 85.49%
| CRM |  | Measurement |  |  | Comparability |  |  |  |  | Ratio
Sample | μ | unc | x | s | n | IRMM | ISO | Wald | t-test | Mood, B=2 | x/μ
MOOS3 | 0.1581 | 0.00115 | 0.15407 | 0.00048067 | 3 | FALSE | FALSE | TRUE | FALSE | FALSE | 0.97452
Note ICM-21-02 for MOOS-3 stability: nitrate (n = 14) w = 0.1535 μg/g, uw = 0.020 μg/g
At this point the MOOS-3 was not a reliable CRM for validation of nitrite. Comparison of the data with the independent UV-vis method is a better alternative for validation
n = 3 as the isotope ratio was measured in triplicate
Comparability tests are for information only
MOOS-3 is expired and not properly informative as a QC sample at this point
4. Determination of silicate by UV-vis w2d2
[5 empty] | Baseline | Blank | Grav | Density | Matrix | uncertainty | Linear model: Gauss's method (GUM's Linear Approximation) |  |  |  | Monte Carlo method |  |  | Ratio Gauss vs Monte Carlo |  | Quadratic | Quadratic | Ratio linear
Silicate signals |  |  |  |  | signal | signal | signal | signal | signal | signal | w(SiO2) | u[w(SiO2)] | uR[w(SiO2)] | R^2 | w(SiO2) | u[w(SiO2)] | uR[w(SiO2)] | w(SiO2) | u[w(SiO2)] | w1(SiO2) | w2(SiO2) | vs quadratic | Z-score | Degree of
Sample ID | Date | 810nm (Abs) | mS | mF | A(810) | A(810) | A(810) | A(810) | A(810) | A(810) | µg/g | µg/g | % |  | µg/g | µg/g | % | % | % | µg/g | µg/g | % | value | matching
BLS-DIW | 3/6/2024 1:16:24 PM | 1.3e-05
BSL-LNWS | 3/6/2024 1:17:07 PM | -0.000185
BSL-M4 | 3/6/2024 1:17:47 PM | -0.000135
BSL-SALT | 3/6/2024 1:18:28 PM | 0.00041
BSL-M3 | 3/6/2024 1:19:12 PM | -0.000192
BLK-DIW | 3/6/2024 1:20:00 PM | 0.000529 |  |  | 0.000516
BLK-LNSW | 3/6/2024 1:20:56 PM | 0.012266 |  |  | 0.012451
CAL1 | 3/6/2024 1:21:48 PM | 0.076466 | 3.0717 | 3.3937 | 0.076651 | 0.064716 | 0.071499 | 0.069809 | 0.069809 | 0.00039968 | 0.21593 | 0.00013933
CAL2 | 3/6/2024 1:22:41 PM | 0.10976 | 3.065 | 3.3879 | 0.10995 | 0.098012 | 0.10834 | 0.10578 | 0.10578 | 0.00060562 | 0.32785 | 0.0001638
CAL3 | 3/6/2024 1:23:31 PM | 0.14093 | 3.0721 | 3.3952 | 0.14112 | 0.12918 | 0.14277 | 0.13939 | 0.13939 | 0.00079807 | 0.433 | 0.00018938
CAL4 | 3/6/2024 1:24:25 PM | 0.20633 | 3.0695 | 3.392 | 0.20652 | 0.19458 | 0.21503 | 0.20995 | 0.20995 | 0.001202 | 0.65023 | 0.00025116
CAL5 | 3/6/2024 1:25:14 PM | 0.26781 | 3.0736 | 3.3961 | 0.268 | 0.25606 | 0.28293 | 0.27624 | 0.27624 | 0.0015816 | 0.85791 | 0.00031453
CAL6 | 3/6/2024 1:38:06 PM | 0.55433 | 3.0633 | 3.386 | 0.55452 | 0.54258 | 0.59974 | 0.58556 | 0.58556 | 0.0033525 | 1.8276 | 0.00062995
052a_1 | 3/6/2024 1:26:28 PM | 0.12388 | 3.0577 | 3.3806 | 0.12402 | 0.12402 | 0.13711 | 0.1343 | 0.1343 | 0.00076888 | 0.4164 | 0.0024244 | 0.5822% | 0.99999 | 0.41636 | 0.0026318 | 0.6321% | 100% | 92.12% | 0.41624 | 241.6907 | 100% | -0.58424 | 96.04%
130a_1 | 3/6/2024 1:27:11 PM | 0.12398 | 3.0636 | 3.3852 | 0.12412 | 0.12412 | 0.13714 | 0.13432 | 0.13432 | 0.00076904 | 0.41649 | 0.0024249 | 0.5822% | 0.99999 | 0.41647 | 0.0026149 | 0.6279% | 100% | 92.73% | 0.41633 | 241.6907 | 100% | -0.54868 | 96.06%
207a_1 | 3/6/2024 1:28:49 PM | 0.12454 | 3.0618 | 3.3846 | 0.12467 | 0.12467 | 0.13782 | 0.13499 | 0.13499 | 0.00077285 | 0.41856 | 0.0024364 | 0.5821% | 0.99999 | 0.41859 | 0.0026073 | 0.6229% | 99.99% | 93.45% | 0.41839 | 241.6886 | 100% | 0.30241 | 96.54%
284a_1 | 3/6/2024 1:29:29 PM | 0.12406 | 3.067 | 3.3904 | 0.12419 | 0.12419 | 0.13729 | 0.13447 | 0.13447 | 0.00076987 | 0.41694 | 0.0024274 | 0.5822% | 0.99999 | 0.41703 | 0.0026467 | 0.6347% | 99.98% | 91.71% | 0.41678 | 241.6902 | 100% | -0.36354 | 96.16%
360a_1 | 3/6/2024 1:30:11 PM | 0.12601 | 3.068 | 3.3908 | 0.12615 | 0.12615 | 0.13941 | 0.13655 | 0.13655 | 0.00078177 | 0.4234 | 0.0024633 | 0.5818% | 0.99999 | 0.42348 | 0.0026653 | 0.6294% | 99.98% | 92.42% | 0.42323 | 241.6838 | 100% | 2.2948 | 97.65%
052a_2 | 3/6/2024 1:30:55 PM | 0.12447 | 3.0578 | 3.3809 | 0.1246 | 0.1246 | 0.13777 | 0.13494 | 0.13494 | 0.00077257 | 0.4184 | 0.0024355 | 0.5821% | 0.99999 | 0.41847 | 0.0026888 | 0.6425% | 99.98% | 90.58% | 0.41824 | 241.6887 | 100% | 0.23839 | 96.5%
130a_2 | 3/6/2024 1:31:41 PM | 0.12427 | 3.0716 | 3.394 | 0.12441 | 0.12441 | 0.13746 | 0.13464 | 0.13464 | 0.00077085 | 0.41747 | 0.0024304 | 0.5822% | 0.99999 | 0.41748 | 0.0026588 | 0.6369% | 100% | 91.41% | 0.4173 | 241.6897 | 100% | -0.1454 | 96.29%
207a_2 | 3/6/2024 1:32:33 PM | 0.12335 | 3.0678 | 3.3891 | 0.12349 | 0.12349 | 0.13642 | 0.13362 | 0.13362 | 0.000765 | 0.4143 | 0.0024128 | 0.5824% | 0.99999 | 0.41429 | 0.002577 | 0.622% | 100% | 93.63% | 0.41414 | 241.6928 | 100% | -1.4513 | 95.56%
284a_2 | 3/6/2024 1:33:24 PM | 0.12411 | 3.0662 | 3.3888 | 0.12425 | 0.12425 | 0.13732 | 0.1345 | 0.1345 | 0.00077004 | 0.41703 | 0.0024279 | 0.5822% | 0.99999 | 0.4171 | 0.0026802 | 0.6426% | 99.98% | 90.59% | 0.41687 | 241.6901 | 100% | -0.32512 | 96.19%
360a_2 | 3/6/2024 1:34:16 PM | 0.1248 | 3.0738 | 3.3963 | 0.12494 | 0.12494 | 0.13805 | 0.13521 | 0.13521 | 0.00077411 | 0.41924 | 0.0024402 | 0.582% | 0.99999 | 0.41932 | 0.0026388 | 0.6293% | 99.98% | 92.47% | 0.41907 | 241.6879 | 100% | 0.58268 | 96.69%
SALT1_309 | 3/6/2024 1:35:30 PM | 0.15579 | 3.071 | 3.3936 | 0.15538 | 0.15538 | 0.17171 | 0.16782 | 0.16782 | 0.00096079 | 0.52053 | 0.0030107 | 0.5784% | 0.99999 | 0.52057 | 0.0033711 | 0.6476% | 99.99% | 89.31% | 0.52031 | 241.5867 | 100%
MOOS3 | 3/6/2024 1:36:25 PM | 0.515 | 3.0834 | 3.4062 | 0.51519 | 0.51519 | 0.56912 | 0.55217 | 0.55217 | 0.0031614 | 1.7221 | 0.010006 | 0.581% | 0.99999 | 1.7221 | 0.012843 | 0.7457% | 100% | 77.91% | 1.7224 | 172.6678 | 99.98%
| CRM |  | Measurement |  |  | Comparability |  |  |  |  | Ratio
Sample | μ | unc | x | s | n | IRMM | ISO | Wald | t-test | Mood, B=2 | x/μ
SALT1_309 | 0.51957 | 0.0083042 | 0.52053 | 0.0030107 | 3 | TRUE | TRUE | TRUE | TRUE | TRUE | 1.0018
MOOS3 | 1.77 | 0.02 | 1.7221 | 0.010006 | 3 | FALSE | TRUE | TRUE | FALSE | TRUE | 0.97291
Note ICM-21-02 for MOOS-3 stability: silicate (n = 7) w = 1.747 μg/g, uw = 0.010 μg/g
n = 3 as the isotope ratio was measured in triplicate
Comparability tests are for information only
MOOS-3 is expired and not properly informative as a QC sample at this point
EtONO and EtONO2 signals on MOOS-4 without derivatization
File | Date | Nitrite 31 | Nitrate 46
MOOS4_NO_DERV.D | 2024-03-06 10:47 | 0 | N/A
MOOS4_NO_DERV.D | 2024-03-05 12:08 | N/A | 0
No endogenous EtONO or EtONO2 were detected in the underivatized MOOS-4 sample
Difference between isotopic composition of sample (RA) and standard (RA*)
1. Isotope ratio Nitrate
File | Date | rAB (nitrate) |  | Sample | rAB/(V/V) |  | t-Test: Paired Two Sample for Means
ST_1.D | 2024-03-05 23:11 | 79.7213 |  | MOOS4_1 | 78.6152
MOOS4_1.D | 2024-03-05 23:21 | 78.6152 |  | MOOS4_2 | 81.3355 |  |  | Variable 1 | Variable 2
ST_2.D | 2024-03-05 23:31 | 82.0991 |  | MOOS4_3 | 82.6065 |  | Mean | 81.554 | 81.7959
MOOS4_2.D | 2024-03-05 23:41 | 81.3355 |  | MOOS4_4 | 81.9053 |  | Variance | 2.0775 | 0.94569
ST_3.D | 2024-03-05 23:51 | 82.2614 |  | MOOS4_5 | 83.0929 |  | Observations | 7 | 7
MOOS4_3.D | 2024-03-06 00:01 | 82.6065 |  | MOOS4_6 | 81.9428 |  | Pearson Correlation | 0.84515
ST_4.D | 2024-03-06 00:11 | 82.67 |  | MOOS4_7 | 81.3797 |  | Hypothesized Mean Difference | 0
MOOS4_4.D | 2024-03-06 00:21 | 81.9053 |  | NIST_1 | 79.7213 |  | df | 6
ST_5.D | 2024-03-06 00:31 | 81.8296 |  | NIST_2 | 82.0991 |  | t Stat | -0.79134
MOOS4_5.D | 2024-03-06 00:41 | 83.0929 |  | NIST_3 | 82.2614 |  | P(T<=t) one-tail | 0.22944
ST_6.D | 2024-03-06 00:51 | 82.3244 |  | NIST_4 | 82.67 |  | t Critical one-tail | 1.9432
MOOS4_6.D | 2024-03-06 01:01 | 81.9428 |  | NIST_5 | 81.8296 |  | P(T<=t) two-tail | 0.45888
ST_7.D | 2024-03-06 01:11 | 81.6653 |  | NIST_6 | 82.3244 |  | t Critical two-tail | 2.4469
MOOS4_7.D | 2024-03-06 01:21 | 81.3797 |  | NIST_7 | 81.6653 |  | No significantive difference alfa = 0.05
2. Isotope ratio Nitrite
File | Date | rAB (nitrate) |  | Sample | rAB/(V/V) |  | Sample | rAB/(V/V) |  | t-Test: Paired Two Sample for Means
ST_1.D | 2024-03-07 02:10 | 159.069 |  | MOOS4_1 | 168.4723 |  | MOOS4_1 | 168.4723
MOOS4_1.D | 2024-03-07 02:24 | 168.4723 |  | MOOS4_2 | 155.014 |  | MOOS4_2 | 155.014 |  |  | Variable 1 | Variable 2
ST_2.D | 2024-03-07 02:37 | 162.906 |  | MOOS4_3 | 136.9964 |  | MOOS4_4 | 159.5065 |  | Mean | 162.7396 | 165.938
MOOS4_2.D | 2024-03-07 02:51 | 155.014 |  | MOOS4_4 | 159.5065 |  | MOOS4_5 | 159.2242 |  | Variance | 33.0054 | 42.6442
ST_3.D | 2024-03-07 03:05 | 170.1123 |  | MOOS4_5 | 159.2242 |  | MOOS4_6 | 164.7165 |  | Observations | 6 | 6
MOOS4_3.D | 2024-03-07 03:19 | 136.9964 |  | MOOS4_6 | 164.7165 |  | MOOS4_7 | 169.5041 |  | Pearson Correlation | -0.46729
ST_4.D | 2024-03-07 03:33 | 177.0429 |  | MOOS4_7 | 169.5041 |  | NIST_1 | 159.069 |  | Hypothesized Mean Difference | 0
MOOS4_4.D | 2024-03-07 03:46 | 159.5065 |  | NIST_1 | 159.069 |  | NIST_2 | 162.906 |  | df | 5
ST_5.D | 2024-03-07 04:00 | 162.3431 |  | NIST_2 | 162.906 |  | NIST_3 | 170.1123 |  | t Stat | -0.74459
MOOS4_5.D | 2024-03-07 04:14 | 159.2242 |  | NIST_3 | 170.1123 |  | NIST_4 | 177.0429 |  | P(T<=t) one-tail | 0.24501
ST_6.D | 2024-03-07 04:28 | 164.1551 |  | NIST_4 | 177.0429 |  | NIST_5 | 162.3431 |  | t Critical one-tail | 2.015
MOOS4_6.D | 2024-03-07 04:41 | 164.7165 |  | NIST_5 | 162.3431 |  | NIST_6 | 164.1551 |  | P(T<=t) two-tail | 0.49001
ST_7.D | 2024-03-07 04:55 | 124.3796 |  | NIST_6 | 164.1551 |  |  |  |  | t Critical two-tail | 2.5706
MOOS4_7.D | 2024-03-07 05:09 | 169.5041 |  | NIST_7 | 124.3796 |  |  |  |  | No significantive difference alfa = 0.05
2024-03-21 Characterization MOOS-4 [Silicate LC-ICPMS Zuzana]: sample preparation (verified 2024-04-24)
Preparation of standard solutions silicate in DIW
| Origin |  | m0 = vial |  | m1 = m0 + standard |  | m2 = m1 + DIW | [7 empty] | Final
ID solution | w(SiO2) | u[w(SiO2)] | load | unload | load | unload | load | unload | m0 | m1 | m2 | u(m0) | u(m1) | u(m2) | w(SiO2) | u[w(SiO2)] | Target | Traciability | Prep. Date | Exp. Date | t | Rel Hum | Analyst | Balance
unique | µg/g | µg/g | g | g | g | g | g | g | g | g | g | g | g | g | µg/g | µg/g | % |  | yyyymmdd | yyyymmdd | °C | % | initials | model
Dilution 1 | 21432.0637 | 0 | 12.8586 |  | 13.1694 |  | 53.1369 |  | 12.8586 | 13.1694 | 53.1369 | 0.0002 | 0.0002 | 0.0002 | 165.3765 | 0.1505 | N/A | NIST | 20240315 | 20240330 | 21.7 | 28.3 | ZG | XPE304
| Origin |  | m0 = vial |  | m1 = m0 + standard |  | m2 = m1 + DIW | [7 empty] | Final
ID solution | w(SiO2) | u[w(SiO2)] | load | unload | load | unload | load | unload | m0 | m1 | m2 | u(m0) | u(m1) | u(m2) | w(SiO2) | u[w(SiO2)] | Target | Traciability | Prep. Date | Exp. Date | t | Rel Hum | Analyst | Balance
unique | µg/g | µg/g | g | g | g | g | g | g | g | g | g | g | g | g | µg/g | µg/g | % |  | yyyymmdd | yyyymmdd | °C | % | initials | model
Dilution 2 | 165.3765 | 0.1505 | 12.82 |  | 13.0764 |  | 52.688 |  | 12.82 | 13.0764 | 52.688 | 0.0002 | 0.0002 | 0.0002 | 1.0636 | 0.001521 | N/A | NIST | 20240315 | 20240330 | 21.7 | 28.3 | ZG | XPE304
| Origin |  | m0 = vial |  | m1 = m0 + standard |  | m2 = m1 + DIW | [7 empty] | Final
ID solution | w(SiO2) | u[w(SiO2)] | load | unload | load | unload | load | unload | m0 | m1 | m2 | u(m0) | u(m1) | u(m2) | w(SiO2) | u[w(SiO2)] | Target | Traciability | Prep. Date | Exp. Date | t | Rel Hum | Analyst | Balance
unique | µg/g | µg/g | g | g | g | g | g | g | g | g | g | g | g | g | µg/g | µg/g | % |  | yyyymmdd | yyyymmdd | °C | % | initials | model
AB1 | 1.0636 | 0.001521 | 6.1597 |  | 8.0966 |  | 16.027 |  | 6.1597 | 8.0966 | 16.027 | 0.0002 | 0.0002 | 0.0002 | 0.20877 | 0.00030018 | N/A | NIST | 20240315 | 20240330 | 21.7 | 28.3 | ZG | XPE304
AB2 | 1.0636 | 0.001521 | 6.1556 |  | 10.0308 |  | 15.9111 |  | 6.1556 | 10.0308 | 15.9111 | 0.0002 | 0.0002 | 0.0002 | 0.42249 | 0.00060511 | N/A | NIST | 20240315 | 20240330 | 21.7 | 28.3 | ZG | XPE304
AB3 | 1.0636 | 0.001521 | 6.1841 |  | 13.9298 |  | 16.2397 |  | 6.1841 | 13.9298 | 16.2397 | 0.0002 | 0.0002 | 0.0002 | 0.81926 | 0.0011722 | N/A | NIST | 20240315 | 20240330 | 21.7 | 28.3 | ZG | XPE304
Preparation of isotope dilution blends for SiO2 analysis by LC ICPMS
|  | m0 = mass of vial |  |  | m1 = m0 + mass of sample/ST |  |  | m2 = m1 + IntST
|  | load | unload | mass | load | unload | mass | load | unload | mass |  | mA | mB
Type | ID | g | g | g | g | g | g | g | g | g | Date | g | g
CAL | AB1 | 0.9306 |  | 0.9306 | 1.4262 |  | 1.4262 | 1.8822 |  | 1.8822 | 2024-03-21 | 0.4956 | 0.456
CAL | AB2A | 0.9359 |  | 0.9359 | 1.4314 |  | 1.4314 | 1.8875 |  | 1.8875 |  | 0.4955 | 0.4561
CAL | AB2B | 0.9502 |  | 0.9502 | 1.4472 |  | 1.4472 | 1.9044 |  | 1.9044 |  | 0.497 | 0.4572
CAL | AB2C | 0.9505 |  | 0.9505 | 1.4485 |  | 1.4485 | 1.9061 |  | 1.9061 |  | 0.498 | 0.4576
CAL | AB3 | 0.9301 |  | 0.9301 | 1.4255 |  | 1.4255 | 1.8828 |  | 1.8828 |  | 0.4954 | 0.4573
SAMPLE | 181_A1 | 0.9362 |  | 0.9362 | 1.4508 |  | 1.4508 | 1.9064 |  | 1.9064 |  | 0.5146 | 0.4556
SAMPLE | 181_A2 | 0.9309 |  | 0.9309 | 1.4221 |  | 1.4221 | 1.8791 |  | 1.8791 |  | 0.4912 | 0.457
SAMPLE | 335_A1 | 0.9302 |  | 0.9302 | 1.442 |  | 1.442 | 1.899 |  | 1.899 |  | 0.5118 | 0.457
SAMPLE | 335_A2 | 0.9418 |  | 0.9418 | 1.4329 |  | 1.4329 | 1.891 |  | 1.891 |  | 0.4911 | 0.4581
SAMPLE | 027_A2 | 0.9311 |  | 0.9311 | 1.4226 |  | 1.4226 | 1.8809 |  | 1.8809 |  | 0.4915 | 0.4583
QC | SALT1_309_1 | 0.936 |  | 0.936 | 1.4434 |  | 1.4434 | 2.036 |  | 2.036 |  | 0.5074 | 0.59258
QC | SALT1_309_2 | 0.9309 |  | 0.9309 | 1.44 |  | 1.44 | 2.0345 |  | 2.0345 | 2024-03-21 | 0.5091 | 0.59453
QC m2 recalculated assuming the same internal standard solution used for MOOS-4 samples (MOOS4_SEAN1\CERT\DATA\20240320_Silicate_LCICPMS_Zuzana)
2024-03-21 Characterization MOOS-4 [Silicate LC-ICPMS Zuzana]: data analysis (verified 2024-05-10)
1. Determination of silicate by IDMS (LC-ICPMS Zuzana)
Sample ID | rA*B | u(rA*B) | Date | wA* | u(wA*) | mA* | mB | wA*(mA*/mB) | u(wA*(mA*/mB))
AB1 | 0.49883 | 0.0072166 | 2024-03-21 | 0.20877 | 0.00030018 | 0.4956 | 0.456 | 0.2269 | 0.00037817
AB2A | 0.98138 | 0.014198 | 2024-03-21 | 0.42249 | 0.00060511 | 0.4955 | 0.4561 | 0.45898 | 0.00076276
AB2B | 0.98352 | 0.014228 | 2024-03-21 | 0.42249 | 0.00060511 | 0.497 | 0.4572 | 0.45926 | 0.0007627
AB2C | 0.98176 | 0.014203 | 2024-03-21 | 0.42249 | 0.00060511 | 0.498 | 0.4576 | 0.45979 | 0.0007633
AB3 | 1.8255 | 0.02641 | 2024-03-21 | 0.81926 | 0.0011722 | 0.4954 | 0.4573 | 0.88751 | 0.0014733
g: Gauss's method (GUM's Linear Approximation); g: Monte Carlo method |  | Ratio wA | Ratio u(wA) | Z-score | Degree of
Sample ID | rAB | u(rAB) | Date | mA | mB | m = mB/mA | u(m) | wA | u(wA) | uR(wA) | R^2 | wA | u(wA) | uR(wA) | GS/MT | GS/MT | value | matching
181_A1 | 1.003 | 0.014511 | 2024-03-21 | 0.5146 | 0.4556 | 0.88535 | 0.0007341 | 0.41572 | 0.0063296 | 1.523% | 1 | 0.41572 | 0.0073112 | 1.759% | 100% | 86.57% | -0.82184 | 102.1%
181_A2 | 0.9695 | 0.014026 | 2024-03-21 | 0.4912 | 0.457 | 0.93037 | 0.00078649 | 0.42157 | 0.0064122 | 1.521% | 1 | 0.42195 | 0.0071418 | 1.693% | 99.91% | 89.78% | 0.090915 | 98.71%
335_A1 | 1.0058 | 0.01455 | 2024-03-21 | 0.5118 | 0.457 | 0.89293 | 0.0007409 | 0.42048 | 0.0064026 | 1.523% | 1 | 0.42056 | 0.0073456 | 1.747% | 99.98% | 87.16% | -0.078884 | 102.4%
335_A2 | 0.99173 | 0.014347 | 2024-03-21 | 0.4911 | 0.4581 | 0.9328 | 0.00078761 | 0.43283 | 0.0065881 | 1.522% | 1 | 0.43339 | 0.0077242 | 1.782% | 99.87% | 85.29% | 1.8503 | 101%
027_A2 | 0.95157 | 0.013766 | 2024-03-21 | 0.4915 | 0.4583 | 0.93245 | 0.00078683 | 0.41432 | 0.0062985 | 1.52% | 1 | 0.41491 | 0.0071041 | 1.712% | 99.86% | 88.66% | -1.0405 | 96.88%
SALT1_309_1 | 0.93493 | 0.013526 | 2024-03-21 | 0.5074 | 0.59258 | 1.1679 | 0.00085706 | 0.50944 | 0.0077376 | 1.519% | 1 | 0.50989 | 0.0086247 | 1.691% | 99.91% | 89.71%
SALT1_309_2 | 0.9419 | 0.013626 | 2024-03-21 | 0.5091 | 0.59453 | 1.1678 | 0.00085417 | 0.51338 | 0.0077991 | 1.519% | 1 | 0.51346 | 0.0089685 | 1.747% | 99.99% | 86.96%
| CRM |  | Measurement |  |  | Comparability |  |  |  |  | Ratio
Sample | μ | unc | x | s | n | IRMM | ISO | Wald | t-test | Mood, B=2 | x/μ
SALT1_309_1 | 0.51957 | 0.0083042 | 0.50944 | 0.0077376 | 3 | TRUE | TRUE | TRUE | TRUE | TRUE | 0.98049
SALT1_309_2 | 0.51957 | 0.0083042 | 0.51338 | 0.0077991 | 3 | TRUE | TRUE | TRUE | TRUE | TRUE | 0.98809
Comparability tests are for information only
MOOS-3 is expired and not properly informative as a QC sample at this point
CHAR summary nitrite (verified 2024-05-08)
Linear pooling: combining data single day and single method
Sample | w(NO2−) | uw(NO2−) | Cal/Method | Uncertainty | Date |  | w(NO2−) | uw(NO2−) |  | w(NO2−) | uw(NO2−) | uRw(NO2−) | average | stdev | RSD | Cal/Method | Date
ID | μg/g | μg/g | ID | ID | YYYYMMDD |  | μg/g | μg/g |  | μg/g | μg/g | % | μg/g | μg/g | % | ID | YYYYMMDD
001_A1 | 0.07595 | 0.00023596 | ID/GCMS | Gauss | 2024-02-21 13:14 |  | 0.07595 | 0.00023596 |  | 0.075606 | 0.00031851 | 0.4213% | 0.075604 | 0.00022618 | 0.2992% | ID/GCMS 2024-02-21 n = 10 | 2024-02-21
078_A1 | 0.075767 | 0.00023547 | ID/GCMS | Gauss | 2024-02-21 13:28 |  | 0.075767 | 0.00023547 |  | 0.075646 | 0.00064255 | 0.8494% | 0.075648 | 0.00034709 | 0.4588% | EX/UV-VIS 2024-02-20 n = 10 | 2024-02-20
155_A1 | 0.075282 | 0.00023421 | ID/GCMS | Gauss | 2024-02-21 13:42 |  | 0.075282 | 0.00023421 |  | 0.075615 | 0.00046709 | 0.6177% | 0.075615 | 0.00011659 | 0.1542% | EX/UV-VIS 2024-02-21 n = 5 | 2024-02-21
232_A1 | 0.075755 | 0.00023544 | ID/GCMS | Gauss | 2024-02-21 13:56 |  | 0.075755 | 0.00023545 |  | 0.075416 | 0.00026394 | 0.35% | 0.075417 | 0.00012003 | 0.1591% | ID/GCMS 2024-03-06 n = 10 | 2024-03-06
310_A1 | 0.075665 | 0.00023527 | ID/GCMS | Gauss | 2024-02-21 14:09 |  | 0.075665 | 0.00023527 |  | *** Data not linked. Calculation performed in R and crossed-validated with https://consensus.nist.gov/app/nicob***
001_A2 | 0.075394 | 0.00023443 | ID/GCMS | Gauss | 2024-02-21 14:23 |  | 0.075394 | 0.00023443
078_A2 | 0.075492 | 0.00023475 | ID/GCMS | Gauss | 2024-02-21 14:37 |  | 0.075492 | 0.00023475 |  | DerSimonian-Laird: combining data different day and/or method
155_A2 | 0.07585 | 0.00023581 | ID/GCMS | Gauss | 2024-02-21 14:51 |  | 0.07585 | 0.00023581 |  | $mu | $s | $tau | w(NO2−) | uw(NO2−)CHAR | uRw(NO2−)CHAR | uw/SD | w/average
232_A2 | 0.075357 | 0.00023446 | ID/GCMS | Gauss | 2024-02-21 15:05 |  | 0.075357 | 0.00023446 |  | μg/g | μg/g | μg/g | μg/g | μg/g | % | % | %
310_A2 | 0.07553 | 0.00023492 | ID/GCMS | Gauss | 2024-02-21 15:19 |  | 0.07553 | 0.00023492 |  | 0.075523 | 0.00017898 | 0 | 0.075523 | 0.00017898 | 0.237% | 72.99% | 99.94%
001a_1 | 0.076063 | 0.00055289 | EX/UV-VIS | Gauss | 2024-02-20 |  | 0.076063 | 0.00055289 |  | *** Data not linked. Calculation performed in R and crossed-validated with https://consensus.nist.gov/app/nicob***
078a_1 | 0.075672 | 0.00055258 | EX/UV-VIS | Gauss | 2024-02-20 |  | 0.075672 | 0.00055258 |  | *** Uncertainty from primary standard was not considered here, but enrolled in CERT tab ***
155a_1 | 0.075506 | 0.00055246 | EX/UV-VIS | Gauss | 2024-02-20 |  | 0.075506 | 0.00055246
232a_1 | 0.075444 | 0.00055241 | EX/UV-VIS | Gauss | 2024-02-20 |  | 0.075444 | 0.00055241
310a_1 | 0.076249 | 0.00055305 | EX/UV-VIS | Gauss | 2024-02-20 |  | 0.076249 | 0.00055305
001a_2 | 0.075382 | 0.00055237 | EX/UV-VIS | Gauss | 2024-02-20 |  | 0.075382 | 0.00055237
078a_2 | 0.076036 | 0.00055287 | EX/UV-VIS | Gauss | 2024-02-20 |  | 0.076036 | 0.00055287
155a_2 | 0.075542 | 0.00055248 | EX/UV-VIS | Gauss | 2024-02-20 |  | 0.075542 | 0.00055248
232a_2 | 0.075327 | 0.00055233 | EX/UV-VIS | Gauss | 2024-02-20 |  | 0.075327 | 0.00055233
310a_2 | 0.075257 | 0.00055228 | EX/UV-VIS | Gauss | 2024-02-20 |  | 0.075257 | 0.00055228
001a_1 | 0.075665 | 0.00045504 | EX/UV-VIS | Gauss | 2024-02-21 |  | 0.075665 | 0.00045504
078a_1 | 0.075666 | 0.00045504 | EX/UV-VIS | Gauss | 2024-02-21 |  | 0.075666 | 0.00045504
155a_1 | 0.075698 | 0.00045511 | EX/UV-VIS | Gauss | 2024-02-21 |  | 0.075698 | 0.00045511
232a_1 | 0.075639 | 0.00045498 | EX/UV-VIS | Gauss | 2024-02-21 |  | 0.075639 | 0.00045498
310a_1 | 0.07541 | 0.00045449 | EX/UV-VIS | Gauss | 2024-02-21 |  | 0.07541 | 0.00045449
052_A1 | 0.075441 | 0.00023785 | ID/GCMS | Gauss | 2024-03-06 12:54 |  | 0.075441 | 0.00023785
130_A1 | 0.075416 | 0.00023783 | ID/GCMS | Gauss | 2024-03-06 13:08 |  | 0.075416 | 0.00023783
207_A1 | 0.075319 | 0.00023757 | ID/GCMS | Gauss | 2024-03-06 13:23 |  | 0.075319 | 0.00023757
284_A1 | 0.075677 | 0.00023865 | ID/GCMS | Gauss | 2024-03-06 13:37 |  | 0.075677 | 0.00023865
360_A1 | 0.075316 | 0.00023754 | ID/GCMS | Gauss | 2024-03-06 13:51 |  | 0.075316 | 0.00023754
052_A2 | 0.075486 | 0.00023806 | ID/GCMS | Gauss | 2024-03-06 14:05 |  | 0.075486 | 0.00023806
130_A2 | 0.075486 | 0.00023809 | ID/GCMS | Gauss | 2024-03-06 14:19 |  | 0.075486 | 0.00023809
207_A2 | 0.075318 | 0.00023757 | ID/GCMS | Gauss | 2024-03-06 14:34 |  | 0.075318 | 0.00023757
284_A2 | 0.07527 | 0.00023746 | ID/GCMS | Gauss | 2024-03-06 14:48 |  | 0.07527 | 0.00023746
360_A2 | 0.075441 | 0.00023793 | ID/GCMS | Gauss | 2024-03-06 15:02 |  | 0.075441 | 0.00023793
CHAR summary nitrate (verified 2024-05-08)
Linear pooling: combining data single day and single method
Sample | w(NO3−) | uw(NO3−) | Cal/Method | Uncertainty | Date |  | w(NO3−) | uw(NO3−) |  | w(NO3−) | uw(NO3−) | uRw(NO3−) | average | stdev | RSD | Cal/Method | Date
ID | μg/g | μg/g | ID | ID | YYYYMMDD |  | μg/g | μg/g |  | μg/g | μg/g | % | μg/g | μg/g | % | ID | YYYYMMDD
078_A1 | 1.0412 | 0.017134 | ID/GCMS | Gauss | 2024-02-20 14:38 |  | 1.0412 | 0.017134 |  | 1.0321 | 0.017639 | 1.709% | 1.032 | 0.0048088 | 0.466% | ID/GCMS 2024-02-20 n = 8 | 2024-02-20
155_A1 | 1.0305 | 0.016954 | ID/GCMS | Gauss | 2024-02-20 14:48 |  | 1.0305 | 0.016954 |  | 1.0519 | 0.019671 | 1.87% | 1.052 | 0.012044 | 1.145% | ID/GCMS 2024-03-05 n = 10 | 2024-03-05
232_A1 | 1.0276 | 0.016909 | ID/GCMS | Gauss | 2024-02-20 14:58 |  | 1.0276 | 0.016909 |  | *** Data not linked. Calculation performed in R and crossed-validated with https://consensus.nist.gov/app/nicob***
310_A1 | 1.0296 | 0.016949 | ID/GCMS | Gauss | 2024-02-20 15:08 |  | 1.0296 | 0.016949
001_A2 | 1.0316 | 0.016987 | ID/GCMS | Gauss | 2024-02-20 15:18 |  | 1.0316 | 0.016987 |  | DerSimonian-Laird: combining data different day and/or method
078_A2 | 1.0271 | 0.016897 | ID/GCMS | Gauss | 2024-02-20 15:28 |  | 1.0271 | 0.016897 |  | $mu | $s | $tau | w(NO3−) | uw(NO3−)CHAR | uRw(NO3−)CHAR | uw/SD | w/average
155_A2 | 1.0371 | 0.017064 | ID/GCMS | Gauss | 2024-02-20 15:38 |  | 1.0371 | 0.017064 |  | μg/g | μg/g | μg/g | μg/g | μg/g | % | % | %
232_A2 | 1.0314 | 0.016968 | ID/GCMS | Gauss | 2024-02-20 15:48 |  | 1.0314 | 0.016968 |  | 1.0409 | 0.013133 | 0 | 1.0409 | 0.013133 | 1.262% | 95% | 99.79%
052_A1 | 1.056 | 0.016089 | ID/GCMS | Gauss | 2024-03-05 13:38 |  | 1.056 | 0.016089 |  | *** Data not linked. Calculation performed in R and crossed-validated with https://consensus.nist.gov/app/nicob***
130_A1 | 1.0435 | 0.015908 | ID/GCMS | Gauss | 2024-03-05 13:48 |  | 1.0435 | 0.015908 |  | *** Uncertainty from primary standard was not considered here, but enrolled in CERT tab ***
207_A1 | 1.0406 | 0.015871 | ID/GCMS | Gauss | 2024-03-05 13:58 |  | 1.0406 | 0.015871
284_A1 | 1.0576 | 0.016119 | ID/GCMS | Gauss | 2024-03-05 14:08 |  | 1.0576 | 0.016119
360_A1 | 1.0435 | 0.015911 | ID/GCMS | Gauss | 2024-03-05 14:18 |  | 1.0435 | 0.015911
052_A2 | 1.0665 | 0.016249 | ID/GCMS | Gauss | 2024-03-05 14:28 |  | 1.0665 | 0.016249
130_A2 | 1.0752 | 0.016379 | ID/GCMS | Gauss | 2024-03-05 14:38 |  | 1.0752 | 0.016379
207_A2 | 1.0546 | 0.016073 | ID/GCMS | Gauss | 2024-03-05 14:49 |  | 1.0546 | 0.016073
284_A2 | 1.0422 | 0.015895 | ID/GCMS | Gauss | 2024-03-05 14:59 |  | 1.0422 | 0.015895
360_A2 | 1.0406 | 0.015869 | ID/GCMS | Gauss | 2024-03-05 15:09 |  | 1.0406 | 0.015869
CHAR summary phosphate (verified 2024-05-08)
Linear pooling: combining data single day and single method
Sample | w(PO43−) | uw(PO43−) | Cal/Method | Uncertainty | Date |  | w(PO43−) | uw(PO43−) |  | w(PO43−) | uw(PO43−) | uRw(PO43−) | average | stdev | RSD | Cal/Method | Date
ID | μg/g | μg/g | ID | ID | YYYYMMDD |  | μg/g | μg/g |  | μg/g | μg/g | % | μg/g | μg/g | % | ID | YYYYMMDD
001a_1 | 0.26245 | 0.0011008 | EX/UV-VIS | Gauss | 2024-02-20 |  | 0.26245 | 0.0011008 |  | 0.26318 | 0.0014695 | 0.5584% | 0.26317 | 0.0010251 | 0.3895% | EX/UV-VIS 2024-02-20 n = 10 | 2024-02-20
078a_1 | 0.2641 | 0.0011064 | EX/UV-VIS | Gauss | 2024-02-20 |  | 0.2641 | 0.0011064 |  | 0.26396 | 0.0010827 | 0.4102% | 0.26396 | 0.00076644 | 0.2904% | EX/UV-VIS 2024-03-05 n = 10 | 2024-03-05
155a_1 | 0.2634 | 0.001104 | EX/UV-VIS | Gauss | 2024-02-20 |  | 0.2634 | 0.001104 |  | *** Data not linked. Calculation performed in R and crossed-validated with https://consensus.nist.gov/app/nicob***
232a_1 | 0.26305 | 0.0011028 | EX/UV-VIS | Gauss | 2024-02-20 |  | 0.26305 | 0.0011028
310a_1 | 0.26436 | 0.0011072 | EX/UV-VIS | Gauss | 2024-02-20 |  | 0.26436 | 0.0011072 |  | DerSimonian-Laird: combining data different day and/or method
001a_2 | 0.26321 | 0.0011033 | EX/UV-VIS | Gauss | 2024-02-20 |  | 0.26321 | 0.0011033 |  | $mu | $s | $tau | w(PO43−) | uw(PO43−)CHAR | uRw(PO43−)CHAR | uw/SD | w/average
078a_2 | 0.26397 | 0.0011059 | EX/UV-VIS | Gauss | 2024-02-20 |  | 0.26397 | 0.0011059 |  | μg/g | μg/g | μg/g | μg/g | μg/g | % | % | %
155a_2 | 0.26332 | 0.0011037 | EX/UV-VIS | Gauss | 2024-02-20 |  | 0.26332 | 0.0011037 |  | 0.26369 | 0.00087165 | 0 | 0.26369 | 0.00087165 | 0.3306% | 89.91% | 100%
232a_2 | 0.26311 | 0.001103 | EX/UV-VIS | Gauss | 2024-02-20 |  | 0.26311 | 0.001103 |  | *** Data not linked. Calculation performed in R and crossed-validated with https://consensus.nist.gov/app/nicob***
310a_2 | 0.26074 | 0.001095 | EX/UV-VIS | Gauss | 2024-02-20 |  | 0.26074 | 0.001095 |  | *** Uncertainty from primary standard was not considered here, but enrolled in CERT tab ***
052a_1 | 0.26335 | 0.00080082 | EX/UV-VIS | Gauss | 2024-03-05 13:38 |  | 0.26335 | 0.00080082
130a_1 | 0.26475 | 0.00080454 | EX/UV-VIS | Gauss | 2024-03-05 13:38 |  | 0.26475 | 0.00080454
207a_1 | 0.2637 | 0.00080175 | EX/UV-VIS | Gauss | 2024-03-05 13:38 |  | 0.2637 | 0.00080175
284a_1 | 0.26388 | 0.00080224 | EX/UV-VIS | Gauss | 2024-03-05 13:38 |  | 0.26388 | 0.00080224
360a_1 | 0.26433 | 0.00080343 | EX/UV-VIS | Gauss | 2024-03-05 13:38 |  | 0.26433 | 0.00080343
052a_2 | 0.2635 | 0.00080123 | EX/UV-VIS | Gauss | 2024-03-05 13:38 |  | 0.2635 | 0.00080123
130a_2 | 0.26521 | 0.00080578 | EX/UV-VIS | Gauss | 2024-03-05 13:38 |  | 0.26521 | 0.00080578
207a_2 | 0.26255 | 0.0007987 | EX/UV-VIS | Gauss | 2024-03-05 13:38 |  | 0.26255 | 0.0007987
284a_2 | 0.26383 | 0.00080209 | EX/UV-VIS | Gauss | 2024-03-05 13:38 |  | 0.26383 | 0.00080209
360a_2 | 0.2645 | 0.00080387 | EX/UV-VIS | Gauss | 2024-03-05 13:38 |  | 0.2645 | 0.00080387
CHAR summary silicate (verified 2024-05-08)
Linear pooling: combining data single day and single method
Sample | w(SiO2) | uw(SiO2) | Cal/Method | Uncertainty | Date |  | w(SiO2) | uw(SiO2) |  | w(SiO2) | uw(SiO2) | uRw(SiO2) | average | stdev | RSD | Cal/Method | Date
ID | μg/g | μg/g | ID | ID | YYYYMMDD |  | μg/g | μg/g |  | μg/g | μg/g | % | μg/g | μg/g | % | ID | YYYYMMDD
001a_1 | 0.42144 | 0.002829 | EX/UV-VIS | Gauss | 2024-02-21 |  | 0.42144 | 0.002829 |  | 0.42596 | 0.0038783 | 0.9105% | 0.42597 | 0.0027756 | 0.6516% | EX/UV-VIS 2024-02-21 n = 10 | 2024-02-21
078a_1 | 0.42354 | 0.0028419 | EX/UV-VIS | Gauss | 2024-02-21 |  | 0.42354 | 0.0028419 |  | 0.41783 | 0.003325 | 0.7958% | 0.41782 | 0.0023967 | 0.5736% | EX/UV-VIS 2024-03-06 n = 10 | 2024-03-06
155a_1 | 0.42728 | 0.0028649 | EX/UV-VIS | Gauss | 2024-02-21 |  | 0.42728 | 0.0028649 |  | 0.42102 | 0.0091489 | 2.173% | 0.42099 | 0.0072973 | 1.733% | ID/LC-ICPMS 2024-03-21 n = 5 | 2024-03-21
232a_1 | 0.42484 | 0.0028499 | EX/UV-VIS | Gauss | 2024-02-21 |  | 0.42484 | 0.0028499 |  | *** Data not linked. Calculation performed in R and crossed-validated with https://consensus.nist.gov/app/nicob***
310a_1 | 0.42896 | 0.0028753 | EX/UV-VIS | Gauss | 2024-02-21 |  | 0.42896 | 0.0028753
001a_2 | 0.4237 | 0.0028429 | EX/UV-VIS | Gauss | 2024-02-21 |  | 0.4237 | 0.0028429 |  | DerSimonian-Laird: combining data different day and/or method
078a_2 | 0.42941 | 0.0028781 | EX/UV-VIS | Gauss | 2024-02-21 |  | 0.42941 | 0.0028781 |  | $mu | $s | $tau | w(SiO2) | uw(SiO2)CHAR | uRw(SiO2)CHAR | uw/SD | w/average
155a_2 | 0.42421 | 0.002846 | EX/UV-VIS | Gauss | 2024-02-21 |  | 0.42421 | 0.002846 |  | μg/g | μg/g | μg/g | μg/g | μg/g | % | % | %
232a_2 | 0.42899 | 0.0028755 | EX/UV-VIS | Gauss | 2024-02-21 |  | 0.42899 | 0.0028755 |  | 0.42142 | 0.0028979 | 0.0023875 | 0.42142 | 0.0037547 | 0.891% | 71.12% | 99.93%
310a_2 | 0.4273 | 0.002865 | EX/UV-VIS | Gauss | 2024-02-21 |  | 0.4273 | 0.002865 |  | *** Data not linked. Calculation performed in R and crossed-validated with https://consensus.nist.gov/app/nicob***
052a_1 | 0.4164 | 0.0024244 | EX/UV-VIS | Gauss | 2024-03-06 |  | 0.4164 | 0.0024244 |  | *** Uncertainty from primary standard was not considered here, but enrolled in CERT tab ***
130a_1 | 0.41649 | 0.0024249 | EX/UV-VIS | Gauss | 2024-03-06 |  | 0.41649 | 0.0024249
207a_1 | 0.41856 | 0.0024364 | EX/UV-VIS | Gauss | 2024-03-06 |  | 0.41856 | 0.0024364
284a_1 | 0.41694 | 0.0024274 | EX/UV-VIS | Gauss | 2024-03-06 |  | 0.41694 | 0.0024274
360a_1 | 0.4234 | 0.0024633 | EX/UV-VIS | Gauss | 2024-03-06 |  | 0.4234 | 0.0024633
052a_2 | 0.4184 | 0.0024355 | EX/UV-VIS | Gauss | 2024-03-06 |  | 0.4184 | 0.0024355
130a_2 | 0.41747 | 0.0024304 | EX/UV-VIS | Gauss | 2024-03-06 |  | 0.41747 | 0.0024304
207a_2 | 0.4143 | 0.0024128 | EX/UV-VIS | Gauss | 2024-03-06 |  | 0.4143 | 0.0024128
284a_2 | 0.41703 | 0.0024279 | EX/UV-VIS | Gauss | 2024-03-06 |  | 0.41703 | 0.0024279
360a_2 | 0.41924 | 0.0024402 | EX/UV-VIS | Gauss | 2024-03-06 |  | 0.41924 | 0.0024402
181_A1 | 0.41572 | 0.0063296 | ID/LC-ICPMS | Gauss | 2024-03-21 |  | 0.41572 | 0.0063296
181_A2 | 0.42157 | 0.0064122 | ID/LC-ICPMS | Gauss | 2024-03-21 |  | 0.42157 | 0.0064122
335_A1 | 0.42048 | 0.0064026 | ID/LC-ICPMS | Gauss | 2024-03-21 |  | 0.42048 | 0.0064026
335_A2 | 0.43283 | 0.0065881 | ID/LC-ICPMS | Gauss | 2024-03-21 |  | 0.43283 | 0.0065881
027_A2 | 0.41432 | 0.0062985 | ID/LC-ICPMS | Gauss | 2024-03-21 |  | 0.41432 | 0.0062985
HOM data phosphate (verified 2024-05-09)
DerSimonian-Laird: combining data different day and/or method
Sample | w(PO43−) | uw(PO43−) | Cal/Method | Uncertainty | Date |  | w(PO43−) | uw(PO43−) |  | $mu | $s | $tau | Date
ID | μg/g | μg/g | ID | ID | YYYYMMDD |  | μg/g | μg/g |  | μg/g | μg/g | μg/g | YYYYMMDD
001a_1 | 0.26245 | 0.0011008 | EX/UV-VIS | Gauss | 2024-02-20 |  | 0.26245 | 0.0011008 |  | 0.26317 | 0.00034886 | 0 | 2024-02-20
078a_1 | 0.2641 | 0.0011064 | EX/UV-VIS | Gauss | 2024-02-20 |  | 0.2641 | 0.0011064 |  | 0.26396 | 0.00025375 | 0 | 2024-03-05
155a_1 | 0.2634 | 0.001104 | EX/UV-VIS | Gauss | 2024-02-20 |  | 0.2634 | 0.001104 |  | *** Data not linked. Calculation performed in R
232a_1 | 0.26305 | 0.0011028 | EX/UV-VIS | Gauss | 2024-02-20 |  | 0.26305 | 0.0011028 |  | and crossed-validated with https://consensus.nist.gov/app/nicob***
310a_1 | 0.26436 | 0.0011072 | EX/UV-VIS | Gauss | 2024-02-20 |  | 0.26436 | 0.0011072
001a_2 | 0.26321 | 0.0011033 | EX/UV-VIS | Gauss | 2024-02-20 |  | 0.26321 | 0.0011033
078a_2 | 0.26397 | 0.0011059 | EX/UV-VIS | Gauss | 2024-02-20 |  | 0.26397 | 0.0011059
155a_2 | 0.26332 | 0.0011037 | EX/UV-VIS | Gauss | 2024-02-20 |  | 0.26332 | 0.0011037
232a_2 | 0.26311 | 0.001103 | EX/UV-VIS | Gauss | 2024-02-20 |  | 0.26311 | 0.001103
310a_2 | 0.26074 | 0.001095 | EX/UV-VIS | Gauss | 2024-02-20 |  | 0.26074 | 0.001095
052a_1 | 0.26335 | 0.00080082 | EX/UV-VIS | Gauss | 2024-03-05 13:38 |  | 0.26335 | 0.00080082
130a_1 | 0.26475 | 0.00080454 | EX/UV-VIS | Gauss | 2024-03-05 13:38 |  | 0.26475 | 0.00080454
207a_1 | 0.2637 | 0.00080175 | EX/UV-VIS | Gauss | 2024-03-05 13:38 |  | 0.2637 | 0.00080175
284a_1 | 0.26388 | 0.00080224 | EX/UV-VIS | Gauss | 2024-03-05 13:38 |  | 0.26388 | 0.00080224
360a_1 | 0.26433 | 0.00080343 | EX/UV-VIS | Gauss | 2024-03-05 13:38 |  | 0.26433 | 0.00080343
052a_2 | 0.2635 | 0.00080123 | EX/UV-VIS | Gauss | 2024-03-05 13:38 |  | 0.2635 | 0.00080123
130a_2 | 0.26521 | 0.00080578 | EX/UV-VIS | Gauss | 2024-03-05 13:38 |  | 0.26521 | 0.00080578
207a_2 | 0.26255 | 0.0007987 | EX/UV-VIS | Gauss | 2024-03-05 13:38 |  | 0.26255 | 0.0007987
284a_2 | 0.26383 | 0.00080209 | EX/UV-VIS | Gauss | 2024-03-05 13:38 |  | 0.26383 | 0.00080209
360a_2 | 0.2645 | 0.00080387 | EX/UV-VIS | Gauss | 2024-03-05 13:38 |  | 0.2645 | 0.00080387
No HOM component was found for phosphate
HOM data nitrite (verified 2024-05-10)
DerSimonian-Laird: combining data different day and/or method
Sample | w(NO2−) | uw(NO2−) | Cal/Method | Uncertainty | Date |  | w(NO2−) | uw(NO2−) |  | $mu | $s | $tau | Date
ID | μg/g | μg/g | ID | ID | YYYYMMDD |  | μg/g | μg/g |  | μg/g | μg/g | μg/g | YYYYMMDD
001_A1 | 0.07595 | 0.00023596 | ID/GCMS | Gauss | 2024-02-21 13:14 |  | 0.07595 | 0.00023596 |  | 0.075603 | 7.4336e-05 | 0 | 2024-02-21
078_A1 | 0.075767 | 0.00023547 | ID/GCMS | Gauss | 2024-02-21 13:28 |  | 0.075767 | 0.00023547 |  | 0.075417 | 7.5216e-05 | 0 | 2024-03-06
155_A1 | 0.075282 | 0.00023421 | ID/GCMS | Gauss | 2024-02-21 13:42 |  | 0.075282 | 0.00023421 |  | *** Data not linked. Calculation performed in R
232_A1 | 0.075755 | 0.00023544 | ID/GCMS | Gauss | 2024-02-21 13:56 |  | 0.075755 | 0.00023545 |  | and crossed-validated with https://consensus.nist.gov/app/nicob***
310_A1 | 0.075665 | 0.00023527 | ID/GCMS | Gauss | 2024-02-21 14:09 |  | 0.075665 | 0.00023527
001_A2 | 0.075394 | 0.00023443 | ID/GCMS | Gauss | 2024-02-21 14:23 |  | 0.075394 | 0.00023443
078_A2 | 0.075492 | 0.00023475 | ID/GCMS | Gauss | 2024-02-21 14:37 |  | 0.075492 | 0.00023475
155_A2 | 0.07585 | 0.00023581 | ID/GCMS | Gauss | 2024-02-21 14:51 |  | 0.07585 | 0.00023581
232_A2 | 0.075357 | 0.00023446 | ID/GCMS | Gauss | 2024-02-21 15:05 |  | 0.075357 | 0.00023446
310_A2 | 0.07553 | 0.00023492 | ID/GCMS | Gauss | 2024-02-21 15:19 |  | 0.07553 | 0.00023492
052_A1 | 0.075441 | 0.00023785 | ID/GCMS | Gauss | 2024-03-06 12:54 |  | 0.075441 | 0.00023785
130_A1 | 0.075416 | 0.00023783 | ID/GCMS | Gauss | 2024-03-06 13:08 |  | 0.075416 | 0.00023783
207_A1 | 0.075319 | 0.00023757 | ID/GCMS | Gauss | 2024-03-06 13:23 |  | 0.075319 | 0.00023757
284_A1 | 0.075677 | 0.00023865 | ID/GCMS | Gauss | 2024-03-06 13:37 |  | 0.075677 | 0.00023865
360_A1 | 0.075316 | 0.00023754 | ID/GCMS | Gauss | 2024-03-06 13:51 |  | 0.075316 | 0.00023754
052_A2 | 0.075486 | 0.00023806 | ID/GCMS | Gauss | 2024-03-06 14:05 |  | 0.075486 | 0.00023806
130_A2 | 0.075486 | 0.00023809 | ID/GCMS | Gauss | 2024-03-06 14:19 |  | 0.075486 | 0.00023809
207_A2 | 0.075318 | 0.00023757 | ID/GCMS | Gauss | 2024-03-06 14:34 |  | 0.075318 | 0.00023757
284_A2 | 0.07527 | 0.00023746 | ID/GCMS | Gauss | 2024-03-06 14:48 |  | 0.07527 | 0.00023746
360_A2 | 0.075441 | 0.00023793 | ID/GCMS | Gauss | 2024-03-06 15:02 |  | 0.075441 | 0.00023793
No HOM component was found for nitrite
HOM data silicate (verified 2024-05-10)
DerSimonian-Laird: combining data different day and/or method
Sample | w(SiO2) | uw(SiO2) | Cal/Method | Uncertainty | Date |  | w(SiO2) | uw(SiO2) |  | $mu | $s | $tau | Date
ID | μg/g | μg/g | ID | ID | YYYYMMDD |  | μg/g | μg/g |  | μg/g | μg/g | μg/g | YYYYMMDD
001a_1 | 0.42144 | 0.002829 | EX/UV-VIS | Gauss | 2024-02-21 |  | 0.42144 | 0.002829 |  | 0.42594 | 0.00090337 | 0 | 2024-02-21
078a_1 | 0.42354 | 0.0028419 | EX/UV-VIS | Gauss | 2024-02-21 |  | 0.42354 | 0.0028419 |  | 0.4178 | 0.00076914 | 0 | 2024-03-06
155a_1 | 0.42728 | 0.0028649 | EX/UV-VIS | Gauss | 2024-02-21 |  | 0.42728 | 0.0028649 |  | *** Data not linked. Calculation performed in R
232a_1 | 0.42484 | 0.0028499 | EX/UV-VIS | Gauss | 2024-02-21 |  | 0.42484 | 0.0028499 |  | and crossed-validated with https://consensus.nist.gov/app/nicob***
310a_1 | 0.42896 | 0.0028753 | EX/UV-VIS | Gauss | 2024-02-21 |  | 0.42896 | 0.0028753
001a_2 | 0.4237 | 0.0028429 | EX/UV-VIS | Gauss | 2024-02-21 |  | 0.4237 | 0.0028429
078a_2 | 0.42941 | 0.0028781 | EX/UV-VIS | Gauss | 2024-02-21 |  | 0.42941 | 0.0028781
155a_2 | 0.42421 | 0.002846 | EX/UV-VIS | Gauss | 2024-02-21 |  | 0.42421 | 0.002846
232a_2 | 0.42899 | 0.0028755 | EX/UV-VIS | Gauss | 2024-02-21 |  | 0.42899 | 0.0028755
310a_2 | 0.4273 | 0.002865 | EX/UV-VIS | Gauss | 2024-02-21 |  | 0.4273 | 0.002865
052a_1 | 0.4164 | 0.0024244 | EX/UV-VIS | Gauss | 2024-03-06 |  | 0.4164 | 0.0024244
130a_1 | 0.41649 | 0.0024249 | EX/UV-VIS | Gauss | 2024-03-06 |  | 0.41649 | 0.0024249
207a_1 | 0.41856 | 0.0024364 | EX/UV-VIS | Gauss | 2024-03-06 |  | 0.41856 | 0.0024364
284a_1 | 0.41694 | 0.0024274 | EX/UV-VIS | Gauss | 2024-03-06 |  | 0.41694 | 0.0024274
360a_1 | 0.4234 | 0.0024633 | EX/UV-VIS | Gauss | 2024-03-06 |  | 0.4234 | 0.0024633
052a_2 | 0.4184 | 0.0024355 | EX/UV-VIS | Gauss | 2024-03-06 |  | 0.4184 | 0.0024355
130a_2 | 0.41747 | 0.0024304 | EX/UV-VIS | Gauss | 2024-03-06 |  | 0.41747 | 0.0024304
207a_2 | 0.4143 | 0.0024128 | EX/UV-VIS | Gauss | 2024-03-06 |  | 0.4143 | 0.0024128
284a_2 | 0.41703 | 0.0024279 | EX/UV-VIS | Gauss | 2024-03-06 |  | 0.41703 | 0.0024279
360a_2 | 0.41924 | 0.0024402 | EX/UV-VIS | Gauss | 2024-03-06 |  | 0.41924 | 0.0024402
No HOM component was found for silicate

## LONG
Index
Line | Description
9 | Long term stability of the MOOS-3: summary (verified 2024-05-09)
66 | Long-term stability phosphate MOOS-4 (verified 2024-05-09)
93 | Monitoring of MOOS-4 and expiry date
Long term stability of the MOOS-3: summary (verified 2024-05-09)
The MOOS-3 material was collected in 1996 and sterilized by gamma radiation:
Clancy, V., Willie, S. Preparation and certification of a reference material for the determination of nutrients in seawater. Anal Bioanal Chem 378, 1239–1242 (2004). https://doi.org/10.1007/s00216-003-2473-1
The MOOS-3 was re-blended in 2011 and released in Novemebr 2013 and distributed until December 2022
a. Nitrite stability data
Date | w(NO2−) | uw(NO2−) | n | Notes
yyyy-mm-dd | μg/g | μg/g
2013-11-01 | 0.1581 | 0.0011498 | 13 | Certification data by ID/GCMS
2014-05-01 | 0.15825 | 0.0050115 | 3 | CM-13-18 by ID/GCMS
2014-12-01 | 0.15828 | 0.0007449 | 9 | CM-14-27 by ID/GCMS
2016-02-01 | 0.15982 | 0.0015452 | 10 | CM-16-03 by ID/GCMS
2017-06-01 | 0.1556 | 0.0067371 | 6 | CM-17-12 by ID/GCMS
2019-09-01 | 0.16429 | 0.0010059 | 6 | CM-19-05 by SA/UV-VIS
2021-06-01 | 0.1575 | 0.005267 | 14 | ICM-21-02 by EX/UV-VIS
b. Nitrate stability data
Date | w(NO3−) | uw(NO3−) | n | Notes
yyyy-mm-dd | μg/g | μg/g
2013-11-01 | 1.3842 | 0.0074611 | 16 | Certification data by ID/GCMS
2014-05-01 | 1.428 | 0.04587 | 3 | CM-13-18 by ID/GCMS
2014-12-01 | 1.4098 | 0.0080294 | 9 | CM-14-27 by ID/GCMS
2015-02-01 | 1.3812 | 0.0066793 | 4 | CM-15-08 by ID/GCMS
2016-02-01 | 1.4003 | 0.026364 | 10 | CM-16-03 by ID/GCMS
2017-08-01 | 1.3866 | 0.022192 | 9 | CM-17-19 by ID/GCMS
2019-09-01 | 1.3734 | 0.016547 | 3 | CM-19-05 by ID/GCMS
2021-06-01 | 1.3887 | 0.021063 | 21 | ICM-21-02 by ID/GCMS
c. Phosphate stability data
Date | w(PO43−) | uw(PO43−) | n | Notes
yyyy-mm-dd | μg/g | μg/g
2013-11-01 | 0.14745 | 0.0067274 | 25 | Certification data by SA/LC-ICPMS and SA/UV-VIS
2016-02-01 | 0.15919 | 0.012363 | 10 | CM-16-03 by SA/UV-VIS
2016-05-01 | 0.15499 | 0.0061199 | 10 | CM-16-05 by SA/UV-VIS
2017-06-01 | 0.15081 | 0.0049205 | 6 | CM-17-12 by SA/UV-VIS
2019-09-01 | 0.15348 | 0.010999 | 5 | CM-19-05 by SA/UV-VIS
2021-06-01 | 0.15628 | 0.010147 | 21 | ICM-21-02 by EX/UV-VIS
d. Silicate stability data
Date | w(SiO2) | uw(SiO2) | n | Notes
yyyy-mm-dd | μg/g | μg/g
2013-11-01 | 1.7713 | 0.021156 | 33 | Certification data by SA/LC-ICPMS, SA/UV-VIS and SA/ICPOES
2016-02-01 | 1.7617 | 0.040927 | 10 | CM-16-03 by SA/UV-VIS
2017-06-01 | 1.7257 | 0.021298 | 2 | CM-17-12 by SA/UV-VIS
2019-09-01 | 1.8092 | 0.1522 | 4 | CM-19-05 by SA/UV-VIS
2021-06-01 | 1.7467 | 0.010037 | 7 | ICM-21-02 by EX/UV-VIS
The MOOS-3 which was produced similarly to the MOOS-4 shows no trends in nutrient
mass fraction during a period of 7+ years
Long-term stability phosphate MOOS-4 (verified 2024-05-09)
Isochronous phosphate (21 °C) |  |  | Isochronous phosphate (40 °C)
Day | PO43− |  | Day | PO43−
number | formation % |  | number | formation %
0 | 0% |  | 0 | 0%
27 | 0.05137% |  | 27 | 0.5241%
35 | 0.05558% |  | 35 | 0.7232%
43 | 0.09246% |  | 43 | 0.9219%
LINEST at 21 °C |  |  | LINEST at 40 °C
1.981e-05 | -2.1473e-05 |  | 0.00021174 | -0.00013496
3.3019e-06 | 0.00010181 |  | 9.7398e-06 | 0.00030032
0.94736 | 0.00010683 |  | 0.99579 | 0.00031512
35.9942 | 2 |  | 472.6072 | 2
4.1078e-07 | 2.2825e-08 |  | 4.6929e-05 | 1.986e-07
3.4796 | Q10 temperature coefficient
v (5 °C) | 2.6942e-06
uR,LONG | 0.3933% | Covering a 4 yr period (k = 1)
Phosphate formation is minor and seems to be an order zero reaction
Phosphate formation could be due to the hydrolysis of non-reactive P-compounds:
Adina Paytan and Karen McLaughlin, the Oceanic Phosphorus Cycle, Chem. Rev. 2007, 107, 563−576
Claudia R. Benitez-Nelson, The biogeochemical cycling of phosphorus in marine systems, Earth-Science Reviews 51 (2000) 109–135
Monitoring of MOOS-4 and expiry date
Certification | March, 2024
Expiry date | March, 2028
Stability | March-September, 2026

## CERT
Index
Line | Description
13 | Value assignment for density in MOOS-4 (verified 2024-05-10)
18 | Value assignment for nitrite in MOOS-4 (verified 2024-05-10)
49 | Value assignment for nitrate in MOOS-4 (verified 2024-05-10)
80 | Value assignment for phosphate in MOOS-4 (verified 2024-05-10)
111 | Value assignment for silicate in MOOS-4 (verified 2024-05-10)
142 | MOOS-4 certified values (verified 2024-05-10)
151 | Comparing the uncertainty with previous CRMs
Value assignment for density in MOOS-4 (verified 2024-05-10)
Analyte | Unit | Density | u | uR | U
Density | g/mL | 1.021 | 0.00050111 | 0.04908% | 0.0010022
Value assignment for nitrite in MOOS-4 (verified 2024-05-10)
a. Mass fraction nitrite: uncertainty
uncertainty | value | notes
uR,CHAR | 0.237% | relative standard uncertainty due to characterization
uR,PRIM | 0.2875% | relative standard uncertainty due to the primary standard
uR,HOM | 0% | relative standard uncertainty due to homogeneiety
uR,SHORT | 0% | relative standard uncertainty due to short-term (transportation) conditions
uR,LONG | 0% | relative standard uncertainty due to long-term (storage at 5 °C) conditions
uR | 0.3726% | relative standard uncertainty (combined)
UR | 0.7451% | relative expanded uncertainty (combined, k = 2)
b. Molar concentration nitrite: uncertainty
uncertainty | value | notes
uR,CHAR | 0.237% | relative standard uncertainty due to characterization
uR,PRIM | 0.2875% | relative standard uncertainty due to the primary standard
uR,HOM | 0% | relative standard uncertainty due to homogeneiety
uR,SHORT | 0% | relative standard uncertainty due to short-term (transportation) conditions
uR,LONG | 0% | relative standard uncertainty due to long-term (storage at 5 °C) conditions
uR,DENSITY | 0.04908% | relative standard uncertainty due to unit conversion (density)
uR | 0.3758% | relative standard uncertainty (combined)
UR | 0.7516% | relative expanded uncertainty (combined, k = 2)
c. Nitrite: value assignment
measurand | value | exp uncert. | unit
w(NO2−) | 0.075523 | 0.00056275 | μg/g
c(NO2−) | 1.6761 | 0.012597 | μmol/L
Value assignment for nitrate in MOOS-4 (verified 2024-05-10)
a. Mass fraction nitrate: uncertainty
uncertainty | value | notes
uR,CHAR | 1.262% | relative standard uncertainty due to characterization
uR,PRIM | 0.09155% | relative standard uncertainty due to the primary standard
uR,HOM | 0.1977% | relative standard uncertainty due to homogeneiety
uR,SHORT | 0% | relative standard uncertainty due to short-term (transportation) conditions
uR,LONG | 0% | relative standard uncertainty due to long-term (storage at 5 °C) conditions
uR | 1.28% | relative standard uncertainty (combined)
UR | 2.561% | relative expanded uncertainty (combined, k = 2)
b. Molar concentration nitrate: uncertainty
uncertainty | value | notes
uR,CHAR | 1.262% | relative standard uncertainty due to characterization
uR,PRIM | 0.09155% | relative standard uncertainty due to the primary standard
uR,HOM | 0.1977% | relative standard uncertainty due to homogeneiety
uR,SHORT | 0% | relative standard uncertainty due to short-term (transportation) conditions
uR,LONG | 0% | relative standard uncertainty due to long-term (storage at 5 °C) conditions
uR,DENSITY | 0.04908% | relative standard uncertainty due to unit conversion (density)
uR | 1.281% | relative standard uncertainty (combined)
UR | 2.562% | relative expanded uncertainty (combined, k = 2)
c. Nitrate: value assignment
measurand | value | exp uncert. | unit
w(NO3−) | 1.0409 | 0.026654 | μg/g
c(NO3−) | 17.1404 | 0.43921 | μmol/L
Value assignment for phosphate in MOOS-4 (verified 2024-05-10)
a. Mass fraction phosphate: uncertainty
uncertainty | value | notes
uR,CHAR | 0.3306% | relative standard uncertainty due to characterization
uR,PRIM | 0.1794% | relative standard uncertainty due to the primary standard
uR,HOM | 0% | relative standard uncertainty due to homogeneiety
uR,SHORT | 0.06085% | relative standard uncertainty due to short-term (transportation) conditions
uR,LONG | 0.3933% | relative standard uncertainty due to long-term (storage at 5 °C) conditions
uR | 0.5476% | relative standard uncertainty (combined)
UR | 1.095% | relative expanded uncertainty (combined, k = 2)
b. Molar concentration phosphate: uncertainty
uncertainty | value | notes
uR,CHAR | 0.3306% | relative standard uncertainty due to characterization
uR,PRIM | 0.1794% | relative standard uncertainty due to the primary standard
uR,HOM | 0% | relative standard uncertainty due to homogeneiety
uR,SHORT | 0.06085% | relative standard uncertainty due to short-term (transportation) conditions
uR,LONG | 0.3933% | relative standard uncertainty due to long-term (storage at 5 °C) conditions
uR,DENSITY | 0.04908% | relative standard uncertainty due to unit conversion (density)
uR | 0.5498% | relative standard uncertainty (combined)
UR | 1.1% | relative expanded uncertainty (combined, k = 2)
c. Phosphate: value assignment
measurand | value | exp uncert. | unit
w(PO43−) | 0.26369 | 0.002888 | μg/g
c(PO43−) | 2.8348 | 0.031172 | μmol/L
Value assignment for silicate in MOOS-4 (verified 2024-05-10)
a. Mass fraction silicate: uncertainty
uncertainty | value | notes
uR,CHAR | 0.891% | relative standard uncertainty due to characterization
uR,PRIM | 0.1238% | relative standard uncertainty due to the primary standard
uR,HOM | 0% | relative standard uncertainty due to homogeneiety
uR,SHORT | 0% | relative standard uncertainty due to short-term (transportation) conditions
uR,LONG | 0% | relative standard uncertainty due to long-term (storage at 5 °C) conditions
uR | 0.8995% | relative standard uncertainty (combined)
UR | 1.799% | relative expanded uncertainty (combined, k = 2)
b. Molar concentration silicate: uncertainty
uncertainty | value | notes
uR,CHAR | 0.891% | relative standard uncertainty due to characterization
uR,PRIM | 0.1238% | relative standard uncertainty due to the primary standard
uR,HOM | 0% | relative standard uncertainty due to homogeneiety
uR,SHORT | 0% | relative standard uncertainty due to short-term (transportation) conditions
uR,LONG | 0% | relative standard uncertainty due to long-term (storage at 5 °C) conditions
uR,DENSITY | 0.04908% | relative standard uncertainty due to unit conversion (density)
uR | 0.9009% | relative standard uncertainty (combined)
UR | 1.802% | relative expanded uncertainty (combined, k = 2)
c. Silicate: value assignment
measurand | value | exp uncert. | unit
w(SiO2) | 0.42142 | 0.0075816 | μg/g
c(SiO2) | 7.1611 | 0.12902 | μmol/L
MOOS-4 certified values (verified 2024-05-10)
Analyte | Chemical | Molar concentration |  | Mass fraction |  | Value
name | formula | μmol/L |  | mg/kg |  | type
Nitrite | NO2− | 1.676 ± 0.013 |  | 0.07552 ± 0.00056 |  | certified
Nitrate | NO3− | 17.14 ± 0.44 |  | 1.041 ± 0.027 |  | certified
Phosphate | PO43− | 2.835 ± 0.031 |  | 0.2637 ± 0.0029 |  | certified
Silicate | SiO2 | 7.16 ± 0.13 |  | 0.4214 ± 0.0076 |  | certified
Comparing the uncertainty with previous CRMs
Analyte | MOOS-3 | SALT-1 | MOOS-4
Nitrite | 1.455% | N/A | 0.7451%
Nitrate | 1.084% | 2.379% | 2.561%
Phosphate | 8.844% | 1.89% | 1.095%
Silicate | 2.26% | 3.558% | 1.799%

## RAW_DATA
Index
Line | Description
31 | 2024-01-15 Nitrate short-term stability by ID HS GC-MS (verified on 2024-04-09)
61 | 2024-01-15 Nitrite short-term stability by UV-vis (verified on 2024-04-09)
88 | 2024-01-15 Phosphate short-term stability by UV-vis (verified 2024-04-09)
115 | 2024-01-15 Silicate short-term stability by UV-vis (verified on 2024-04-09)
142 | 2024-02-06 Standardization of Nitrite standard from Sigma Aldrich vs NIST 3185 by IDMS (verified on 2024-04-09)
242 | 2024-02-13 Standardization of Nitrite standard from Sigma Aldrich vs NIST 3185 by IDMS: verify the RA to RA' difference (verified 2024-04-09)
261 | 2024-02-14 Validation of nitrite and nitrate method and preliminary results on MOOS-4 and SEAN-1 (verified on 2024-04-09)
283 | 2024-02-20 CHAR MOOS4 NO3 w1d1 - GCMS (verified on 2024-04-09)
389 | 2024-02-21 CHAR MOOS4 NO2 w1d2 - GCMS (verified on 2024-04-09)
482 | 2024-02-20 CHAR MOOS4 NO2 w1d1 - spectrophotometry (verified on 2024-04-09)
508 | 2024-02-20 CHAR MOOS4 PO4 w1d1 - spectrophotometry (verified on 2024-04-09)
536 | 2024-02-21 CHAR MOOS4 SiO2 w1d2 - spectrophotometry (verified on 2024-04-09)
565 | 2024-02-21 CHAR MOOS4 NO2 w1d2 - spectrophotometry (verified on 2024-04-09)
586 | 2024-03-05 CHAR MOOS4 NO3 w2d1 - GCMS (verified on 2024-04-09)
707 | 2024-03-05 CHAR MOOS4 PO4 w2d1 - spectrophotometry (verified on 2024-04-09)
735 | 2024-03-06 CHAR MOOS4 NO2 w2d2 - GCMS (verified on 2024-04-09)
841 | 2024-03-06 CHAR MOOS4 SiO2 w2d2 - spectrophotometry (verified on 2024-04-09)
870 | 2024-03-21 CHAR MOOS4 SiO2 - LC-ICPMS (verified on 2024-04-24)
886 | 2024-05-02 Nitrite matrix effects (verified 2024-05-03)
898 | 2024-05-02 Phosphate matrix effects (verified 2024-05-03)
910 | 2024-05-02 Silicate matrix effects (verified 2024-05-03)
922 | 2024-05-06 Silicate interference on phosphate data (verified 2024-05-10)
933 | 2024-05-07 Temperature effect on nitrite determination (verified 2024-05-10)
943 | 2024-05-07 Temperature effect on phosphate determination (verified 2024-05-10)
953 | 2024-05-07 Temperature effect on silicate determination (verified 2024-05-10)
2024-01-15 Nitrate short-term stability by ID HS GC-MS (verified on 2024-04-09)
Sample |  | Nitrate_46 Results |  | Nitrate_47 Results |  | Ratio | sample
Data File | Acq. Date-Time | RT | Area | RT | Area | 46/47 | name
VOID.D | 2024-01-15 09:56
BLK.D | 2024-01-15 10:06 | 4.0326 | 25.145 | [5 empty] | sample | Ratio
BLK-IS.D | 2024-01-15 10:16 | 4.0327 | 81.5985 | 4.0347 | 11203.9725 |  |  |  | name | 46/47
MOOS237A1.D | 2024-01-15 10:26 | 4.0305 | 5822.2781 | 4.0325 | 5962.0385 | 0.97656 | MOOS237a1 |  | MOOS009a1 | 0.98799
MOOS351A2.D | 2024-01-15 10:36 | 4.0286 | 6366.658 | 4.0286 | 6551.38 | 0.9718 | MOOS351a2 |  | MOOS009a2 | 0.96314
MOOS161A1.D | 2024-01-15 10:45 | 4.0345 | 5508.9056 | 4.0345 | 5645.4301 | 0.97582 | MOOS161a1 |  | MOOS047a1 | 0.97974
MOOS123A2.D | 2024-01-15 10:55 | 4.0286 | 6118.8699 | 4.0306 | 6276.8512 | 0.97483 | MOOS123a2 |  | MOOS047a2 | 0.96659
MOOS161A2.D | 2024-01-15 11:05 | 4.0346 | 5392.288 | 4.0346 | 5571.5593 | 0.96782 | MOOS161a2 |  | MOOS085a1 | 0.97655
MOOS275A1.D | 2024-01-15 11:15 | 4.0346 | 5496.3211 | 4.0326 | 5626.555 | 0.97685 | MOOS275a1 |  | MOOS085a2 | 0.96849
MOOS009A2.D | 2024-01-15 11:25 | 4.0306 | 5878.3358 | 4.0306 | 6103.3016 | 0.96314 | MOOS009a2 |  | MOOS123a1 | 0.97738
MOOS313A2.D | 2024-01-15 11:35 | 4.0266 | 6046.7857 | 4.0266 | 6247.8105 | 0.96782 | MOOS313a2 |  | MOOS123a2 | 0.97483
MOOS199A1.D | 2024-01-15 11:45 | 4.0346 | 5328.3006 | 4.0346 | 5487.1325 | 0.97105 | MOOS199a1 |  | MOOS161a1 | 0.97582
MOOS237A2.D | 2024-01-15 11:54 | 4.0306 | 5767.5165 | 4.0306 | 5993.6324 | 0.96227 | MOOS237a2 |  | MOOS161a2 | 0.96782
MOOS047A2.D | 2024-01-15 12:04 | 4.0325 | 5742.0048 | 4.0325 | 5940.4705 | 0.96659 | MOOS047a2 |  | MOOS199a1 | 0.97105
MOOS085A1.D | 2024-01-15 12:14 | 4.0307 | 5946.7395 | 4.0327 | 6089.5452 | 0.97655 | MOOS085a1 |  | MOOS199a2 | 0.9654
MOOS199A2.D | 2024-01-15 12:24 | 4.0346 | 5465.3351 | 4.0326 | 5661.1918 | 0.9654 | MOOS199a2 |  | MOOS237a1 | 0.97656
MOOS123A1.D | 2024-01-15 12:34 | 4.0327 | 5565.7525 | 4.0327 | 5694.5661 | 0.97738 | MOOS123a1 |  | MOOS237a2 | 0.96227
MOOS009A1.D | 2024-01-15 12:44 | 4.0325 | 5420.4952 | 4.0325 | 5486.3846 | 0.98799 | MOOS009a1 |  | MOOS275a1 | 0.97685
MOOS313A1.D | 2024-01-15 12:54 | 4.0286 | 6036.265 | 4.0286 | 6161.5655 | 0.97966 | MOOS313a1 |  | MOOS275a2 | 0.96234
MOOS275A2.D | 2024-01-15 13:03 | 4.0306 | 5538.154 | 4.0326 | 5754.906 | 0.96234 | MOOS275a2 |  | MOOS313a1 | 0.97966
MOOS047A1.D | 2024-01-15 13:13 | 4.0286 | 5887.251 | 4.0286 | 6009.021 | 0.97974 | MOOS047a1 |  | MOOS313a2 | 0.96782
MOOS351A1.D | 2024-01-15 13:23 | 4.0325 | 5204.2018 | 4.0325 | 5315.1337 | 0.97913 | MOOS351a1 |  | MOOS351a1 | 0.97913
MOOS085A2.D | 2024-01-15 13:33 | 4.0327 | 5506.9673 | 4.0327 | 5686.1655 | 0.96849 | MOOS085a2 |  | MOOS351a2 | 0.9718
2024-01-16 Correction because of miss-labeling
blank contribution | 0.7283%
2024-01-15 Nitrite short-term stability by UV-vis (verified on 2024-04-09)
Sample ID | Date and Time | 541nm (Abs) |  | sample | signal
BLK | 1/15/2024 11:36:45 AM | 0.000175
BLK-IS | 1/15/2024 11:37:40 AM | 0.000668
MOOS351a1 | 1/15/2024 11:38:52 AM | 0.084748 |  | MOOS009a1 | 0.085949
MOOS313a2 | 1/15/2024 11:39:47 AM | 0.084852 |  | MOOS009a2 | 0.086121
MOOS085a1 | 1/15/2024 11:40:38 AM | 0.084193 |  | MOOS047a1 | 0.085613
MOOS237a2 | 1/15/2024 11:41:33 AM | 0.084691 |  | MOOS047a2 | 0.085773
MOOS351a2 | 1/15/2024 11:42:28 AM | 0.085299 |  | MOOS085a1 | 0.084193
MOOS047a2 | 1/15/2024 11:43:20 AM | 0.085773 |  | MOOS085a2 | 0.085289
MOOS275a1 | 1/15/2024 11:44:14 AM | 0.086173 |  | MOOS123a1 | 0.086719
MOOS009a2 | 1/15/2024 11:45:23 AM | 0.086121 |  | MOOS123a2 | 0.085916
MOOS161a2 | 1/15/2024 11:46:17 AM | 0.086119 |  | MOOS161a1 | 0.08647
MOOS313a1 | 1/15/2024 11:47:11 AM | 0.086039 |  | MOOS161a2 | 0.086119
MOOS123a2 | 1/15/2024 11:48:04 AM | 0.085916 |  | MOOS199a1 | 0.086315
MOOS199a1 | 1/15/2024 11:48:57 AM | 0.086315 |  | MOOS199a2 | 0.085523
MOOS275a2 | 1/15/2024 11:49:57 AM | 0.086322 |  | MOOS237a1 | 0.085775
MOOS123a1 | 1/15/2024 11:51:03 AM | 0.086719 |  | MOOS237a2 | 0.084691
MOOS009a1 | 1/15/2024 11:51:59 AM | 0.085949 |  | MOOS275a1 | 0.086173
MOOS085a2 | 1/15/2024 11:52:55 AM | 0.085289 |  | MOOS275a2 | 0.086322
MOOS237a1 | 1/15/2024 11:53:56 AM | 0.085775 |  | MOOS313a1 | 0.086039
MOOS161a1 | 1/15/2024 11:55:00 AM | 0.08647 |  | MOOS313a2 | 0.084852
MOOS199a2 | 1/15/2024 11:56:03 AM | 0.085523 |  | MOOS351a1 | 0.084748
MOOS047a1 | 1/15/2024 11:57:02 AM | 0.085613 |  | MOOS351a2 | 0.085299
Blank + matrix | 0.4919%
2024-01-15 Phosphate short-term stability by UV-vis (verified 2024-04-09)
Sample ID | Date and Time | 890nm (Abs) |  | sample | signal
BLK | 1/15/2024 1:28:07 PM | -0.005606
BLK-IS | 1/15/2024 1:29:24 PM | -0.005439
MOOS199a2 | 1/15/2024 1:30:57 PM | 0.27091 |  | MOOS009a1 | 0.26948
MOOS085a1 | 1/15/2024 1:32:06 PM | 0.27006 |  | MOOS009a2 | 0.2708
MOOS085a2 | 1/15/2024 1:33:08 PM | 0.27059 |  | MOOS047a1 | 0.27071
MOOS313a1 | 1/15/2024 1:34:03 PM | 0.2727 |  | MOOS047a2 | 0.27071
MOOS047a2 | 1/15/2024 1:34:52 PM | 0.27071 |  | MOOS085a1 | 0.27006
MOOS047a1 | 1/15/2024 1:35:44 PM | 0.27071 |  | MOOS085a2 | 0.27059
MOOS351a1 | 1/15/2024 1:36:51 PM | 0.27294 |  | MOOS123a1 | 0.27003
MOOS237a2 | 1/15/2024 1:37:47 PM | 0.27077 |  | MOOS123a2 | 0.27058
MOOS313a2 | 1/15/2024 1:38:43 PM | 0.27231 |  | MOOS161a1 | 0.27256
MOOS351a2 | 1/15/2024 1:39:42 PM | 0.27325 |  | MOOS161a2 | 0.27167
MOOS275a2 | 1/15/2024 1:40:47 PM | 0.27083 |  | MOOS199a1 | 0.27053
MOOS199a1 | 1/15/2024 1:41:54 PM | 0.27053 |  | MOOS199a2 | 0.27091
MOOS237a1 | 1/15/2024 1:42:47 PM | 0.27094 |  | MOOS237a1 | 0.27094
MOOS161a1 | 1/15/2024 1:43:31 PM | 0.27256 |  | MOOS237a2 | 0.27077
MOOS123a1 | 1/15/2024 1:44:23 PM | 0.27003 |  | MOOS275a1 | 0.27059
MOOS009a2 | 1/15/2024 1:45:11 PM | 0.2708 |  | MOOS275a2 | 0.27083
MOOS009a1 | 1/15/2024 1:46:11 PM | 0.26948 |  | MOOS313a1 | 0.2727
MOOS123a2 | 1/15/2024 1:47:20 PM | 0.27058 |  | MOOS313a2 | 0.27231
MOOS161a2 | 1/15/2024 1:48:16 PM | 0.27167 |  | MOOS351a1 | 0.27294
MOOS275a1 | 1/15/2024 1:49:13 PM | 0.27059 |  | MOOS351a2 | 0.27325
Blank + matrix | -2.037%
2024-01-15 Silicate short-term stability by UV-vis (verified on 2024-04-09)
Sample ID | Date and Time | 810nm (Abs) |  | sample | signal
BLK | 1/15/2024 3:59:38 PM | 0.011716
BLK-IS | 1/15/2024 4:00:43 PM | 0.012798
MOOS313a1 | 1/15/2024 4:01:44 PM | 0.12022 |  | MOOS009a1 | 0.11918
MOOS009a2 | 1/15/2024 4:02:36 PM | 0.12198 |  | MOOS009a2 | 0.12198
MOOS199a2 | 1/15/2024 4:03:17 PM | 0.11901 |  | MOOS047a1 | 0.12144
MOOS047a2 | 1/15/2024 4:04:05 PM | 0.12087 |  | MOOS047a2 | 0.12087
MOOS085a2 | 1/15/2024 4:04:43 PM | 0.1203 |  | MOOS085a1 | 0.12169
MOOS085a1 | 1/15/2024 4:05:28 PM | 0.12169 |  | MOOS085a2 | 0.1203
MOOS009a1 | 1/15/2024 4:06:19 PM | 0.11918 |  | MOOS123a1 | 0.12115
MOOS161a1 | 1/15/2024 4:07:14 PM | 0.12055 |  | MOOS123a2 | 0.12104
MOOS123a2 | 1/15/2024 4:08:07 PM | 0.12104 |  | MOOS161a1 | 0.12055
MOOS275a2 | 1/15/2024 4:09:03 PM | 0.11995 |  | MOOS161a2 | 0.11883
MOOS161a2 | 1/15/2024 4:10:08 PM | 0.11883 |  | MOOS199a1 | 0.12064
MOOS237a2 | 1/15/2024 4:11:03 PM | 0.12 |  | MOOS199a2 | 0.11901
MOOS199a1 | 1/15/2024 4:12:15 PM | 0.12064 |  | MOOS237a1 | 0.1204
MOOS351a2 | 1/15/2024 4:13:05 PM | 0.11969 |  | MOOS237a2 | 0.12
MOOS275a1 | 1/15/2024 4:13:57 PM | 0.11773 |  | MOOS275a1 | 0.11773
MOOS237a1 | 1/15/2024 4:14:50 PM | 0.1204 |  | MOOS275a2 | 0.11995
MOOS313a2 | 1/15/2024 4:15:38 PM | 0.11918 |  | MOOS313a1 | 0.12022
MOOS351a1 | 1/15/2024 4:16:29 PM | 0.11916 |  | MOOS313a2 | 0.11918
MOOS047a1 | 1/15/2024 4:17:16 PM | 0.12144 |  | MOOS351a1 | 0.11916
MOOS123a1 | 1/15/2024 4:18:05 PM | 0.12115 |  | MOOS351a2 | 0.11969
Blank + matrix | 10.2%
2024-02-06 Standardization of Nitrite standard from Sigma Aldrich vs NIST 3185 by IDMS (verified on 2024-04-09)
Sample |  | Nitrite_45 Results |  | Nitrate_46 Results |  | Nitrate_47 Results |  | Ratio | sample |  |  | sample | Ratio_I | Ratio_II | Ratio_III | Ratio_IV | Ratio_V | Average | STDEV | RSD | u(Average)
Data File | Acq. Date-Time | RT | Area | RT | Area | RT | Area | 46/47 | name |  |  | name | 46/47 | 46/47 | 46/47 | 46/47 | 46/47 | 46/47 | 46/47 | 46/47 | 46/47
VOID_I.D | 2024-02-06 12:37 | [6 empty] | #VALUE! | VOID |  |  | VOID | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
WATER_I.D | 2024-02-06 12:51 | 2.4559 | 21.3963 | 5.8057 | 35.9045 |  |  | #VALUE! | WATER |  |  | WATER | #VALUE! | 5.2216 | 5.4863 | 4.5667 | 3.4785 | #VALUE! | #VALUE! | #VALUE!
BLK1_I.D | 2024-02-06 13:04 |  |  | 5.8037 | 273.8877 | 5.8057 | 15543.63 | 0.017621 | BLK1 |  |  | BLK1 | 0.017621 | 0.018128 | 0.017959 | 0.017961 | 0.018131 | 0.01796 | 0.00020777 | 1.157%
BLK2_I.D | 2024-02-06 13:18 | 2.4474 | 11.4046 | 5.7998 | 241.0911 | 5.8018 | 17208.7494 | 0.01401 | BLK2 |  |  | BLK2 | 0.01401 | 0.0144 | 0.014455 | 0.014352 | 0.01445 | 0.014333 | 0.00018552 | 1.294%
BLK3_I.D | 2024-02-06 13:31 | 2.4646 | 6.8748 | 5.8018 | 242.141 | 5.8018 | 15982.6332 | 0.01515 | BLK3 |  |  | BLK3 | 0.01515 | 0.015208 | 0.015205 | 0.015362 | 0.015447 | 0.015274 | 0.00012445 | 0.8148%
SULF_BLK1_I.D | 2024-02-06 13:45 | 2.4589 | 9.3628 | 5.8038 | 354.4114 | 5.8018 | 20366.9702 | 0.017401 | SULF_BLK1 |  |  | SULF_BLK1 | 0.017401 | 0.017567 | 0.017697 | 0.017725 | 0.01763 | 0.017604 | 0.00012888 | 0.7321%
SULF_BLK2_I.D | 2024-02-06 13:59 | 2.4675 | 6.3303 | 5.8039 | 330.39 | 5.8039 | 19511.5615 | 0.016933 | SULF_BLK2 |  |  | SULF_BLK2 | 0.016933 | 0.017111 | 0.017216 | 0.017176 | 0.017168 | 0.017121 | 0.0001114 | 0.6507%
SULF_BLK3_I.D | 2024-02-06 14:12 |  |  | 5.8037 | 274.4905 | 5.8037 | 17766.2925 | 0.01545 | SULF_BLK3 |  |  | SULF_BLK3 | 0.01545 | 0.01551 | 0.01558 | 0.015677 | 0.015902 | 0.015624 | 0.00017707 | 1.133%
AB1_I.D | 2024-02-06 14:26 | 2.4617 | 11.9207 | 5.7997 | 8757.859 | 5.7997 | 16588.2599 | 0.52796 | AB1 |  |  | AB1 | 0.52796 | 0.52842 | 0.52814 | 0.52777 | 0.52876 | 0.52821 | 0.00039165 | 0.07415% | 0.00017515
AB2A_I.D | 2024-02-06 14:39 | 2.4504 | 22.3117 | 5.7999 | 17972.8705 | 5.7999 | 17132.6102 | 1.049 | AB2A |  |  | AB2A | 1.049 | 1.0503 | 1.0462 | 1.0523 | 1.0503 | 1.0496 | 0.0022327 | 0.2127% | 0.00099849
AB2B_I.D | 2024-02-06 14:53 | 2.4647 | 13.7652 | 5.8019 | 17031.8885 | 5.8019 | 16326.094 | 1.0432 | AB2B |  |  | AB2B | 1.0432 | 1.042 | 1.0451 | 1.0436 | 1.0461 | 1.044 | 0.0016169 | 0.1549% | 0.0007231
AB2C_I.D | 2024-02-06 15:07 | 2.4645 | 15.5642 | 5.8046 | 15698.0605 | 5.8046 | 15093.1669 | 1.0401 | AB2C |  |  | AB2C | 1.0401 | 1.0389 | 1.0377 | 1.0383 | 1.0397 | 1.0389 | 0.00098008 | 0.09434% | 0.00043831
AB3_I.D | 2024-02-06 15:21 | 2.4675 | 20.4664 | 5.8038 | 31784.9999 | 5.8038 | 15097.5141 | 2.1053 | AB3 |  |  | AB3 | 2.1053 | 2.1024 | 2.1062 | 2.1063 | 2.1083 | 2.1057 | 0.0021503 | 0.1021% | 0.00096165
S1_I.D | 2024-02-06 15:35 | 2.4617 | 22.8758 | 5.8038 | 15884.112 | 5.8038 | 15056.4418 | 1.055 | S1 |  |  | S1 | 1.055 | 1.0543 | 1.0567 | 1.0558 | 1.0559 | 1.0555 | 0.00091546 | 0.08673% | 0.00040941
S2_I.D | 2024-02-06 15:49 | 2.4616 | 23.6908 | 5.8037 | 17242.5055 | 5.8037 | 16269.8585 | 1.0598 | S2 |  |  | S2 | 1.0598 | 1.062 | 1.0613 | 1.0622 | 1.0612 | 1.0613 | 0.00093658 | 0.08825% | 0.00041885
S3_I.D | 2024-02-06 16:03 | 2.456 | 19.4895 | 5.8018 | 17804.2037 | 5.8018 | 16748.827 | 1.063 | S3 |  |  | S3 | 1.063 | 1.0661 | 1.0656 | 1.0665 | 1.0621 | 1.0647 | 0.0019693 | 0.185% | 0.00088072
SULF_S1_I.D | 2024-02-06 16:16 | 2.4733 | 18.7427 | 5.8039 | 258.458 | 5.8039 | 18777.2845 | 0.013764 | SULF_S1 |  |  | SULF_S1 | 0.013764 | 0.013778 | 0.013928 | 0.013947 | 0.014007 | 0.013885 | 0.00010776 | 0.7761%
SULF_S2_I.D | 2024-02-06 16:30 | 2.4502 | 27.8141 | 5.7997 | 377.3748 | 5.7997 | 22269.472 | 0.016946 | SULF_S2 |  |  | SULF_S2 | 0.016946 | 0.017361 | 0.017349 | 0.0172 | 0.017544 | 0.01728 | 0.00022307 | 1.291%
SULF_S3_I.D | 2024-02-06 16:44 | 2.4646 | 19.8644 | 5.8038 | 353.462 | 5.8018 | 20736.2356 | 0.017046 | SULF_S3 |  |  | SULF_S3 | 0.017046 | 0.016976 | 0.017286 | 0.017194 | 0.017115 | 0.017123 | 0.00012189 | 0.7119%
VOID_II.D | 2024-02-06 16:58 |  |  |  |  | 5.8059 | 7.011 | #VALUE! | VOID
WATER_II.D | 2024-02-06 17:12 | 2.4703 | 10.8111 | 5.8037 | 29.0427 | 5.8078 | 5.562 | 5.2216 | WATER
BLK1_II.D | 2024-02-06 17:26 | 2.453 | 17.0036 | 5.8005 | 329.298 | 5.8005 | 18165.0225 | 0.018128 | BLK1
BLK2_II.D | 2024-02-06 17:39 | 2.4532 | 12.7188 | 5.7999 | 263.299 | 5.7999 | 18284.5593 | 0.0144 | BLK2
BLK3_II.D | 2024-02-06 17:53 | 2.4502 | 8.6998 | 5.8017 | 279.5881 | 5.7997 | 18384.7845 | 0.015208 | BLK3
SULF_BLK1_II.D | 2024-02-06 18:07 | 2.4588 | 18.4914 | 5.8018 | 383.7015 | 5.8018 | 21841.5645 | 0.017567 | SULF_BLK1
SULF_BLK2_II.D | 2024-02-06 18:21 | 2.4645 | 8.5767 | 5.8037 | 332.4807 | 5.8037 | 19430.249 | 0.017111 | SULF_BLK2
SULF_BLK3_II.D | 2024-02-06 18:35 | 2.4703 | 8.6988 | 5.8057 | 289.1967 | 5.8037 | 18645.658 | 0.01551 | SULF_BLK3
AB1_II.D | 2024-02-06 18:49 | 2.4589 | 19.2429 | 5.8018 | 9623.3795 | 5.8018 | 18211.7348 | 0.52842 | AB1
AB2A_II.D | 2024-02-06 19:02 | 2.4732 | 5.7112 | 5.8038 | 16045.8635 | 5.8038 | 15277.497 | 1.0503 | AB2A
AB2B_II.D | 2024-02-06 19:16 | 2.4676 | 4.423 | 5.8039 | 16222.9435 | 5.8039 | 15569.1746 | 1.042 | AB2B
AB2C_II.D | 2024-02-06 19:30 | 2.4618 | 4.5412 | 5.8019 | 18022.6265 | 5.8019 | 17347.358 | 1.0389 | AB2C
AB3_II.D | 2024-02-06 19:44 | 2.4617 | 6.1849 | 5.8017 | 35592.4564 | 5.8017 | 16929.753 | 2.1024 | AB3
S1_II.D | 2024-02-06 19:58 | 2.4789 | 5.5569 | 5.8038 | 15107 | 5.8058 | 14329.0443 | 1.0543 | S1
S2_II.D | 2024-02-06 20:12 | 2.4646 | 11.0591 | 5.8018 | 18383.0275 | 5.8018 | 17310.2347 | 1.062 | S2
S3_II.D | 2024-02-06 20:25 | 2.4589 | 11.913 | 5.8006 | 19526.4546 | 5.8006 | 18315.2091 | 1.0661 | S3
SULF_S1_II.D | 2024-02-06 20:39 | 2.4701 | 6.37 | 5.8036 | 270.9605 | 5.8036 | 19666.0807 | 0.013778 | SULF_S1
SULF_S2_II.D | 2024-02-06 20:53 | 2.4789 | 8.0435 | 5.8026 | 353.295 | 5.8046 | 20349.6382 | 0.017361 | SULF_S2
SULF_S3_II.D | 2024-02-06 21:07 | 2.4644 | 10.4043 | 5.8036 | 373.4159 | 5.8016 | 21997.1475 | 0.016976 | SULF_S3
VOID_III.D | 2024-02-06 21:21 |  |  |  |  | 5.8079 | 8.825 | #VALUE! | VOID
WATER_III.D | 2024-02-06 21:35 | 2.4731 | 16.3512 | 5.8037 | 32.8247 | 5.8118 | 5.983 | 5.4863 | WATER
BLK1_III.D | 2024-02-06 21:49 | 2.4532 | 11.1287 | 5.8018 | 341.153 | 5.7998 | 18995.7222 | 0.017959 | BLK1
BLK2_III.D | 2024-02-06 22:02 | 2.4561 | 10.8609 | 5.8019 | 261.4845 | 5.8019 | 18090.1514 | 0.014455 | BLK2
BLK3_III.D | 2024-02-06 22:16 | 2.4647 | 5.3078 | 5.8019 | 235.0495 | 5.8039 | 15458.5088 | 0.015205 | BLK3
SULF_BLK1_III.D | 2024-02-06 22:30 |  |  | 5.8018 | 394.5755 | 5.8039 | 22296.164 | 0.017697 | SULF_BLK1
SULF_BLK2_III.D | 2024-02-06 22:44 | 2.4561 | 26.1354 | 5.7978 | 383.9935 | 5.8019 | 22304.372 | 0.017216 | SULF_BLK2
SULF_BLK3_III.D | 2024-02-06 22:58 | 2.4645 | 18.9335 | 5.8037 | 321.469 | 5.8017 | 20632.9435 | 0.01558 | SULF_BLK3
AB1_III.D | 2024-02-06 23:12 | 2.459 | 11.9573 | 5.8019 | 9192.782 | 5.8019 | 17406.0483 | 0.52814 | AB1
AB2A_III.D | 2024-02-06 23:26 | 2.4645 | 7.5133 | 5.8037 | 17994.9755 | 5.8037 | 17200.006 | 1.0462 | AB2A
AB2B_III.D | 2024-02-06 23:40 | 2.4732 | 5.2904 | 5.8038 | 17198.054 | 5.8038 | 16456.3889 | 1.0451 | AB2B
AB2C_III.D | 2024-02-06 23:53 | 2.4561 | 21.0975 | 5.8039 | 17456.9726 | 5.8019 | 16822.9225 | 1.0377 | AB2C
AB3_III.D | 2024-02-07 00:07 | 2.4673 | 6.8703 | 5.8037 | 35489.7059 | 5.8037 | 16850.1905 | 2.1062 | AB3
S1_III.D | 2024-02-07 00:21 | 2.4588 | 8.524 | 5.8017 | 18779.0319 | 5.8017 | 17771.8599 | 1.0567 | S1
S2_III.D | 2024-02-07 00:35 |  |  | 5.8039 | 15591.531 | 5.8039 | 14690.3252 | 1.0613 | S2
S3_III.D | 2024-02-07 00:49 | 2.4616 | 11.4515 | 5.8017 | 19611.541 | 5.8017 | 18404.1704 | 1.0656 | S3
SULF_S1_III.D | 2024-02-07 01:03 | 2.4646 | 8.034 | 5.8038 | 285.425 | 5.8038 | 20493.5267 | 0.013928 | SULF_S1
SULF_S2_III.D | 2024-02-07 01:17 | 2.4588 | 8.6952 | 5.8017 | 409.572 | 5.8017 | 23607.209 | 0.017349 | SULF_S2
SULF_S3_III.D | 2024-02-07 01:31 | 2.456 | 15.9125 | 5.8046 | 392.5573 | 5.8026 | 22709.15 | 0.017286 | SULF_S3
VOID_IV.D | 2024-02-07 01:44 | [6 empty] | #VALUE! | VOID
WATER_IV.D | 2024-02-07 01:58 | 2.4673 | 18.7886 | 5.8017 | 33.1244 | 5.8117 | 7.2535 | 4.5667 | WATER
BLK1_IV.D | 2024-02-07 02:12 | 2.456 | 9.9067 | 5.7998 | 332.8073 | 5.8018 | 18529.804 | 0.017961 | BLK1
BLK2_IV.D | 2024-02-07 02:26 | 2.4559 | 19.7605 | 5.7996 | 264.0225 | 5.8017 | 18396.1876 | 0.014352 | BLK2
BLK3_IV.D | 2024-02-07 02:40 | 2.4646 | 30.6937 | 5.8018 | 248.814 | 5.8038 | 16196.9641 | 0.015362 | BLK3
SULF_BLK1_IV.D | 2024-02-07 02:54 | 2.4644 | 24.0094 | 5.8037 | 371.0922 | 5.8037 | 20935.7731 | 0.017725 | SULF_BLK1
SULF_BLK2_IV.D | 2024-02-07 03:08 | 2.4588 | 11.2747 | 5.7997 | 347.3805 | 5.8017 | 20225.018 | 0.017176 | SULF_BLK2
SULF_BLK3_IV.D | 2024-02-07 03:21 | 2.4559 | 9.4392 | 5.7997 | 328.754 | 5.8017 | 20969.965 | 0.015677 | SULF_BLK3
AB1_IV.D | 2024-02-07 03:35 | 2.4646 | 19.7342 | 5.8018 | 8553.984 | 5.8039 | 16207.7985 | 0.52777 | AB1
AB2A_IV.D | 2024-02-07 03:49 | 2.4561 | 11.1298 | 5.8039 | 16630.747 | 5.8039 | 15804.218 | 1.0523 | AB2A
AB2B_IV.D | 2024-02-07 04:03 | 2.4588 | 7.328 | 5.8017 | 18977.469 | 5.8017 | 18183.9913 | 1.0436 | AB2B
AB2C_IV.D | 2024-02-07 04:17 | 2.4531 | 24.007 | 5.8018 | 19385.95 | 5.8018 | 18671.4155 | 1.0383 | AB2C
AB3_IV.D | 2024-02-07 04:31 | 2.4616 | 29.0786 | 5.8037 | 31974.1575 | 5.8037 | 15180.34 | 2.1063 | AB3
S1_IV.D | 2024-02-07 04:44 | 2.4561 | 13.2483 | 5.8019 | 17906.6455 | 5.8019 | 16960.2706 | 1.0558 | S1
S2_IV.D | 2024-02-07 04:58 |  |  | 5.8019 | 18284.5318 | 5.8019 | 17214.452 | 1.0622 | S2
S3_IV.D | 2024-02-07 05:12 | 2.4618 | 9.8891 | 5.8019 | 18024.086 | 5.8019 | 16900.623 | 1.0665 | S3
SULF_S1_IV.D | 2024-02-07 05:26 |  |  | 5.8038 | 289.2915 | 5.8038 | 20742.4585 | 0.013947 | SULF_S1
SULF_S2_IV.D | 2024-02-07 05:40 | 2.4561 | 7.4693 | 5.7999 | 397.8482 | 5.8019 | 23131.1331 | 0.0172 | SULF_S2
SULF_S3_IV.D | 2024-02-07 05:54 | 2.4645 | 21.4049 | 5.8017 | 368.1585 | 5.8037 | 21412.445 | 0.017194 | SULF_S3
VOID_V.D | 2024-02-07 06:08 |  |  |  |  | 5.8038 | 11.7127 | #VALUE! | VOID
WATER_V.D | 2024-02-07 06:21 | 2.4475 | 24.1715 | 5.7999 | 38.6882 | 5.8039 | 11.122 | 3.4785 | WATER
BLK1_V.D | 2024-02-07 06:35 | 2.4589 | 11.2664 | 5.7998 | 306.391 | 5.7998 | 16898.8875 | 0.018131 | BLK1
BLK2_V.D | 2024-02-07 06:49 |  |  | 5.8057 | 198.7119 | 5.8037 | 13752.1225 | 0.01445 | BLK2
BLK3_V.D | 2024-02-07 07:03 | 2.4475 | 12.0703 | 5.8018 | 283.5359 | 5.7998 | 18355.8872 | 0.015447 | BLK3
SULF_BLK1_V.D | 2024-02-07 07:17 | 2.4503 | 77.6565 | 5.8018 | 399.6141 | 5.8018 | 22667.195 | 0.01763 | SULF_BLK1
SULF_BLK2_V.D | 2024-02-07 07:31 | 2.4647 | 30.3054 | 5.8027 | 297.6741 | 5.8027 | 17339.18 | 0.017168 | SULF_BLK2
SULF_BLK3_V.D | 2024-02-07 07:44 | 2.4617 | 11.0383 | 5.7997 | 308.095 | 5.8017 | 19374.2475 | 0.015902 | SULF_BLK3
AB1_V.D | 2024-02-07 07:58 |  |  | 5.8037 | 6318.206 | 5.8037 | 11949.0102 | 0.52876 | AB1
AB2A_V.D | 2024-02-07 08:12 | 2.4474 | 8.3486 | 5.7997 | 15427.698 | 5.7997 | 14689.3317 | 1.0503 | AB2A
AB2B_V.D | 2024-02-07 08:26 | 2.4532 | 7.1375 | 5.8018 | 19160.7853 | 5.8018 | 18315.7887 | 1.0461 | AB2B
AB2C_V.D | 2024-02-07 08:40 | 2.4532 | 13.8811 | 5.8019 | 19059.4585 | 5.8019 | 18332.1855 | 1.0397 | AB2C
AB3_V.D | 2024-02-07 08:54 | 2.4561 | 13.2532 | 5.8019 | 32550.0505 | 5.8019 | 15439.209 | 2.1083 | AB3
S1_V.D | 2024-02-07 09:08 | 2.4589 | 7.2947 | 5.8039 | 12928.7202 | 5.8039 | 12244.3662 | 1.0559 | S1
S2_V.D | 2024-02-07 09:21 | 2.4589 | 41.3768 | 5.8007 | 18232.817 | 5.8027 | 17180.864 | 1.0612 | S2
S3_V.D | 2024-02-07 09:35 | 2.4646 | 14.9878 | 5.8038 | 15890.0975 | 5.8038 | 14960.717 | 1.0621 | S3
SULF_S1_V.D | 2024-02-07 09:49 | 2.4531 | 9.4871 | 5.7998 | 280.763 | 5.8018 | 20044.7988 | 0.014007 | SULF_S1
SULF_S2_V.D | 2024-02-07 10:03 | 2.4617 | 21.2575 | 5.8017 | 408.7845 | 5.8017 | 23301.043 | 0.017544 | SULF_S2
SULF_S3_V.D | 2024-02-07 10:17 | 2.4617 | 66.8668 | 5.8018 | 403.4718 | 5.8018 | 23574.723 | 0.017115 | SULF_S3
2024-02-13 Standardization of Nitrite standard from Sigma Aldrich vs NIST 3185 by IDMS: verify the RA to RA' difference (verified 2024-04-09)
Sample |  | Nitrate_46 Results |  | Nitrate_47 Results |  | Ratio
Data File | Acq. Date-Time | RT | Area | RT | Area | 46/47 |  | Sample | rAB/(V/V)
NO2_SIGMA_1.D | 2024-02-13 11:00 | 4.0346 | 17402.2325 | 4.0346 | 202.495 | 85.9391 |  | Sigma_1 | 85.9391
NO3_NIST_1.D | 2024-02-13 11:10 | 4.0365 | 15385.1938 | 4.0365 | 183.5085 | 83.8391 |  | Sigma_2 | 83.5038
NO2_SIGMA_2.D | 2024-02-13 11:20 | 4.0346 | 15449.6588 | 4.0325 | 185.0175 | 83.5038 |  | Sigma_3 | 84.722
NO3_NIST_2.D | 2024-02-13 11:29 | 4.0307 | 18135.2803 | 4.0327 | 215.557 | 84.1322 |  | Sigma_4 | 84.6215
NO2_SIGMA_3.D | 2024-02-13 11:39 | 4.0305 | 17348.7896 | 4.0305 | 204.7732 | 84.722 |  | Sigma_5 | 82.2031
NO3_NIST_3.D | 2024-02-13 11:49 | 4.0325 | 17171.6775 | 4.0325 | 207.6805 | 82.6831 |  | Sigma_6 | 83.7705
NO2_SIGMA_4.D | 2024-02-13 11:59 | 4.0307 | 17472.009 | 4.0327 | 206.4725 | 84.6215 |  | Sigma_7 | 83.3213
NO3_NIST_4.D | 2024-02-13 12:08 | 4.0346 | 15550.4042 | 4.0366 | 187.252 | 83.0453 |  | NIST_1 | 83.8391
NO2_SIGMA_5.D | 2024-02-13 12:18 | 4.0345 | 16535.7242 | 4.0366 | 201.157 | 82.2031 |  | NIST_2 | 84.1322
NO3_NIST_5.D | 2024-02-13 12:28 | 4.0326 | 17620.8386 | 4.0347 | 212.279 | 83.0079 |  | NIST_3 | 82.6831
NO2_SIGMA_6.D | 2024-02-13 12:38 | 4.0347 | 15737.2925 | 4.0327 | 187.862 | 83.7705 |  | NIST_4 | 83.0453
NO3_NIST_6.D | 2024-02-13 12:48 | 4.0305 | 17664.5265 | 4.0305 | 213.272 | 82.8263 |  | NIST_5 | 83.0079
NO2_SIGMA_7.D | 2024-02-13 12:57 | 4.0345 | 16347.38 | 4.0345 | 196.197 | 83.3213 |  | NIST_6 | 82.8263
NO3_NIST_7.D | 2024-02-13 13:07 | 4.0306 | 17797.8995 | 4.0327 | 214.4525 | 82.9923 |  | NIST_7 | 82.9923
2024-02-14 Validation of nitrite and nitrate method and preliminary results on MOOS-4 and SEAN-1 (verified on 2024-04-09)
Sample |  | Nitrite_31 Results |  | Nitrite_32 Results |  | Nitrate_46 Results |  | Nitrate_47 Results
Data File | Acq. Date-Time | RT | Area | RT | Area | RT | Area | RT | Area
VOID.D | 2024-02-14 12:17
BLK_M1.D | 2024-02-14 12:31 | 2.4548 | 33.1798 |  |  | 5.8034 | 63.3455
BLK_1.D | 2024-02-14 12:45 | 2.4528 | 4.0593 |  |  | 5.8014 | 31.2491
BLK_M2.D | 2024-02-14 12:59 | 2.4628 | 26.4855 |  |  | 5.8034 | 39.108
BLK_2.D | 2024-02-14 13:13 | 2.4365 | 3.9274 |  |  | 5.8073 | 30.6688
BLK_M3.D | 2024-02-14 13:26 | 2.4547 | 26.03 |  |  | 5.8053 | 41.947
BLK_3.D | 2024-02-14 13:40 | 2.4367 | 3.6349 |  |  | 5.7995 | 29.6838
LOD_1.D | 2024-02-14 13:54 | 2.4588 | 44.2184 |  |  | 5.8034 | 255.805
LOD_2.D | 2024-02-14 14:08 | 2.4506 | 46.1822 |  |  | 5.8013 | 282.1934
LOD_3.D | 2024-02-14 14:22 | 2.4609 | 45.3955 |  |  | 5.8055 | 284.542 | 5.8015 | 4.11
CAL_5.D | 2024-02-14 14:36 | 2.4587 | 420.0154 | 2.4607 | 3048.9851 | 5.8033 | 2196.9592 | 5.8033 | 18759.6954
CAL_4.D | 2024-02-14 14:50 | 2.4529 | 983.7853 | 2.4529 | 3057.1758 | 5.8015 | 4876.6445 | 5.8015 | 18735.4776
CAL_3.D | 2024-02-14 15:04 | 2.4606 | 1962.4017 | 2.4606 | 3148.3171 | 5.8032 | 9719.639 | 5.8032 | 19328.18
CAL_2.D | 2024-02-14 15:18 | 2.4627 | 2713.2296 | 2.4607 | 2931.6985 | 5.8033 | 12998.8232 | 5.8033 | 17489.1105
CAL_1.D | 2024-02-14 15:32 | 2.4648 | 3559.5854 | 2.4668 | 2906.59 | 5.8054 | 17380.055 | 5.8034 | 17406.2999
SEAN1797.D | 2024-02-14 15:46 | 2.4506 | 878.6904 | 2.4527 | 2375.5602 | 5.8013 | 4811.443 | 5.8013 | 9018.527
MOOS2157.D | 2024-02-14 16:00 | 2.4567 | 2229.7591 | 2.4587 | 2308.6759 | 5.8013 | 8910.6326 | 5.8013 | 8814.154
2024-02-20 CHAR MOOS4 NO3 w1d1 - GCMS (verified on 2024-04-09)
Sample |  | Nitrate_46 Results |  | Nitrate_47 Results |  | Nitrate_46 blk corr |  | BLK no | BLK yes | Isotope ratio
Data File | Acq. Date-Time | RT | Area | RT | Area | RT | Area | rAB | rAB | rAB | u[rAB]
VOID_I.D | 2024-02-20 12:57
BLK_DIW_I.D | 2024-02-20 13:07 | 4.0327 | 180.1218
BLK_NACL_I.D | 2024-02-20 13:17 | 4.0266 | 116.8747
BLK_IS_DIW_I.D | 2024-02-20 13:27 | 4.0266 | 264.567 | 4.0286 | 17760.5546
BLK_IS_NACL_I.D | 2024-02-20 13:37 | 4.0285 | 748.6685 | 4.0285 | 7589.8613 | [7 empty] | Sample ID | rAB | u[rAB] | Date and Time
AB1_I.D | 2024-02-20 13:48 | 4.0306 | 8740.67 | 4.0306 | 17745.877 |  |  | 0.49255 | 0.49255 | 0.49255 | 0.010377 |  | AB1 | 0.49291 | 0.010385 | 2024-02-20 13:48
AB2A_I.D | 2024-02-20 13:58 | 4.0325 | 17371.5239 | 4.0325 | 17601.8058 |  |  | 0.98692 | 0.98692 | 0.98692 | 0.020793 |  | AB2A | 0.98737 | 0.020803 | 2024-02-20 13:58
AB2B_I.D | 2024-02-20 14:08 | 4.0345 | 16346.8661 | 4.0345 | 17071.1165 |  |  | 0.95757 | 0.95757 | 0.95757 | 0.020175 |  | AB2B | 0.9571 | 0.020165 | 2024-02-20 14:08
AB3_I.D | 2024-02-20 14:18 | 4.0286 | 32586.7743 | 4.0286 | 17437.9985 |  |  | 1.8687 | 1.8687 | 1.8687 | 0.039372 |  | AB3 | 1.8695 | 0.039388 | 2024-02-20 14:18
001_A1_I.D | 2024-02-20 14:28 | 4.0326 | 8977.5153 | 4.0326 | 8476.789 |  | 8944.638 | 1.0591 | 1.0552 | 1.0571 | 0.0095515 |  | 001_A1 | 1.0578 | 0.0095499 | 2024-02-20 14:28
078_A1_I.D | 2024-02-20 14:38 | 4.0305 | 8266.7175 | 4.0305 | 8132.9368 |  | 8233.8403 | 1.0164 | 1.0124 | 1.0144 | 0.0091996 |  | 078_A1 | 1.0146 | 0.0091909 | 2024-02-20 14:38
155_A1_I.D | 2024-02-20 14:48 | 4.0264 | 8459.0005 | 4.0264 | 8414.8134 |  | 8426.1233 | 1.0053 | 1.0013 | 1.0033 | 0.0090887 |  | 155_A1 | 1.0015 | 0.0090629 | 2024-02-20 14:48
232_A1_I.D | 2024-02-20 14:58 | 4.0327 | 8261.2909 | 4.0327 | 8247.2732 |  | 8228.4137 | 1.0017 | 0.99771 | 0.99971 | 0.0090664 |  | 232_A1 | 1.0016 | 0.0090679 | 2024-02-20 14:58
310_A1_I.D | 2024-02-20 15:08 | 4.0286 | 8327.8528 | 4.0286 | 8317.2193 |  | 8294.9755 | 1.0013 | 0.99733 | 0.9993 | 0.0090592 |  | 310_A1 | 1.0003 | 0.0090689 | 2024-02-20 15:08
001_A2_I.D | 2024-02-20 15:18 | 4.0305 | 8107.0173 | 4.0305 | 8064.664 |  | 8074.1401 | 1.0053 | 1.0012 | 1.0032 | 0.0091066 |  | 001_A2 | 1.0054 | 0.0091219 | 2024-02-20 15:18
078_A2_I.D | 2024-02-20 15:28 | 4.0306 | 8471.0005 | 4.0306 | 8465.372 |  | 8438.1233 | 1.0007 | 0.99678 | 0.99872 | 0.0090467 |  | 078_A2 | 0.99937 | 0.0090416 | 2024-02-20 15:28
155_A2_I.D | 2024-02-20 15:38 | 4.0286 | 8384.0465 | 4.0286 | 8311.1918 |  | 8351.1693 | 1.0088 | 1.0048 | 1.0068 | 0.0091241 |  | 155_A2 | 1.0058 | 0.0091044 | 2024-02-20 15:38
232_A2_I.D | 2024-02-20 15:48 | 4.0285 | 8512.6958 | 4.0285 | 8482.0318 |  | 8479.8186 | 1.0036 | 0.99974 | 1.0017 | 0.0090714 |  | 232_A2 | 1.0005 | 0.0090531 | 2024-02-20 15:48
310_A2_I.D | 2024-02-20 15:58 | 4.0286 | 9064.2432 | 4.0286 | 8478.2165 |  | 9031.3659 | 1.0691 | 1.0652 | 1.0672 | 0.0096385 |  | 310_A2 | 1.0689 | 0.0096427 | 2024-02-20 15:58
SALT1_113_I.D | 2024-02-20 16:08 | 4.0286 | 8883.0777 | 4.0306 | 8092.556 |  | 8848.7872 | 1.0977 | 1.0934 | 1.0956 | 0.0099215 |  | SALT1_113 | 1.0973 | 0.009934 | 2024-02-20 16:08
MOOS3_I.D | 2024-02-20 16:18 | 4.0325 | 9632.4345 | 4.0346 | 7304.3466 |  | 9596.2246 | 1.3187 | 1.3138 | 1.3162 | 0.011906 |  | MOOS3 | 1.3124 | 0.011853 | 2024-02-20 16:18
AB2C_I.D | 2024-02-20 16:28 | 4.0305 | 16577.1126 | 4.0305 | 16567.3802 |  |  | 1.0006 | 1.0006 | 1.0006 | 0.021081 |  | AB2C | 1.0012 | 0.021095 | 2024-02-20 16:28
VOID_II.D | 2024-02-20 16:38 | 4.0367 | 5.1793 | 4.0387 | 6.769
BLK_DIW_II.D | 2024-02-20 16:48 | 4.0284 | 175.2454 | 4.0345 | 6.2865
BLK_NACL_II.D | 2024-02-20 16:58 | 4.0285 | 110.6493
BLK_IS_DIW_II.D | 2024-02-20 17:08 | 4.0305 | 273.2675 | 4.0305 | 17786.384
BLK_IS_NACL_II.D | 2024-02-20 17:18 | 4.0345 | 718.4182 | 4.0345 | 7394.9142
AB1_II.D | 2024-02-20 17:28 | 4.0287 | 8832.235 | 4.0287 | 17899.5066 |  |  | 0.49343 | 0.49343 | 0.49343 | 0.010396
AB2A_II.D | 2024-02-20 17:38 | 4.0345 | 17480.646 | 4.0345 | 17702.1315 |  |  | 0.98749 | 0.98749 | 0.98749 | 0.020805
AB2B_II.D | 2024-02-20 17:48 | 4.0326 | 16897.5356 | 4.0326 | 17656.0804 |  |  | 0.95704 | 0.95704 | 0.95704 | 0.020164
AB3_II.D | 2024-02-20 17:58 | 4.0325 | 32443.3995 | 4.0325 | 17366.1908 |  |  | 1.8682 | 1.8682 | 1.8682 | 0.03936
001_A1_II.D | 2024-02-20 18:08 | 4.0305 | 9073.529 | 4.0325 | 8535.1225 |  | 9040.6518 | 1.0631 | 1.0592 | 1.0612 | 0.0095837
078_A1_II.D | 2024-02-20 18:18 | 4.0327 | 8312.713 | 4.0307 | 8175.8182 |  | 8279.8358 | 1.0167 | 1.0127 | 1.0147 | 0.0091999
155_A1_II.D | 2024-02-20 18:29 | 4.0286 | 8721.915 | 4.0266 | 8688.3851 |  | 8689.0378 | 1.0039 | 1.0001 | 1.002 | 0.0090642
232_A1_II.D | 2024-02-20 18:39 | 4.0286 | 8686.768 | 4.0286 | 8642.3419 |  | 8653.8908 | 1.0051 | 1.0013 | 1.0032 | 0.0090773
310_A1_II.D | 2024-02-20 18:49 | 4.0327 | 8144.8067 | 4.0327 | 8133.9135 |  | 8111.9295 | 1.0013 | 0.9973 | 0.99932 | 0.0090692
001_A2_II.D | 2024-02-20 18:59 | 4.0286 | 8331.1063 | 4.0286 | 8265.3725 |  | 8298.2291 | 1.008 | 1.004 | 1.006 | 0.0091194
078_A2_II.D | 2024-02-20 19:09 | 4.0286 | 8733.6335 | 4.0286 | 8700.7342 |  | 8700.7563 | 1.0038 | 1 | 1.0019 | 0.009063
155_A2_II.D | 2024-02-20 19:19 | 4.0325 | 8353.9675 | 4.0325 | 8291.9106 |  | 8321.0903 | 1.0075 | 1.0035 | 1.0055 | 0.009114
232_A2_II.D | 2024-02-20 19:29 | 4.0287 | 8710.4708 | 4.0287 | 8687.288 |  | 8677.5936 | 1.0027 | 0.99888 | 1.0008 | 0.0090539
310_A2_II.D | 2024-02-20 19:39 | 4.0285 | 9366.062 | 4.0285 | 8744.481 |  | 9333.1848 | 1.0711 | 1.0673 | 1.0692 | 0.0096444
SALT1_113_II.D | 2024-02-20 19:49 | 4.0326 | 8874.5065 | 4.0326 | 8058.9675 |  | 8840.216 | 1.1012 | 1.0969 | 1.0991 | 0.0099536
MOOS3_II.D | 2024-02-20 19:59 | 4.0326 | 10050.2679 | 4.0326 | 7661.3499 |  | 10014.0581 | 1.3118 | 1.3071 | 1.3095 | 0.011823
AB2C_II.D | 2024-02-20 20:09 | 4.0345 | 16725.2636 | 4.0345 | 16719.9727 |  |  | 1.0003 | 1.0003 | 1.0003 | 0.021075
VOID_III.D | 2024-02-20 20:19 | 4.0366 | 7.0774
BLK_DIW_III.D | 2024-02-20 20:29 | 4.0285 | 184.668 | 4.0365 | 6.891
BLK_NACL_III.D | 2024-02-20 20:39 | 4.0306 | 111.155 | 4.0245 | 5.1985
BLK_IS_DIW_III.D | 2024-02-20 20:49 | 4.0327 | 280.5215 | 4.0347 | 18105.325
BLK_IS_NACL_III.D | 2024-02-20 20:59 | 4.0306 | 753.092 | 4.0327 | 7720.7578
AB1_III.D | 2024-02-20 21:09 | 4.0284 | 9085.1699 | 4.0284 | 18438.2365 |  |  | 0.49274 | 0.49274 | 0.49274 | 0.010381
AB2A_III.D | 2024-02-20 21:19 | 4.0266 | 18625.2578 | 4.0266 | 18857.3245 |  |  | 0.98769 | 0.98769 | 0.98769 | 0.020809
AB2B_III.D | 2024-02-20 21:29 | 4.0345 | 17044.1433 | 4.0325 | 17815.784 |  |  | 0.95669 | 0.95669 | 0.95669 | 0.020156
AB3_III.D | 2024-02-20 21:39 | 4.0326 | 33899.5653 | 4.0326 | 18113.0231 |  |  | 1.8716 | 1.8716 | 1.8716 | 0.039431
001_A1_III.D | 2024-02-20 21:49 | 4.0286 | 9432.7143 | 4.0266 | 8924.731 |  | 9399.837 | 1.0569 | 1.0532 | 1.0551 | 0.0095144
078_A1_III.D | 2024-02-20 21:59 | 4.0286 | 8855.3124 | 4.0286 | 8710.924 |  | 8822.4352 | 1.0166 | 1.0128 | 1.0147 | 0.0091733
155_A1_III.D | 2024-02-20 22:09 | 4.0306 | 8797.988 | 4.0306 | 8788.669 |  | 8765.1107 | 1.0011 | 0.99732 | 0.99919 | 0.0090357
232_A1_III.D | 2024-02-20 22:19 | 4.0327 | 8791.9585 | 4.0327 | 8759.2593 |  | 8759.0813 | 1.0037 | 0.99998 | 1.0019 | 0.0090601
310_A1_III.D | 2024-02-20 22:30 | 4.0306 | 8495.8785 | 4.0327 | 8459.6173 |  | 8463.0012 | 1.0043 | 1.0004 | 1.0023 | 0.0090783
001_A2_III.D | 2024-02-20 22:40 | 4.0345 | 8134.8926 | 4.0345 | 8061.8475 |  | 8102.0154 | 1.0091 | 1.005 | 1.007 | 0.0091396
078_A2_III.D | 2024-02-20 22:50 | 4.0325 | 8921.567 | 4.0345 | 8927.4425 |  | 8888.6898 | 0.99934 | 0.99566 | 0.9975 | 0.0090151
155_A2_III.D | 2024-02-20 23:00 | 4.0285 | 9124.5865 | 4.0285 | 9062.1891 |  | 9091.7093 | 1.0069 | 1.0033 | 1.0051 | 0.0090751
232_A2_III.D | 2024-02-20 23:10 | 4.0286 | 8815.8895 | 4.0286 | 8807.5765 |  | 8783.0123 | 1.0009 | 0.99721 | 0.99908 | 0.0090339
310_A2_III.D | 2024-02-20 23:20 | 4.0285 | 9643.4359 | 4.0285 | 8993.5312 |  | 9610.5586 | 1.0723 | 1.0686 | 1.0704 | 0.0096451
SALT1_113_III.D | 2024-02-20 23:30 | 4.0325 | 9061.7708 | 4.0325 | 8243.7972 |  | 9027.4803 | 1.0992 | 1.0951 | 1.0971 | 0.0099269
MOOS3_III.D | 2024-02-20 23:40 | 4.0327 | 10304.1515 | 4.0347 | 7842.7995 |  | 10267.9416 | 1.3138 | 1.3092 | 1.3115 | 0.011831
AB2C_III.D | 2024-02-20 23:50 | 4.0306 | 17552.5148 | 4.0306 | 17503.854 |  |  | 1.0028 | 1.0028 | 1.0028 | 0.021127
Evaluation of potential signal that can be due to preformed EtONO2 | [7 empty] | RSD | RSD | RSD
α is the signal suppression due to matrix Signal (matrix) = α * Signal (standard) | [7 empty] | 0.09491% | 0.10319 | 0.098579
BLK_1 = A0 + S(EtONO) | [7 empty] | 0.04075% | 0.098926 | 0.095382
BLK_2 = α * A0 + S(EtONO) | [7 empty] | 0.04666% | 0.097954 | 0.099644
BLK_1 - BLK_2 = A0 * ( 1 - α ) | [7 empty] | 0.09677% | 0.097678
S(matrix) = α * S(standard) | [7 empty] | 0.2952% | 0.09787
A0 is the nitrate blank present as NO3 at the moment of the derivatization | [7 empty] | 0.01454% | 0.09805
[8 empty] | 0.2127% | 0.097641
α → | 0.42318 | 1.522% | A0 | S(EtONO2) |  |  |  | 0.1724% | 0.098571
BLK 1-2 | 67.1187 | 8.311% | 116.3601 | 63.6517 |  |  |  | 0.1715% | 0.098043
BLK 3-4 | 0 | n/a | 0 | 272.7853 |  |  |  | 0.1945% | 0.10466
0.2276%
Average (BLK EtONO2 in water) |  |  |  | 63.6517 |  |  |  | 0.09533%
0.1351%
Correction of the sample signals: MOOS-4 SAMPLES | [7 empty] | 0.1482%
0.1601%
α → | 0.48245 | 0.48581 | 0.48219 | 0.48348 |  |  |  | 0.2703%
Average (BLK EtONO2 in sample) |  |  |  | 30.7745 |  |  |  | 0.1349% |  | 2.266%
46 area correction for samples |  |  |  | -32.8772 |  |  |  | 0.08557% | 0.8847% | 2.107%
Relevance of the blank |  |  |  | -0.3789% |  |  |  |  | MOOS | STANDARD
Correction of the sample signals: SALT-1 SAMPLE
α → | 0.46819 | 0.46134 | 0.45431 | 0.46128
Average (BLK EtONO2 in sample) |  |  |  | 29.3612
46 area correction for samples |  |  |  | -34.2905
Relevance of the blank |  |  |  | -0.3836%
Correction of the sample signals: MOOS-3 SAMPLE
α → | 0.42259 | 0.43857 | 0.43221 | 0.43112
Average (BLK EtONO2 in sample) |  |  |  | 27.4418
46 area correction for samples |  |  |  | -36.2099
Relevance of the blank |  |  |  | -0.3623%
2024-02-21 CHAR MOOS4 NO2 w1d2 - GCMS (verified on 2024-04-09)
Sample |  | Nitrite_31 Results |  | Nitrite_32 Results |  | Isotope ratio
Data File | Acq. Date-Time | RT | Area | RT | Area | rAB | u[rAB]
VOID_I.D | 2024-02-21 10:42
VOIDX.D | 2024-02-21 10:55
VOIDXI.D | 2024-02-21 11:09
BLK_DIW_I.D | 2024-02-21 11:23 | 2.4567 | 5.472
BLK_NACL_I.D | 2024-02-21 11:37 | 2.4568 | 60.0865
BLK_IS_DIW_I.D | 2024-02-21 11:51 | 2.4508 | 36.7923 | 2.4508 | 2753.8684
BLK_IS_NACL_I.D | 2024-02-21 12:04 | 2.4566 | 107.597 | 2.4587 | 2650.9543 |  |  |  | Sample ID | rAB | u[rAB] | Date and Time
AB1_I.D | 2024-02-21 12:18 | 2.4546 | 1357.5246 | 2.4566 | 2772.9864 | 0.48955 | 0.0013719 |  | AB1 | 0.491 | 0.001376 | 2024-02-21 12:18
AB2A_I.D | 2024-02-21 12:32 | 2.4668 | 2624.1745 | 2.4648 | 2713.2169 | 0.96718 | 0.0027104 |  | AB2A | 0.9635 | 0.0027001 | 2024-02-21 12:32
AB2B_I.D | 2024-02-21 12:46 | 2.4568 | 2747.4528 | 2.4568 | 2833.1055 | 0.96977 | 0.0027177 |  | AB2B | 0.96822 | 0.0027133 | 2024-02-21 12:46
AB3_I.D | 2024-02-21 13:00 | 2.4648 | 5326.5344 | 2.4648 | 2828.1439 | 1.8834 | 0.005278 |  | AB3 | 1.8787 | 0.0052648 | 2024-02-21 13:00
001_A1_I.D | 2024-02-21 13:14 | 2.4608 | 2313.6412 | 2.4628 | 2474.3455 | 0.93505 | 0.0026204 |  | 001_A1 | 0.93544 | 0.0026215 | 2024-02-21 13:14
078_A1_I.D | 2024-02-21 13:28 | 2.4608 | 2289.0586 | 2.4628 | 2449.182 | 0.93462 | 0.0026192 |  | 078_A1 | 0.93369 | 0.0026166 | 2024-02-21 13:28
155_A1_I.D | 2024-02-21 13:42 | 2.4488 | 2428.4961 | 2.4508 | 2628.919 | 0.92376 | 0.0025887 |  | 155_A1 | 0.92497 | 0.0025921 | 2024-02-21 13:42
232_A1_I.D | 2024-02-21 13:56 | 2.465 | 2269.0478 | 2.465 | 2438.152 | 0.93064 | 0.002608 |  | 232_A1 | 0.93336 | 0.0026156 | 2024-02-21 13:56
310_A1_I.D | 2024-02-21 14:09 | 2.4546 | 2458.6167 | 2.4587 | 2652.1825 | 0.92702 | 0.0025979 |  | 310_A1 | 0.92931 | 0.0026043 | 2024-02-21 14:09
001_A2_I.D | 2024-02-21 14:23 | 2.4507 | 2373.4034 | 2.4487 | 2545.008 | 0.93257 | 0.0026134 |  | 001_A2 | 0.92897 | 0.0026033 | 2024-02-21 14:23
078_A2_I.D | 2024-02-21 14:37 | 2.4486 | 2380.7468 | 2.4486 | 2567.511 | 0.92726 | 0.0025985 |  | 078_A2 | 0.9285 | 0.002602 | 2024-02-21 14:37
155_A2_I.D | 2024-02-21 14:51 | 2.4506 | 2377.659 | 2.4506 | 2557.599 | 0.92964 | 0.0026052 |  | 155_A2 | 0.92999 | 0.0026062 | 2024-02-21 14:51
232_A2_I.D | 2024-02-21 15:05 | 2.4647 | 2271.1222 | 2.4647 | 2464.5204 | 0.92153 | 0.0025825 |  | 232_A2 | 0.92417 | 0.0025899 | 2024-02-21 15:05
310_A2_I.D | 2024-02-21 15:19 | 2.4668 | 2170.4083 | 2.4647 | 2338.3046 | 0.9282 | 0.0026012 |  | 310_A2 | 0.92699 | 0.0025978 | 2024-02-21 15:19
AB2C_I.D | 2024-02-21 15:33 | 2.4549 | 2593.0738 | 2.4589 | 2685.4315 | 0.96561 | 0.002706 |  | AB2C | 0.9674 | 0.002711 | 2024-02-21 15:33
MOOS3_I.D | 2024-02-21 15:47 | 2.4609 | 4257.687 | 2.4609 | 2317.636 | 1.8371 | 0.0051482 |  | MOOS3 | 1.8423 | 0.0051628 | 2024-02-21 15:47
AB4_I.D | 2024-02-21 16:01 | 2.4547 | 8551.9483 | 2.4547 | 2991.1092 | 2.8591 | 0.0080124 |  | AB4 | 2.8467 | 0.0079775 | 2024-02-21 16:01
VOID_II.D | 2024-02-21 16:15
BLK_DIW_II.D | 2024-02-21 16:29 | 2.4567 | 10.2971
BLK_NACL_II.D | 2024-02-21 16:43 | 2.4548 | 13.4486
BLK_IS_DIW_II.D | 2024-02-21 16:56 | 2.4589 | 34.1817 | 2.463 | 2225.877
BLK_IS_NACL_II.D | 2024-02-21 17:10 | 2.4486 | 121.5038 | 2.4486 | 2752.0005
AB1_II.D | 2024-02-21 17:24 | 2.4547 | 1306.0185 | 2.4547 | 2657.4315 | 0.49146 | 0.0013773
AB2A_II.D | 2024-02-21 17:38 | 2.4587 | 2628.9562 | 2.4607 | 2727.8482 | 0.96375 | 0.0027008
AB2B_II.D | 2024-02-21 17:52 | 2.4569 | 2656.8307 | 2.4569 | 2738.5479 | 0.97016 | 0.0027188
AB3_II.D | 2024-02-21 18:06 | 2.4566 | 5483.299 | 2.4566 | 2923.166 | 1.8758 | 0.0052567
001_A1_II.D | 2024-02-21 18:20 | 2.4589 | 2309.3627 | 2.4609 | 2476.451 | 0.93253 | 0.0026133
078_A1_II.D | 2024-02-21 18:34 | 2.4607 | 2292.6486 | 2.4628 | 2450.171 | 0.93571 | 0.0026222
155_A1_II.D | 2024-02-21 18:48 | 2.463 | 2301.9677 | 2.465 | 2487.166 | 0.92554 | 0.0025937
232_A1_II.D | 2024-02-21 19:02 | 2.4506 | 2394.7682 | 2.4527 | 2562.4205 | 0.93457 | 0.002619
310_A1_II.D | 2024-02-21 19:15 | 2.4486 | 2419.9531 | 2.4486 | 2604.502 | 0.92914 | 0.0026038
001_A2_II.D | 2024-02-21 19:30 | 2.4467 | 2333.6788 | 2.4487 | 2536.9035 | 0.91989 | 0.0025779
078_A2_II.D | 2024-02-21 19:43 | 2.4649 | 2189.3901 | 2.4669 | 2348.7705 | 0.93214 | 0.0026122
155_A2_II.D | 2024-02-21 19:57 | 2.4649 | 2229.9902 | 2.4649 | 2397.598 | 0.93009 | 0.0026065
232_A2_II.D | 2024-02-21 20:11 | 2.4567 | 2358.6624 | 2.4587 | 2543.372 | 0.92738 | 0.0025989
310_A2_II.D | 2024-02-21 20:25 | 2.4607 | 2223.296 | 2.4587 | 2391.4055 | 0.9297 | 0.0026054
AB2C_II.D | 2024-02-21 20:39 | 2.4547 | 2596.703 | 2.4547 | 2685.3 | 0.96701 | 0.0027099
MOOS3_II.D | 2024-02-21 20:53 | 2.4609 | 4286.0517 | 2.4609 | 2324.993 | 1.8435 | 0.0051661
AB4_II.D | 2024-02-21 21:07 | 2.4648 | 8107.1701 | 2.4648 | 2861.078 | 2.8336 | 0.0079409
VOID_III.D | 2024-02-21 21:21
BLK_DIW_III.D | 2024-02-21 21:35 | 2.4549 | 8.7035
BLK_NACL_III.D | 2024-02-21 21:48 | 2.4508 | 5.5595
BLK_IS_DIW_III.D | 2024-02-21 22:02 | 2.4446 | 36.4505 | 2.4567 | 2268.7195
BLK_IS_NACL_III.D | 2024-02-21 22:16 | 2.4589 | 111.7125 | 2.4589 | 2659.4185
AB1_III.D | 2024-02-21 22:30 | 2.4549 | 1308.1494 | 2.4549 | 2658.927 | 0.49198 | 0.0013787
AB2A_III.D | 2024-02-21 22:44 | 2.4648 | 2560.8882 | 2.4648 | 2668.797 | 0.95957 | 0.0026891
AB2B_III.D | 2024-02-21 22:58 | 2.4527 | 2706.7442 | 2.4506 | 2805.7125 | 0.96473 | 0.0027035
AB3_III.D | 2024-02-21 23:12 | 2.4587 | 5437.8068 | 2.4587 | 2897.3975 | 1.8768 | 0.0052595
001_A1_III.D | 2024-02-21 23:26 | 2.4507 | 2265.407 | 2.4507 | 2413.2475 | 0.93874 | 0.0026307
078_A1_III.D | 2024-02-21 23:40 | 2.4548 | 2233.8924 | 2.4548 | 2400.144 | 0.93073 | 0.0026083
155_A1_III.D | 2024-02-21 23:54 | 2.4509 | 2319.2666 | 2.4509 | 2505.6615 | 0.92561 | 0.0025939
232_A1_III.D | 2024-02-22 00:08 | 2.4487 | 2280.3021 | 2.4507 | 2439.212 | 0.93485 | 0.0026198
310_A1_III.D | 2024-02-22 00:21 | 2.4588 | 2238.9915 | 2.4588 | 2402.952 | 0.93177 | 0.0026112
001_A2_III.D | 2024-02-22 00:35 | 2.4488 | 2243.7806 | 2.4508 | 2401.158 | 0.93446 | 0.0026187
078_A2_III.D | 2024-02-22 00:49 | 2.4508 | 2201.2306 | 2.4508 | 2376.9225 | 0.92608 | 0.0025952
155_A2_III.D | 2024-02-22 01:03 | 2.4508 | 2290.8125 | 2.4528 | 2462.6325 | 0.93023 | 0.0026069
232_A2_III.D | 2024-02-22 01:17 | 2.4568 | 2263.8833 | 2.4608 | 2451.1395 | 0.9236 | 0.0025883
310_A2_III.D | 2024-02-22 01:31 | 2.4587 | 2113.8586 | 2.4607 | 2290.008 | 0.92308 | 0.0025868
AB2C_III.D | 2024-02-22 01:45 | 2.4568 | 2631.6116 | 2.4588 | 2714.1528 | 0.96959 | 0.0027172
MOOS3_III.D | 2024-02-22 01:59 | 2.4507 | 4438.9497 | 2.4507 | 2404.1695 | 1.8464 | 0.0051742
AB4_III.D | 2024-02-22 02:13 | 2.4608 | 8348.2218 | 2.4608 | 2931.9561 | 2.8473 | 0.0079793
Effect of blank | [5 empty] | RSD | RSD
0.2345% | [5 empty] | 0.2605% | 0.091253
[6 empty] | 0.3958% | 0.091036
[6 empty] | 0.313% | 0.090471
[6 empty] | 0.2199% | 0.091022
[6 empty] | 0.3338% | 0.090922
[6 empty] | 0.2802% | 0.090597
[6 empty] | 0.1132% | 0.090716
[6 empty] | 0.2522% | 0.091143
[6 empty] | 0.2561% | 0.090562
[6 empty] | 0.8527% | 0.090765
0.3461%
0.03288%
0.3208%
0.3746%
0.2088%
0.2576%
0.4486%
[6 empty] | 0.2802% | 0.2921%
2024-02-20 CHAR MOOS4 NO2 w1d1 - spectrophotometry (verified on 2024-04-09)
Sample ID | Date | 541nm (Abs)
BLS-DIW | 2/20/2024 3:31:31 PM | -0.000252
BSL-LNWS | 2/20/2024 3:32:18 PM | -0.000417
BSL-M4 | 2/20/2024 3:33:13 PM | -0.000417
BSL-M3 | 2/20/2024 3:33:58 PM | -0.000377
BLK-DIW | 2/20/2024 3:34:54 PM | -0.000218
BLK-LNSW | 2/20/2024 3:35:48 PM | 2.3e-05
CAL1 | 2/20/2024 3:37:24 PM | 0.045523
CAL2 | 2/20/2024 3:38:23 PM | 0.064783
CAL3 | 2/20/2024 3:39:20 PM | 0.086985
CAL4 | 2/20/2024 3:40:15 PM | 0.1271
CAL5 | 2/20/2024 3:41:06 PM | 0.17165
001a_1 | 2/20/2024 3:42:46 PM | 0.083905
078a_1 | 2/20/2024 3:43:32 PM | 0.083475
155a_1 | 2/20/2024 3:44:12 PM | 0.083304
232a_1 | 2/20/2024 3:44:55 PM | 0.083229
310a_1 | 2/20/2024 3:45:37 PM | 0.084108
001a_2 | 2/20/2024 3:46:18 PM | 0.083172
078a_2 | 2/20/2024 3:46:59 PM | 0.083875
155a_2 | 2/20/2024 3:47:41 PM | 0.083345
232a_2 | 2/20/2024 3:48:24 PM | 0.0831
310a_2 | 2/20/2024 3:49:07 PM | 0.083043
MOOS3 | 2/20/2024 3:50:04 PM | 0.1634
2024-02-20 CHAR MOOS4 PO4 w1d1 - spectrophotometry (verified on 2024-04-09)
Sample ID | Date | 890nm (Abs)
BLS-DIW | 2/20/2024 5:17:55 PM | -0.000594
BSL-LNWS | 2/20/2024 5:18:37 PM | -0.00262
BSL-M4 | 2/20/2024 5:19:59 PM | -0.001955
BSL-SALT | 2/20/2024 5:20:43 PM | 0.000197
BSL-M3 | 2/20/2024 5:21:31 PM | -0.002503
BLK-DIW | 2/20/2024 5:22:30 PM | -0.000605
BLK-LNSW | 2/20/2024 5:23:42 PM | -0.002192
CAL1 | 2/20/2024 5:24:54 PM | 0.1337
CAL2 | 2/20/2024 5:25:55 PM | 0.20198
CAL3 | 2/20/2024 5:26:51 PM | 0.26679
CAL4 | 2/20/2024 5:28:05 PM | 0.3908
CAL5 | 2/20/2024 5:29:08 PM | 0.53541
001a_1 | 2/20/2024 5:30:41 PM | 0.27373
078a_1 | 2/20/2024 5:31:27 PM | 0.27549
155a_1 | 2/20/2024 5:32:15 PM | 0.27478
232a_1 | 2/20/2024 5:33:00 PM | 0.27465
310a_1 | 2/20/2024 5:33:44 PM | 0.27584
001a_2 | 2/20/2024 5:34:23 PM | 0.27459
078a_2 | 2/20/2024 5:35:06 PM | 0.27548
155a_2 | 2/20/2024 5:36:02 PM | 0.2747
232a_2 | 2/20/2024 5:36:55 PM | 0.27456
310a_2 | 2/20/2024 5:37:45 PM | 0.27196
SALT1_113 | 2/20/2024 5:39:17 PM | 0.15677
MOOS3 | 2/20/2024 5:40:31 PM | 0.15339
2024-02-21 CHAR MOOS4 SiO2 w1d2 - spectrophotometry (verified on 2024-04-09)
Sample ID | Date | 810nm (Abs)
BLS-DIW | 2/21/2024 11:43:54 AM | 0.000828
BSL-LNWS | 2/21/2024 11:44:34 AM | -0.000282
BSL-M4 | 2/21/2024 11:45:09 AM | -0.000234
BSL-SALT | 2/21/2024 11:45:45 AM | 6e-06
BSL-M3 | 2/21/2024 11:46:19 AM | 3.8e-05
BLK-DIW | 2/21/2024 11:47:01 AM | 0.000217
BLK-LNSW | 2/21/2024 11:47:41 AM | 0.011604
CAL1 | 2/21/2024 11:48:24 AM | 0.071961
CAL2 | 2/21/2024 11:49:04 AM | 0.10009
CAL3 | 2/21/2024 11:49:47 AM | 0.12859
CAL4 | 2/21/2024 11:50:34 AM | 0.18786
CAL5 | 2/21/2024 11:51:30 AM | 0.24244
001a_1 | 2/21/2024 11:52:29 AM | 0.11764
078a_1 | 2/21/2024 11:53:08 AM | 0.11819
155a_1 | 2/21/2024 11:53:53 AM | 0.11924
232a_1 | 2/21/2024 11:54:31 AM | 0.11856
310a_1 | 2/21/2024 11:55:10 AM | 0.11979
001a_2 | 2/21/2024 11:55:49 AM | 0.11821
078a_2 | 2/21/2024 11:56:26 AM | 0.11983
155a_2 | 2/21/2024 11:57:02 AM | 0.11837
232a_2 | 2/21/2024 11:57:35 AM | 0.11974
310a_2 | 2/21/2024 11:58:10 AM | 0.11927
SALT1_113 | 2/21/2024 11:59:02 AM | 0.1498
MOOS3 | 2/21/2024 11:59:42 AM | 0.49205
CAL6 | 2/21/2024 12:00:27 PM | 0.66193
2024-02-21 CHAR MOOS4 NO2 w1d2 - spectrophotometry (verified on 2024-04-09)
Sample ID | Date | 541nm (Abs)
BLS-DIW | 2/21/2024 1:35:24 PM | 1.2e-05
BSL-LNWS | 2/21/2024 1:36:29 PM | -8.7e-05
BSL-M4 | 2/21/2024 1:37:26 PM | -0.000158
BSL-M3 | 2/21/2024 1:38:07 PM | 4.3e-05
BLK-DIW | 2/21/2024 1:38:43 PM | 9.1e-05
BLK-LNSW | 2/21/2024 1:39:25 PM | 5e-06
CAL1 | 2/21/2024 1:40:11 PM | 0.044309
CAL2 | 2/21/2024 1:40:59 PM | 0.065183
CAL3 | 2/21/2024 1:41:49 PM | 0.087086
CAL4 | 2/21/2024 1:42:41 PM | 0.13109
CAL5 | 2/21/2024 1:43:29 PM | 0.2008
001a_1 | 2/21/2024 1:44:41 PM | 0.084117
078a_1 | 2/21/2024 1:45:23 PM | 0.08412
155a_1 | 2/21/2024 1:46:03 PM | 0.084168
232a_1 | 2/21/2024 1:46:44 PM | 0.084109
310a_1 | 2/21/2024 1:47:23 PM | 0.083853
MOOS3 | 2/21/2024 1:48:24 PM | 0.16705
2024-03-05 CHAR MOOS4 NO3 w2d1 - GCMS (verified on 2024-04-09)
Sample |  | Nitrate_46 Results |  | Nitrate_47 Results |  | Nitrate_46 blk corr |  | BLK no | BLK yes | Isotope ratio
Data File | Acq. Date-Time | RT | Area | RT | Area | RT | Area | rAB | rAB | rAB | u[rAB]
VOID_I.D | 2024-03-05 11:58
MOOS4_NO_DERV.D | 2024-03-05 12:08 |  | 0
BLK_DIW_I.D | 2024-03-05 12:18 | 4.0347 | 20.1428
BLK_NACL_I.D | 2024-03-05 12:28 | 4.0345 | 66.8515
BLK_IS_DIW_I.D | 2024-03-05 12:38 | 4.0305 | 483.076 | 4.0305 | 18184.7255
BLK_IS_NACL_I.D | 2024-03-05 12:48 | 4.0347 | 137.9721 | 4.0327 | 7788.9873 | [7 empty] | Sample ID | rAB | u[rAB] | Date and Time
AB1_I.D | 2024-03-05 12:58 | 4.0326 | 9368.3376 | 4.0326 | 17759.5463 |  |  | 0.52751 | 0.52751 | 0.52751 | 0.011114 |  | AB1 | 0.5277 | 0.011118 | 2024-03-05 12:58
AB2A_I.D | 2024-03-05 13:08 | 4.0347 | 17065.936 | 4.0347 | 17300.6997 |  |  | 0.98643 | 0.98643 | 0.98643 | 0.020783 |  | AB2A | 0.98614 | 0.020777 | 2024-03-05 13:08
AB2B_I.D | 2024-03-05 13:18 | 4.0306 | 17454.8887 | 4.0306 | 17550.6675 |  |  | 0.99454 | 0.99454 | 0.99454 | 0.020954 |  | AB2B | 0.99424 | 0.020947 | 2024-03-05 13:18
AB3_I.D | 2024-03-05 13:28 | 4.0306 | 33779.9181 | 4.0306 | 17404.317 |  |  | 1.9409 | 1.9409 | 1.9409 | 0.040892 |  | AB3 | 1.9432 | 0.04094 | 2024-03-05 13:28
052_A1_I.D | 2024-03-05 13:38 | 4.0346 | 8515.9152 | 4.0346 | 8356.201 |  | 8503.5284 | 1.0191 | 1.0176 | 1.0184 | 0.0090401 |  | 052_A1 | 1.0189 | 0.0090444 | 2024-03-05 13:38
130_A1_I.D | 2024-03-05 13:48 | 4.0286 | 9276.1209 | 4.0286 | 9241.117 |  | 9263.7341 | 1.0038 | 1.0024 | 1.0031 | 0.0089 |  | 130_A1 | 1.0023 | 0.0088959 | 2024-03-05 13:48
207_A1_I.D | 2024-03-05 13:58 | 4.0327 | 8269.5835 | 4.0327 | 8328.4577 |  | 8257.1967 | 0.99293 | 0.99144 | 0.99219 | 0.0088095 |  | 207_A1 | 0.99549 | 0.0088377 | 2024-03-05 13:58
284_A1_I.D | 2024-03-05 14:08 | 4.0287 | 8555.05 | 4.0287 | 8438.338 |  | 8542.6631 | 1.0138 | 1.0124 | 1.0131 | 0.008993 |  | 284_A1 | 1.0113 | 0.0089772 | 2024-03-05 14:08
360_A1_I.D | 2024-03-05 14:18 | 4.0286 | 8291.049 | 4.0306 | 8278.7795 |  | 8278.6622 | 1.0015 | 0.99999 | 1.0007 | 0.0088852 |  | 360_A1 | 1.0015 | 0.0088916 | 2024-03-05 14:18
052_A2_I.D | 2024-03-05 14:28 | 4.0326 | 8490.5817 | 4.0326 | 8327.3685 |  | 8478.1949 | 1.0196 | 1.0181 | 1.0189 | 0.0090446 |  | 052_A2 | 1.0212 | 0.0090647 | 2024-03-05 14:28
130_A2_I.D | 2024-03-05 14:38 | 4.0326 | 8438.769 | 4.0326 | 8233.5745 |  | 8426.3822 | 1.0249 | 1.0234 | 1.0242 | 0.0090921 |  | 130_A2 | 1.0254 | 0.0091022 | 2024-03-05 14:38
207_A2_I.D | 2024-03-05 14:49 | 4.0347 | 8399.1765 | 4.0347 | 8318.0753 |  | 8386.7897 | 1.0097 | 1.0083 | 1.009 | 0.0089578 |  | 207_A2 | 1.009 | 0.0089556 | 2024-03-05 14:49
284_A2_I.D | 2024-03-05 14:59 | 4.0325 | 8413.3935 | 4.0305 | 8455.0871 |  | 8401.0067 | 0.99507 | 0.9936 | 0.99434 | 0.0088275 |  | 284_A2 | 0.9937 | 0.0088201 | 2024-03-05 14:59
360_A2_I.D | 2024-03-05 15:09 | 4.0286 | 8266.963 | 4.0306 | 8270.0722 |  | 8254.5762 | 0.99962 | 0.99813 | 0.99888 | 0.0088689 |  | 360_A2 | 0.99699 | 0.0088513 | 2024-03-05 15:09
SALT1_309_I.D | 2024-03-05 15:19 | 4.0306 | 8740.0863 | 4.0327 | 7750.537 |  | 8726.8697 | 1.1277 | 1.126 | 1.1268 | 0.010006 |  | SALT1_309 | 1.1249 | 0.0099878 | 2024-03-05 15:19
MOOS3_I.D | 2024-03-05 15:29 | 4.0325 | 10688.542 | 4.0325 | 8008.5015 |  | 10675.4145 | 1.3346 | 1.333 | 1.3338 | 0.011829 |  | MOOS3 | 1.3327 | 0.01182 | 2024-03-05 15:29
AB2C_I.D | 2024-03-05 15:39 | 4.0347 | 17665.498 | 4.0347 | 17211.6165 |  |  | 1.0264 | 1.0264 | 1.0264 | 0.021624 |  | AB2C | 1.0272 | 0.021641 | 2024-03-05 15:39
VOID_II.D | 2024-03-05 15:49 | 4.0367 | 8.341
BLK_DIW_II.D | 2024-03-05 15:59 | 4.0327 | 25.7505
BLK_NACL_II.D | 2024-03-05 16:09 | 4.0265 | 66.972
BLK_IS_DIW_II.D | 2024-03-05 16:19 | 4.0366 | 470.079 | 4.0346 | 17319.0985
BLK_IS_NACL_II.D | 2024-03-05 16:29 | 4.0347 | 144.5865 | 4.0347 | 7771.6199
AB1_II.D | 2024-03-05 16:39 | 4.0347 | 9210.556 | 4.0347 | 17453.7505 |  |  | 0.52771 | 0.52771 | 0.52771 | 0.011118
AB2A_II.D | 2024-03-05 16:49 | 4.0285 | 17321.0183 | 4.0285 | 17563.4248 |  |  | 0.9862 | 0.9862 | 0.9862 | 0.020778
AB2B_II.D | 2024-03-05 16:59 | 4.0327 | 17448.881 | 4.0327 | 17535.4012 |  |  | 0.99507 | 0.99507 | 0.99507 | 0.020965
AB3_II.D | 2024-03-05 17:09 | 4.0325 | 34043.4565 | 4.0325 | 17501.8205 |  |  | 1.9451 | 1.9451 | 1.9451 | 0.040982
052_A1_II.D | 2024-03-05 17:19 | 4.0347 | 8670.5169 | 4.0347 | 8505.977 |  | 8658.13 | 1.0193 | 1.0179 | 1.0186 | 0.0090412
130_A1_II.D | 2024-03-05 17:29 | 4.0287 | 8779.443 | 4.0307 | 8745.5925 |  | 8767.0562 | 1.0039 | 1.0025 | 1.0032 | 0.0089033
207_A1_II.D | 2024-03-05 17:39 | 4.0305 | 8437.987 | 4.0305 | 8452.9473 |  | 8425.6002 | 0.99823 | 0.99676 | 0.9975 | 0.0088554
284_A1_II.D | 2024-03-05 17:49 | 4.0365 | 8474.7311 | 4.0365 | 8390.5895 |  | 8462.3442 | 1.01 | 1.0086 | 1.0093 | 0.0089598
360_A1_II.D | 2024-03-05 17:59 | 4.0347 | 8179.7605 | 4.0347 | 8141.7255 |  | 8167.3737 | 1.0047 | 1.0032 | 1.0039 | 0.0089143
052_A2_II.D | 2024-03-05 18:09 | 4.0326 | 8466.083 | 4.0326 | 8271.3035 |  | 8453.6962 | 1.0235 | 1.0221 | 1.0228 | 0.0090798
130_A2_II.D | 2024-03-05 18:19 | 4.0345 | 8488.502 | 4.0325 | 8266.9015 |  | 8476.1152 | 1.0268 | 1.0253 | 1.0261 | 0.0091085
207_A2_II.D | 2024-03-05 18:29 | 4.0285 | 8622.497 | 4.0305 | 8534.6583 |  | 8610.1102 | 1.0103 | 1.0088 | 1.0096 | 0.0089612
284_A2_II.D | 2024-03-05 18:40 | 4.0286 | 8717.1465 | 4.0286 | 8752.4331 |  | 8704.7597 | 0.99597 | 0.99455 | 0.99526 | 0.0088336
360_A2_II.D | 2024-03-05 18:50 | 4.0325 | 8339.2586 | 4.0325 | 8385.513 |  | 8326.8717 | 0.99448 | 0.99301 | 0.99375 | 0.0088228
SALT1_309_II.D | 2024-03-05 19:00 | 4.0287 | 8826.274 | 4.0287 | 7837.1785 |  | 8813.0574 | 1.1262 | 1.1245 | 1.1254 | 0.0099919
MOOS3_II.D | 2024-03-05 19:10 | 4.0346 | 10462.808 | 4.0346 | 7846.644 |  | 10449.6805 | 1.3334 | 1.3317 | 1.3326 | 0.011819
AB2C_II.D | 2024-03-05 19:20 | 4.0306 | 17953.8155 | 4.0306 | 17463.898 |  |  | 1.0281 | 1.0281 | 1.0281 | 0.02166
VOID_III.D | 2024-03-05 19:30 | 4.0386 | 10.1917
BLK_DIW_III.D | 2024-03-05 19:40 | 4.0305 | 26.1132
BLK_NACL_III.D | 2024-03-05 19:50 | 4.0327 | 67.335
BLK_IS_DIW_III.D | 2024-03-05 20:00 | 4.0367 | 485.495 | 4.0347 | 17753.1074
BLK_IS_NACL_III.D | 2024-03-05 20:10 | 4.0285 | 147.369 | 4.0305 | 8036.188
AB1_III.D | 2024-03-05 20:20 | 4.0346 | 9252.397 | 4.0346 | 17527.2298 |  |  | 0.52789 | 0.52789 | 0.52789 | 0.011122
AB2A_III.D | 2024-03-05 20:30 | 4.0305 | 17113.5845 | 4.0305 | 17360.2629 |  |  | 0.98579 | 0.98579 | 0.98579 | 0.020769
AB2B_III.D | 2024-03-05 20:40 | 4.0346 | 17179.5494 | 4.0346 | 17298.63 |  |  | 0.99312 | 0.99312 | 0.99312 | 0.020924
AB3_III.D | 2024-03-05 20:50 | 4.0305 | 35625.0105 | 4.0305 | 18330.6375 |  |  | 1.9435 | 1.9435 | 1.9435 | 0.040946
052_A1_III.D | 2024-03-05 21:00 | 4.0306 | 8664.036 | 4.0306 | 8489.669 |  | 8651.6492 | 1.0205 | 1.0191 | 1.0198 | 0.0090518
130_A1_III.D | 2024-03-05 21:10 | 4.0327 | 8124.263 | 4.0327 | 8113.985 |  | 8111.8761 | 1.0013 | 0.99974 | 1.0005 | 0.0088844
207_A1_III.D | 2024-03-05 21:20 | 4.0307 | 8521.6 | 4.0307 | 8542.9005 |  | 8509.2132 | 0.99751 | 0.99606 | 0.99678 | 0.0088484
284_A1_III.D | 2024-03-05 21:30 | 4.0305 | 8573.72 | 4.0305 | 8470.0615 |  | 8561.3332 | 1.0122 | 1.0108 | 1.0115 | 0.0089788
360_A1_III.D | 2024-03-05 21:40 | 4.0307 | 8519.531 | 4.0307 | 8515.037 |  | 8507.1442 | 1.0005 | 0.99907 | 0.9998 | 0.0088752
052_A2_III.D | 2024-03-05 21:50 | 4.0325 | 8760.0998 | 4.0325 | 8566.3825 |  | 8747.713 | 1.0226 | 1.0212 | 1.0219 | 0.0090697
130_A2_III.D | 2024-03-05 22:00 | 4.0307 | 8792.505 | 4.0307 | 8563.613 |  | 8780.1182 | 1.0267 | 1.0253 | 1.026 | 0.009106
207_A2_III.D | 2024-03-05 22:10 | 4.0306 | 8989.816 | 4.0306 | 8909.3962 |  | 8977.4292 | 1.009 | 1.0076 | 1.0083 | 0.0089479
284_A2_III.D | 2024-03-05 22:21 | 4.0286 | 8883.6258 | 4.0286 | 8953.4745 |  | 8871.2389 | 0.9922 | 0.99082 | 0.99151 | 0.0087992
360_A2_III.D | 2024-03-05 22:31 | 4.0327 | 8512.2206 | 4.0327 | 8520.1545 |  | 8499.8338 | 0.99907 | 0.99761 | 0.99834 | 0.0088623
SALT1_309_III.D | 2024-03-05 22:41 | 4.0286 | 9021.3766 | 4.0286 | 8030.0286 |  | 9008.16 | 1.1235 | 1.1218 | 1.1226 | 0.0099661
MOOS3_III.D | 2024-03-05 22:51 | 4.0306 | 10602.823 | 4.0306 | 7956.285 |  | 10589.6955 | 1.3326 | 1.331 | 1.3318 | 0.011812
AB2C_III.D | 2024-03-05 23:01 | 4.0305 | 17996.497 | 4.0305 | 17521.4411 |  |  | 1.0271 | 1.0271 | 1.0271 | 0.02164
ST_1.D | 2024-03-05 23:11 | 4.0346 | 20240.3973 | 4.0366 | 253.8895 |  |  |  |  | 79.7213
MOOS4_1.D | 2024-03-05 23:21 | 4.0366 | 9427.4549 | 4.0366 | 119.919 |  |  |  |  | 78.6152
ST_2.D | 2024-03-05 23:31 | 4.0325 | 22271.5691 | 4.0305 | 271.2765 |  |  |  |  | 82.0991
MOOS4_2.D | 2024-03-05 23:41 | 4.0305 | 10294.881 | 4.0325 | 126.573 |  |  |  |  | 81.3355
ST_3.D | 2024-03-05 23:51 | 4.0346 | 23408.5558 | 4.0325 | 284.563 |  |  |  |  | 82.2614
MOOS4_3.D | 2024-03-06 00:01 | 4.0305 | 10629.9735 | 4.0285 | 128.682 |  |  |  |  | 82.6065
ST_4.D | 2024-03-06 00:11 | 4.0286 | 24944.831 | 4.0286 | 301.74 |  |  |  |  | 82.67
MOOS4_4.D | 2024-03-06 00:21 | 4.0326 | 11109.473 | 4.0347 | 135.638 |  |  |  |  | 81.9053
ST_5.D | 2024-03-06 00:31 | 4.0287 | 25719.859 | 4.0287 | 314.31 |  |  |  |  | 81.8296
MOOS4_5.D | 2024-03-06 00:41 | 4.0305 | 11958.7675 | 4.0305 | 143.9205 |  |  |  |  | 83.0929
ST_6.D | 2024-03-06 00:51 | 4.0306 | 25068.9683 | 4.0286 | 304.5145 |  |  |  |  | 82.3244
MOOS4_6.D | 2024-03-06 01:01 | 4.0347 | 11772.674 | 4.0347 | 143.6695 |  |  |  |  | 81.9428
ST_7.D | 2024-03-06 01:11 | 4.0327 | 24459.3586 | 4.0307 | 299.5075 |  |  |  |  | 81.6653
MOOS4_7.D | 2024-03-06 01:21 | 4.0346 | 11657.9305 | 4.0326 | 143.2535 |  |  |  |  | 81.3797
Evaluation of potential signal that can be due to preformed EtONO2 | [7 empty] | RSD | RSD | RSD
α is the signal suppression due to matrix Signal (matrix) = α * Signal (standard) | [7 empty] | 0.03576% | 0.098925 | 0.098301
BLK_1 = A0 + S(EtONO) | [7 empty] | 0.03284% | 0.097845 | 0.098582
BLK_2 = α * A0 + S(EtONO) | [7 empty] | 0.1015% | 0.097617 | 0.10179
BLK_1 - BLK_2 = A0 * ( 1 - α ) | [7 empty] | 0.1101% | 0.099124
S(matrix) = α * S(standard) | [7 empty] | 0.07504% | 0.097856
A0 is the nitrate blank present as NO3 at the moment of the derivatization | [7 empty] | 0.1476% | 0.099901
[8 empty] | 0.2884% | 0.10069
α → | 0.44324 | 2.948% | A0 | S(EtONO2) |  |  |  | 0.1887% | 0.098851
BLK 1-2 | 0 | n/a | 0 | 24.0022 |  |  |  | 0.2165% | 0.097776
BLK 3-4 | 336.2408 | 2.956% | 603.9244 | 0 |  |  |  | 0.202% | 0.097606
0.1039%
Average (BLK EtONO2 in water) |  |  |  | 24.0022 |  |  |  | 0.06289%
0.198%
Correction of the sample signals: MOOS-4 SAMPLES | [7 empty] | 0.2827%
0.1903%
α → | 0.48292 | 0.48246 | 0.48641 | 0.48393 |  |  |  | 0.07619%
Average (BLK EtONO2 in sample) |  |  |  | 11.6153 |  |  |  | 0.08209% |  | 1.947%
46 area correction for samples |  |  |  | -12.3868 |  |  |  | 0.06354% | 0.8847% | 2.107%
Relevance of the blank |  |  |  | -0.1449% |  |  |  |  | MOOS | STANDARD
Correction of the sample signals: SALT-1 SAMPLE
α → | 0.44427 | 0.44775 | 0.45605 | 0.44936
Average (BLK EtONO2 in sample) |  |  |  | 10.7856
46 area correction for samples |  |  |  | -13.2166
Relevance of the blank |  |  |  | -0.1491%
Correction of the sample signals: MOOS-3 SAMPLE
α → | 0.45906 | 0.44829 | 0.45187 | 0.45307
Average (BLK EtONO2 in sample) |  |  |  | 10.8747
46 area correction for samples |  |  |  | -13.1275
Relevance of the blank |  |  |  | -0.124%
2024-03-05 CHAR MOOS4 PO4 w2d1 - spectrophotometry (verified on 2024-04-09)
Sample ID | Date | 890nm (Abs)
BLS-DIW | 3/5/2024 2:15:27 PM | -0.000859
BSL-LNWS | 3/5/2024 2:16:17 PM | -0.001958
BSL-M4 | 3/5/2024 2:16:59 PM | -0.001641
BSL-SALT | 3/5/2024 2:17:47 PM | 0.000148
BSL-M3 | 3/5/2024 2:18:34 PM | -0.002501
BLK-DIW | 3/5/2024 2:19:35 PM | -0.000337
BLK-LNSW | 3/5/2024 2:20:26 PM | -0.002227
CAL1 | 3/5/2024 2:21:31 PM | 0.13287
CAL2 | 3/5/2024 2:22:33 PM | 0.20334
CAL3 | 3/5/2024 2:23:35 PM | 0.26915
CAL4 | 3/5/2024 2:24:35 PM | 0.40547
CAL5 | 3/5/2024 2:25:25 PM | 0.53436
052a_1 | 3/5/2024 2:26:44 PM | 0.27287
130a_1 | 3/5/2024 2:27:38 PM | 0.27442
207a_1 | 3/5/2024 2:28:21 PM | 0.27327
284a_1 | 3/5/2024 2:29:01 PM | 0.27351
360a_1 | 3/5/2024 2:29:44 PM | 0.27404
052a_2 | 3/5/2024 2:30:29 PM | 0.27306
130a_2 | 3/5/2024 2:31:12 PM | 0.27565
207a_2 | 3/5/2024 2:32:00 PM | 0.27207
284a_2 | 3/5/2024 2:32:40 PM | 0.27345
360a_2 | 3/5/2024 2:33:28 PM | 0.27409
SALT1_309 | 3/5/2024 2:34:57 PM | 0.1569
MOOS3 | 3/5/2024 2:35:55 PM | 0.16153
2024-03-06 CHAR MOOS4 NO2 w2d2 - GCMS (verified on 2024-04-09)
Sample |  | Nitrite_31 Results |  | Nitrite_32 Results |  | Isotope ratio
Data File | Acq. Date-Time | RT | Area | RT | Area | rAB | u[rAB]
VOID_I.D | 2024-03-06 10:33
MOOS4_NO_DERV.D | 2024-03-06 10:47 |  | 0
BLK_DIW_I.D | 2024-03-06 11:01 | 2.4627 | 3.9503
BLK_NACL_I.D | 2024-03-06 11:15 | 2.4608 | 95.406
BLK_IS_DIW_I.D | 2024-03-06 11:29 | 2.4567 | 44.702 | 2.4567 | 2976.0438
BLK_IS_NACL_I.D | 2024-03-06 11:43 | 2.463 | 124.0509 | 2.465 | 2841.3794 |  |  |  | Sample ID | rAB | u[rAB] | Date and Time
AB1_I.D | 2024-03-06 11:57 | 2.4649 | 1393.909 | 2.467 | 2885.1459 | 0.48313 | 0.0014428 |  | AB1 | 0.48329 | 0.0014433 | 2024-03-06 11:57
AB2A_I.D | 2024-03-06 12:12 | 2.4588 | 2932.0292 | 2.4628 | 3110.4967 | 0.94262 | 0.0028149 |  | AB2A | 0.94403 | 0.0028191 | 2024-03-06 12:12
AB2B_I.D | 2024-03-06 12:26 | 2.4608 | 2882.3253 | 2.4648 | 3027.2642 | 0.95212 | 0.0028433 |  | AB2B | 0.94992 | 0.0028367 | 2024-03-06 12:26
AB3_I.D | 2024-03-06 12:40 | 2.4627 | 5840.5197 | 2.4627 | 3149.5442 | 1.8544 | 0.0055378 |  | AB3 | 1.8523 | 0.0055315 | 2024-03-06 12:40
052_A1_I.D | 2024-03-06 12:54 | 2.4527 | 2617.7567 | 2.4547 | 2783.8825 | 0.94033 | 0.0028081 |  | 052_A1 | 0.93791 | 0.0028009 | 2024-03-06 12:54
130_A1_I.D | 2024-03-06 13:08 | 2.4607 | 2454.4622 | 2.4627 | 2630.223 | 0.93318 | 0.0027867 |  | 130_A1 | 0.93257 | 0.0027849 | 2024-03-06 13:08
207_A1_I.D | 2024-03-06 13:23 | 2.4468 | 2621.3855 | 2.4509 | 2819.6675 | 0.92968 | 0.0027763 |  | 207_A1 | 0.92738 | 0.0027694 | 2024-03-06 13:23
284_A1_I.D | 2024-03-06 13:37 | 2.4569 | 2660.9755 | 2.4589 | 2857 | 0.93139 | 0.0027814 |  | 284_A1 | 0.93208 | 0.0027834 | 2024-03-06 13:37
360_A1_I.D | 2024-03-06 13:51 | 2.4568 | 2555.1229 | 2.4548 | 2745.8442 | 0.93054 | 0.0027789 |  | 360_A1 | 0.93057 | 0.0027789 | 2024-03-06 13:51
052_A2_I.D | 2024-03-06 14:05 | 2.4628 | 2531.978 | 2.4628 | 2722.5632 | 0.93 | 0.0027772 |  | 052_A2 | 0.93151 | 0.0027818 | 2024-03-06 14:05
130_A2_I.D | 2024-03-06 14:19 | 2.4587 | 2499.5222 | 2.4607 | 2684.4605 | 0.93111 | 0.0027805 |  | 130_A2 | 0.92812 | 0.0027716 | 2024-03-06 14:19
207_A2_I.D | 2024-03-06 14:34 | 2.4508 | 2594.4672 | 2.4508 | 2795.6745 | 0.92803 | 0.0027714 |  | 207_A2 | 0.92816 | 0.0027718 | 2024-03-06 14:34
284_A2_I.D | 2024-03-06 14:48 | 2.4528 | 2568.2673 | 2.4528 | 2798.4056 | 0.91776 | 0.0027407 |  | 284_A2 | 0.92371 | 0.0027584 | 2024-03-06 14:48
360_A2_I.D | 2024-03-06 15:02 | 2.4569 | 2550.4294 | 2.4569 | 2750.3012 | 0.92733 | 0.0027693 |  | 360_A2 | 0.93032 | 0.0027782 | 2024-03-06 15:02
AB2C_I.D | 2024-03-06 15:16 | 2.4546 | 2849.9921 | 2.4526 | 3006.1216 | 0.94806 | 0.0028312 |  | AB2C | 0.94856 | 0.0028327 | 2024-03-06 15:16
MOOS3_I.D | 2024-03-06 15:31 | 2.4506 | 4969.0687 | 2.4506 | 2636.7845 | 1.8845 | 0.0056277 |  | MOOS3 | 1.8899 | 0.0056438 | 2024-03-06 15:31
AB4_I.D | 2024-03-06 15:45 | 2.4567 | 8781.8994 | 2.4587 | 3167.0175 | 2.7729 | 0.0082807 |  | AB4 | 2.7626 | 0.0082498 | 2024-03-06 15:45
VOID_II.D | 2024-03-06 15:59
BLK_DIW_II.D | 2024-03-06 16:13 | 2.4649 | 8.0667
BLK_NACL_II.D | 2024-03-06 16:27 | 2.4648 | 67.76
BLK_IS_DIW_II.D | 2024-03-06 16:41 | 2.4567 | 38.9147 | 2.4587 | 2496.8412
BLK_IS_NACL_II.D | 2024-03-06 16:56 | 2.4548 | 126.5563 | 2.4528 | 2923.7445
AB1_II.D | 2024-03-06 17:10 | 2.4688 | 1287.9545 | 2.4688 | 2659.6496 | 0.48426 | 0.0014461
AB2A_II.D | 2024-03-06 17:24 | 2.4589 | 2836.0705 | 2.463 | 3004.6249 | 0.9439 | 0.0028188
AB2B_II.D | 2024-03-06 17:38 | 2.4608 | 2859.573 | 2.4628 | 2988.1856 | 0.95696 | 0.0028578
AB3_II.D | 2024-03-06 17:52 | 2.4609 | 5677.6305 | 2.4629 | 3077.407 | 1.8449 | 0.0055095
052_A1_II.D | 2024-03-06 18:06 | 2.4609 | 2453.3874 | 2.463 | 2609.1091 | 0.94032 | 0.002808
130_A1_II.D | 2024-03-06 18:20 | 2.4567 | 2467.746 | 2.4567 | 2641.5875 | 0.93419 | 0.0027898
207_A1_II.D | 2024-03-06 18:34 | 2.4628 | 2432.1463 | 2.4628 | 2634.3637 | 0.92324 | 0.0027571
284_A1_II.D | 2024-03-06 18:48 | 2.4607 | 2542.7923 | 2.4647 | 2733.9947 | 0.93006 | 0.0027774
360_A1_II.D | 2024-03-06 19:02 | 2.4667 | 2349.6127 | 2.4646 | 2537.143 | 0.92609 | 0.0027656
052_A2_II.D | 2024-03-06 19:15 | 2.4588 | 2497.966 | 2.4608 | 2679.474 | 0.93226 | 0.002784
130_A2_II.D | 2024-03-06 19:30 | 2.4487 | 2470.364 | 2.4487 | 2661.8983 | 0.92805 | 0.0027714
207_A2_II.D | 2024-03-06 19:43 | 2.4548 | 2498.6339 | 2.4548 | 2697.6121 | 0.92624 | 0.002766
284_A2_II.D | 2024-03-06 19:57 | 2.4486 | 2474.2794 | 2.4527 | 2677.2787 | 0.92418 | 0.0027599
360_A2_II.D | 2024-03-06 20:11 | 2.4528 | 2540.5774 | 2.4528 | 2728.5502 | 0.93111 | 0.0027806
AB2C_II.D | 2024-03-06 20:25 | 2.4549 | 2808.439 | 2.4549 | 2964.3443 | 0.94741 | 0.0028292
MOOS3_II.D | 2024-03-06 20:39 | 2.4609 | 4765.1157 | 2.4609 | 2521.9477 | 1.8895 | 0.0056425
AB4_II.D | 2024-03-06 20:53 | 2.4547 | 8781.6359 | 2.4547 | 3183.4865 | 2.7585 | 0.0082376
VOID_III.D | 2024-03-06 21:07
BLK_DIW_III.D | 2024-03-06 21:20 | 2.4548 | 7.071
BLK_NACL_III.D | 2024-03-06 21:34 | 2.4527 | 72.6701
BLK_IS_DIW_III.D | 2024-03-06 21:48 | 2.4648 | 42.375 | 2.4608 | 2456.0149
BLK_IS_NACL_III.D | 2024-03-06 22:02 | 2.4527 | 134.7972 | 2.4567 | 2857.917
AB1_III.D | 2024-03-06 22:15 | 2.4508 | 1371.5449 | 2.4488 | 2842.6128 | 0.48249 | 0.0014409
AB2A_III.D | 2024-03-06 22:29 | 2.4708 | 2662.8825 | 2.4688 | 2816.1641 | 0.94557 | 0.0028237
AB2B_III.D | 2024-03-06 22:43 | 2.4607 | 2769.5703 | 2.4607 | 2944.2435 | 0.94067 | 0.0028091
AB3_III.D | 2024-03-06 22:57 | 2.4527 | 5822.977 | 2.4548 | 3134.7899 | 1.8575 | 0.0055471
052_A1_III.D | 2024-03-06 23:11 | 2.4647 | 2332.1474 | 2.4667 | 2499.383 | 0.93309 | 0.0027865
130_A1_III.D | 2024-03-06 23:24 | 2.4629 | 2257.0194 | 2.4649 | 2425.9728 | 0.93036 | 0.0027783
207_A1_III.D | 2024-03-06 23:38 | 2.4627 | 2289.6509 | 2.4627 | 2464.0281 | 0.92923 | 0.0027749
284_A1_III.D | 2024-03-06 23:52 | 2.4487 | 2498.237 | 2.4507 | 2672.549 | 0.93478 | 0.0027915
360_A1_III.D | 2024-03-07 00:06 | 2.4505 | 2443.1889 | 2.4505 | 2612.812 | 0.93508 | 0.0027924
052_A2_III.D | 2024-03-07 00:20 | 2.4526 | 2442.7149 | 2.4526 | 2620.1364 | 0.93229 | 0.0027841
130_A2_III.D | 2024-03-07 00:33 | 2.4466 | 2362.2791 | 2.4486 | 2553.2855 | 0.92519 | 0.0027629
207_A2_III.D | 2024-03-07 00:47 | 2.4569 | 2372.6379 | 2.4569 | 2550.6261 | 0.93022 | 0.0027779
284_A2_III.D | 2024-03-07 01:01 | 2.4488 | 2387.9385 | 2.4488 | 2569.9419 | 0.92918 | 0.0027748
360_A2_III.D | 2024-03-07 01:15 | 2.4508 | 2440.7672 | 2.4508 | 2617.4167 | 0.93251 | 0.0027847
AB2C_III.D | 2024-03-07 01:28 | 2.4486 | 2831.6252 | 2.4486 | 2979.997 | 0.95021 | 0.0028376
MOOS3_III.D | 2024-03-07 01:42 | 2.4627 | 4516.8909 | 2.4647 | 2382.6056 | 1.8958 | 0.0056613
AB4_III.D | 2024-03-07 01:56 | 2.4526 | 8664.8315 | 2.4546 | 3143.7219 | 2.7562 | 0.0082309
ST_1.D | 2024-03-07 02:10 | 2.4668 | 2143.2893 | 2.4708 | 13.474 | 159.069
MOOS4_1.D | 2024-03-07 02:24 | 2.4546 | 2199.5481 | 2.4546 | 13.0558 | 168.4723
ST_2.D | 2024-03-07 02:37 | 2.4647 | 2312.6194 | 2.4607 | 14.196 | 162.906
MOOS4_2.D | 2024-03-07 02:51 | 2.4566 | 2283.8094 | 2.4587 | 14.7329 | 155.014
ST_3.D | 2024-03-07 03:05 | 2.4568 | 2518.5121 | 2.4689 | 14.805 | 170.1123
MOOS4_3.D | 2024-03-07 03:19 | 2.4486 | 2401.2048 | 2.4486 | 17.5275 | 136.9964
ST_4.D | 2024-03-07 03:33 | 2.4587 | 2658.9183 | 2.4627 | 15.0185 | 177.0429
MOOS4_4.D | 2024-03-07 03:46 | 2.4547 | 2506.7605 | 2.4608 | 15.7157 | 159.5065
ST_5.D | 2024-03-07 04:00 | 2.4488 | 2474.3839 | 2.4549 | 15.2417 | 162.3431
MOOS4_5.D | 2024-03-07 04:14 | 2.4608 | 2442.8663 | 2.4588 | 15.3423 | 159.2242
ST_6.D | 2024-03-07 04:28 | 2.4588 | 2672.8106 | 2.4608 | 16.2822 | 164.1551
MOOS4_6.D | 2024-03-07 04:41 | 2.4629 | 2591.7248 | 2.4629 | 15.7345 | 164.7165
ST_7.D | 2024-03-07 04:55 | 2.4548 | 2713.9004 | 2.4609 | 21.8195 | 124.3796
MOOS4_7.D | 2024-03-07 05:09 | 2.4527 | 2655.9279 | 2.4527 | 15.6688 | 169.5041
Effect of blank | [5 empty] | RSD | RSD
0.154% | [5 empty] | 0.1847% | 0.091059
[6 empty] | 0.1565% | 0.091042
[6 empty] | 0.8805% | 0.090938
[6 empty] | 0.354% | 0.091359
[6 empty] | 0.4452% | 0.090927
[6 empty] | 0.213% | 0.091129
[6 empty] | 0.3878% | 0.091139
[6 empty] | 0.2607% | 0.090935
[6 empty] | 0.4833% | 0.090889
[6 empty] | 0.141% | 0.091078
0.3188%
0.2147%
0.6197%
0.2882%
0.1546%
0.2986%
0.3277%
[6 empty] | 0.2986% | 0.1541%
2024-03-06 CHAR MOOS4 SiO2 w2d2 - spectrophotometry (verified on 2024-04-09)
Sample ID | Date | 810nm (Abs)
BLS-DIW | 3/6/2024 1:16:24 PM | 1.3e-05
BSL-LNWS | 3/6/2024 1:17:07 PM | -0.000185
BSL-M4 | 3/6/2024 1:17:47 PM | -0.000135
BSL-SALT | 3/6/2024 1:18:28 PM | 0.00041
BSL-M3 | 3/6/2024 1:19:12 PM | -0.000192
BLK-DIW | 3/6/2024 1:20:00 PM | 0.000529
BLK-LNSW | 3/6/2024 1:20:56 PM | 0.012266
CAL1 | 3/6/2024 1:21:48 PM | 0.076466
CAL2 | 3/6/2024 1:22:41 PM | 0.10976
CAL3 | 3/6/2024 1:23:31 PM | 0.14093
CAL4 | 3/6/2024 1:24:25 PM | 0.20633
CAL5 | 3/6/2024 1:25:14 PM | 0.26781
052a_1 | 3/6/2024 1:26:28 PM | 0.12388
130a_1 | 3/6/2024 1:27:11 PM | 0.12398
207a_1 | 3/6/2024 1:28:49 PM | 0.12454
284a_1 | 3/6/2024 1:29:29 PM | 0.12406
360a_1 | 3/6/2024 1:30:11 PM | 0.12601
052a_2 | 3/6/2024 1:30:55 PM | 0.12447
130a_2 | 3/6/2024 1:31:41 PM | 0.12427
207a_2 | 3/6/2024 1:32:33 PM | 0.12335
284a_2 | 3/6/2024 1:33:24 PM | 0.12411
360a_2 | 3/6/2024 1:34:16 PM | 0.1248
SALT1_309 | 3/6/2024 1:35:30 PM | 0.15579
MOOS3 | 3/6/2024 1:36:25 PM | 0.515
CAL6 | 3/6/2024 1:38:06 PM | 0.55433
2024-03-21 CHAR MOOS4 SiO2 - LC-ICPMS (verified on 2024-04-24)
Sample ID | r28/30 | ur28/30
AB1 | 0.49883 | 0.0072166
AB2A | 0.98138 | 0.014198
AB2B | 0.98352 | 0.014228
AB2C | 0.98176 | 0.014203
AB3 | 1.8255 | 0.02641
181_A1 | 1.003 | 0.014511
181_A2 | 0.9695 | 0.014026
335_A1 | 1.0058 | 0.01455
335_A2 | 0.99173 | 0.014347
027_A2 | 0.95157 | 0.013766
SALT1_309_1 | 0.93493 | 0.013526
SALT1_309_2 | 0.9419 | 0.013626
2024-05-02 Nitrite matrix effects (verified 2024-05-03)
Sample ID | Date | 541nm (Abs)
LNSW-N0 | 5/2/2024 1:45:33 PM | 0.083995 |  | LNSW-N0 | 5/2/2024 1:45:33 PM | 0.083995
MOOS-N0 | 5/2/2024 1:46:24 PM | 0.084394 |  | LNSW-N1 | 5/2/2024 1:47:21 PM | 0.16804
LNSW-N1 | 5/2/2024 1:47:21 PM | 0.16804 |  | LNSW-N2 | 5/2/2024 1:49:04 PM | 0.24606
MOOS-N1 | 5/2/2024 1:48:10 PM | 0.16864 |  | LNSW-N3 | 5/2/2024 1:51:00 PM | 0.32271
LNSW-N2 | 5/2/2024 1:49:04 PM | 0.24606 |  | MOOS-N0 | 5/2/2024 1:46:24 PM | 0.084394
MOOS-N2 | 5/2/2024 1:50:09 PM | 0.24819 |  | MOOS-N1 | 5/2/2024 1:48:10 PM | 0.16864
LNSW-N3 | 5/2/2024 1:51:00 PM | 0.32271 |  | MOOS-N2 | 5/2/2024 1:50:09 PM | 0.24819
MOOS-N3 | 5/2/2024 1:51:52 PM | 0.32585 |  | MOOS-N3 | 5/2/2024 1:51:52 PM | 0.32585
2024-05-02 Phosphate matrix effects (verified 2024-05-03)
Sample ID | Date | 890nm (Abs)
LNSW-P0 | 5/2/2024 2:34:04 PM | 0.27951 |  | LNSW-P0 | 5/2/2024 2:34:04 PM | 0.27951
MOOS-P0 | 5/2/2024 2:34:55 PM | 0.27619 |  | LNSW-P1 | 5/2/2024 2:35:58 PM | 0.55881
LNSW-P1 | 5/2/2024 2:35:58 PM | 0.55881 |  | LNSW-P2 | 5/2/2024 2:37:49 PM | 0.81784
MOOS-P1 | 5/2/2024 2:36:52 PM | 0.55584 |  | LNSW-P3 | 5/2/2024 2:40:06 PM | 1.0643
LNSW-P2 | 5/2/2024 2:37:49 PM | 0.81784 |  | MOOS-P0 | 5/2/2024 2:34:55 PM | 0.27619
MOOS-P2 | 5/2/2024 2:38:43 PM | 0.81985 |  | MOOS-P1 | 5/2/2024 2:36:52 PM | 0.55584
LNSW-P3 | 5/2/2024 2:40:06 PM | 1.0643 |  | MOOS-P2 | 5/2/2024 2:38:43 PM | 0.81985
MOOS-P3 | 5/2/2024 2:41:16 PM | 1.0636 |  | MOOS-P3 | 5/2/2024 2:41:16 PM | 1.0636
2024-05-02 Silicate matrix effects (verified 2024-05-03)
Sample ID | Date | 810nm (Abs)
LNSW-Si0 | 5/2/2024 4:09:51 PM | 0.12779 |  | LNSW-Si0 | 5/2/2024 4:09:51 PM | 0.12779
MOOS-Si0 | 5/2/2024 4:10:44 PM | 0.12651 |  | LNSW-Si1 | 5/2/2024 4:11:37 PM | 0.2425
LNSW-Si1 | 5/2/2024 4:11:37 PM | 0.2425 |  | LNSW-Si2 | 5/2/2024 4:13:29 PM | 0.35417
MOOS-Si1 | 5/2/2024 4:12:32 PM | 0.2419 |  | LNSW-Si3 | 5/2/2024 4:15:19 PM | 0.46216
LNSW-Si2 | 5/2/2024 4:13:29 PM | 0.35417 |  | MOOS-Si0 | 5/2/2024 4:10:44 PM | 0.12651
MOOS-Si2 | 5/2/2024 4:14:25 PM | 0.35136 |  | MOOS-Si1 | 5/2/2024 4:12:32 PM | 0.2419
LNSW-Si3 | 5/2/2024 4:15:19 PM | 0.46216 |  | MOOS-Si2 | 5/2/2024 4:14:25 PM | 0.35136
MOOS-Si3 | 5/2/2024 4:16:09 PM | 0.46098 |  | MOOS-Si3 | 5/2/2024 4:16:09 PM | 0.46098
2024-05-06 Silicate interference on phosphate data (verified 2024-05-10)
Sample ID | Date | 890nm (Abs)
0 | 5/6/2024 1:20:17 PM | -0.000756
0.1 | 5/6/2024 1:21:01 PM | -0.000761
0.4 | 5/6/2024 1:21:47 PM | -0.000758
1 | 5/6/2024 1:22:32 PM | -0.00068
1.6 | 5/6/2024 1:23:21 PM | -0.000789
2.2 | 5/6/2024 1:24:04 PM | -0.00058
3 | 5/6/2024 1:24:49 PM | -0.000679
2024-05-07 Temperature effect on nitrite determination (verified 2024-05-10)
Sample ID | Date | 541nm (Abs)
RT-N1 | 5/7/2024 11:27:36 AM | 0.085787 |  | RT-N1 | 5/7/2024 11:27:36 AM | 0.085787
Fridge-N1 | 5/7/2024 11:28:15 AM | 0.085043 |  | RT-N2 | 5/7/2024 11:28:56 AM | 0.085791
RT-N2 | 5/7/2024 11:28:56 AM | 0.085791 |  | RT-N3 | 5/7/2024 11:30:20 AM | 0.085553
Fridge-N2 | 5/7/2024 11:29:37 AM | 0.084657 |  | Fridge-N1 | 5/7/2024 11:28:15 AM | 0.085043
RT-N3 | 5/7/2024 11:30:20 AM | 0.085553 |  | Fridge-N2 | 5/7/2024 11:29:37 AM | 0.084657
Fridge-N3 | 5/7/2024 11:30:59 AM | 0.085815 |  | Fridge-N3 | 5/7/2024 11:30:59 AM | 0.085815
2024-05-07 Temperature effect on phosphate determination (verified 2024-05-10)
Sample ID | Date | 890nm (Abs)
RT-P1 | 5/7/2024 12:21:47 PM | 0.28332 |  | RT-P1 | 5/7/2024 12:21:47 PM | 0.28332
Fridge-P1 | 5/7/2024 12:22:40 PM | 0.28616 |  | RT-P2 | 5/7/2024 12:23:27 PM | 0.28377
RT-P2 | 5/7/2024 12:23:27 PM | 0.28377 |  | RT-P3 | 5/7/2024 12:24:56 PM | 0.28408
Fridge-P2 | 5/7/2024 12:24:09 PM | 0.28465 |  | Fridge-P1 | 5/7/2024 12:22:40 PM | 0.28616
RT-P3 | 5/7/2024 12:24:56 PM | 0.28408 |  | Fridge-P2 | 5/7/2024 12:24:09 PM | 0.28465
Fridge-P3 | 5/7/2024 12:25:40 PM | 0.28383 |  | Fridge-P3 | 5/7/2024 12:25:40 PM | 0.28383
2024-05-07 Temperature effect on silicate determination (verified 2024-05-10)
Sample ID | Date | 810nm (Abs)
RT-Si1 | 5/7/2024 1:13:58 PM | 0.12907 |  | RT-Si1 | 5/7/2024 1:13:58 PM | 0.12907
Fridge-Si1 | 5/7/2024 1:14:43 PM | 0.10326 |  | RT-Si2 | 5/7/2024 1:15:26 PM | 0.12767
RT-Si2 | 5/7/2024 1:15:26 PM | 0.12767 |  | RT-Si3 | 5/7/2024 1:16:53 PM | 0.12823
Fridge-Si2 | 5/7/2024 1:16:10 PM | 0.098004 |  | Fridge-Si1 | 5/7/2024 1:14:43 PM | 0.10326
RT-Si3 | 5/7/2024 1:16:53 PM | 0.12823 |  | Fridge-Si2 | 5/7/2024 1:16:10 PM | 0.098004
Fridge-Si3 | 5/7/2024 1:17:38 PM | 0.093779 |  | Fridge-Si3 | 5/7/2024 1:17:38 PM | 0.093779
Formulas:
Ratio_IV = [Acq. Date-Time]
sample = [Acq. Date-Time]

## SOFTWARE
Index
Line | Description
13 | CM-Form-P208_Macro_ID4MS. Macro used to perform ID4MS calculation (result; uncertainty)
58 | CM-Form-051_density_of_liquids. Macro used to calculate density of liquids by pycnometry (result; uncertainty)
97 | Gravimetric dilutions WnA1S macro: dilution of 1 solution containing n analytes
137 | Gravimetric dilutions WnAnS macro: mixing n solutions containing 1 analytes each and further dilute
176 | Macro for the calculation of external calibration results (linear model) by Gauss method (ResultEXLN_GS) and by Monte Carlo (ResultEXLN_MC)
236 | Macro for the calculation of IDMS result by fitting a rational function to the calibration plot. Gauss method (ResultIDMS_GS) and Monte Carlo (ResultIDMS_MC)
288 | Macro for the calculation of external calibration results using a quadratic model (ResultEXQD)
CM-Form-P208_Macro_ID4MS. Macro used to perform ID4MS calculation (result; uncertainty)
Validation can be found:
- In this document
- DOC_MANAGER\Archive\Software\CM-Form-P208_Macro_ID4MS
- Within project SALT-1 and CCQM K161
Source code can be found in Module 1 (ALT + F11)
Quadruple isotope dilution | [8 empty] | Percentage variance components of the combined standard uncertainty, uc(wA): First-order Taylor expansion
Calibrators | mA*(A*B)/g | u[mA*(A*B)]/g | mB(A*B)/g | u[mB(A*B)]/g | rA*B/(V/V) | u[rA*B]/(V/V) | wA*/(µg/g) | u[wA*]/(µg/g)
A*B-1 | 1.0265 | 3e-05 | 0.05992 | 3e-05 | 6.143 | 0.018429 | 97.44 | 0.09744
A*B-2 | 2.0031 | 3e-05 | 0.06013 | 3e-05 | 4.215 | 0.012645 | 97.44 | 0.09744
A*B-3 | 4.1054 | 3e-05 | 0.05989 | 3e-05 | 2.764 | 0.008292 | 97.44 | 0.09744
Samples | mA(AB)/g | u[mA(AB)]/g | mB(AB)/g | u[mB(AB)]/g | rAB/(V/V) | u[rAB]/(V/V) | wA/(µg/g) | u[wA]/(µg/g) | rAB | rA*B-1 | rA*B-2 | rA*B-3 | wA*-1 | wA*-2 | wA*-3 | mA(AB) | mB(AB) | mA*(A*B)-1 | mA*(A*B)-2 | mA*(A*B)-3 | mB(A*B)-1 | mB(A*B)-2 | mB(A*B)-3
AB | 2.0081 | 3e-05 | 0.06002 | 3e-05 | 4.226 | 0.012678 | 96.5924 | 0.70439 | 48.63% | 0.0006262% | 48.55% | 0.0003848% | 1.956e-05% | 1.878% | 1.357e-05% | 0.0004197% | 0.4698% | 1.671e-08% | 0.0004213% | 7.245e-10% | 4.904e-06% | 0.4676% | 3.404e-06%
CM-Form-051_density_of_liquids. Macro used to calculate density of liquids by pycnometry (result; uncertainty)
Validation can be found:
- In this document
- DOC_MANAGER\Archive\Software\CM-Form-51_Macro_density
- Documentation associated to CM-Form-051 rev 4 (Jenn Bates)
Source code can be found in Module 2 (ALT + F11)
t /°C | 22 | temperature
RH /% | 50 | relative humidity
Determination of density of aqueous liquids | [11 empty] | Percentage variance components: First-order Taylor expansion
ρH2O | u[ρH2O] | ρair | u[ρair] | m0 | u[m0] | m1 | u[m1] | m2 | u[m2] | ρsample | u[ρsample] | ρH2O | ρair | m0 | m1 | m2
kg/m3 | kg/m3 | kg/m3 | kg/m3 | g | g | g | g | g | g | kg/m3 | kg/m3
997.773 | 0.43224 | 1.2052 | 0.006 | 16.6311 | 0.00013538 | 26.8514 | 0.0013652 | 27.1347 | 0.0014854 | 1025.3889 | 0.48685 | 83.25% | 1.166e-05% | 5.645e-05% | 7.896% | 8.85%
The validation of DenLiq module is achieved by comparison of LiqDen results (L5 and M5) with an indipendent calculation strategy (B23 and B28)
ρsample |  | =(B5-D5)*(J5-F5)/(H5-F5)+D5
kg/m3 | [5 empty] | Ratio of cell L5 and B23
1025.3889 | [5 empty] | 1
u[ρsample] |  | =SQRT((D5-B5)^2*(J5-H5)^2*G5^2/(H5-F5)^4+(D5-B5)^2*(J5-F5)^2*I5^2/(H5-F5)^4+(D5-B5)^2*K5^2/(H5-F5)^2+(J5-F5)^2*C5^2/(H5-F5)^2+(J5-H5)^2*E5^2/(H5-F5)^2)
kg/m3 | [5 empty] | Ratio of cell M5 and B28
0.48685 | [5 empty] | 1
Gravimetric dilutions WnA1S macro: dilution of 1 solution containing n analytes
Validation can be found:
- In this document (validation was also obtained by comparing results produced by previous macros used for CCQM-K161 and SALT-1)
- DOC_MANAGER\Archive\Software\Gravimetric_dilutions
Source code can be found in Module 3 (ALT + F11)
Dilution of one solution containing n analytes
1. Case 1: dilution of one solution containing one analyte
w1 | u(w1) | m0 | m1 | m2 | u(m0) | u(m1) | u(m2) | W1 | u(W1)
µg/g | µg/g | g | g | g | g | g | g | µg/g | µg/g
16.7047 | 0.034693 | 8.2119 | 9.904 | 35.0968 | 0.1 | 0.2 | 5 | 1.0514 | 0.23991
[7 empty] | Manual | 1.0514 | 0.23991
[7 empty] | Verify | TRUE | TRUE
2024-04-05: https://uncertainty.nist.gov/ | [6 empty] | NIST unc | 1.05 | 0.24
[7 empty] | Verify | 100.1% | 99.96%
2. Case 2: dilution of one solution containing two analytes
w1 | u(w1) | w2 | u(w2) | m0 | m1 | m2 | u(m0) | u(m1) | u(m2) | W1 | u(W1) | W2 | u(W2)
µg/g | µg/g | µg/g | µg/g | g | g | g | g | g | g | µg/g | µg/g | µg/g | µg/g
16.7047 | 0.034693 | 54.17 | 0.097712 | 8.2119 | 9.904 | 35.0968 | 0.1 | 0.2 | 5 | 1.0514 | 0.23991 | 3.4094 | 0.77797
[9 empty] | Manual | 1.0514 | 0.23991 | 3.4094 | 0.77797
[9 empty] | Verify | TRUE | TRUE | TRUE | TRUE
2024-04-08: https://uncertainty.nist.gov/ | [8 empty] | NIST unc | 1.05 | 0.24 | 3.41 | 0.77
[9 empty] | Verify | 100.1% | 99.96% | 99.98% | 101%
3. Case 3: dilution of one solution containing three analytes
w1 | u(w1) | w2 | u(w2) | w3 | u(w3) | m0 | m1 | m2 | u(m0) | u(m1) | u(m2) | W1 | u(W1) | W2 | u(W2) | W3 | u(W3)
µg/g | µg/g | µg/g | µg/g | µg/g | µg/g | g | g | g | g | g | g | µg/g | µg/g | µg/g | µg/g | µg/g | µg/g
16.7047 | 0.034693 | 54.17 | 0.097712 | 92.3356 | 0.1174 | 8.2119 | 9.904 | 35.0968 | 0.1 | 0.2 | 5 | 1.0514 | 0.23991 | 3.4094 | 0.77797 | 5.8115 | 1.3261
[11 empty] | Manual | 1.0514 | 0.23991 | 3.4094 | 0.77797 | 5.8115 | 1.3261
[11 empty] | Verify | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE
2024-04-08: https://uncertainty.nist.gov/ | [10 empty] | NIST unc | 1.05 | 0.24 | 3.41 | 0.77 | 5.81 | 1.3
[11 empty] | Verify | 100.1% | 99.96% | 99.98% | 101% | 100% | 102%
Gravimetric dilutions WnAnS macro: mixing n solutions containing 1 analytes each and further dilute
Validation can be found:
- In this document (validation was also obtained by comparing results produced by previous macros used for CCQM-K161 and SALT-1)
- DOC_MANAGER\Archive\Software\Gravimetric_dilutions
Source code can be found in Module 3 (ALT + F11)
Dilution of one solution containing n analytes
1. Case 1: dilution of one solution containing one analyte
w1 | u(w1) | m0 | m1 | m2 | um0 | um1 | um2 | W1 | u(W1)
µg/g | µg/g | g | g | g | g | g | g | µg/g | µg/g
1001 | 2 | 16.4315 | 16.9341 | 46.549 | 0.1 | 0.2 | 1 | 16.7047 | 7.4529
[7 empty] | Manual | 16.7047 | 7.4529
[7 empty] | Verify | TRUE | TRUE
2024-04-08: https://uncertainty.nist.gov/ | [6 empty] | NIST unc | 16.7 | 7.43
[7 empty] | Verify | 100% | 100.3%
2. Case 2: dilution of two solutions containing one analyte
w1 | u(w1) | w2 | u(w2) | m0 | m1 | m2 | m3 | um0 | um1 | um2 | um3 | W1 | u(W1) | W2 | u(W2)
µg/g | µg/g | µg/g | µg/g | g | g | g | g | g | g | g | g | µg/g | µg/g | µg/g | µg/g
1001 | 2 | 1000.5 | 1.7951 | 16.4315 | 16.9341 | 18.5648 | 46.549 | 0.1 | 0.2 | 0.3 | 1 | 16.7047 | 7.4529 | 54.17 | 12.1136
[11 empty] | Manual | 16.7047 | 7.4529 | 54.17 | 12.1136
[11 empty] | Verify | TRUE | TRUE | TRUE | TRUE
[11 empty] | NIST unc | 16.7 | 7.43 | 54.2 | 12.1
[11 empty] | Verify | 100% | 100.3% | 99.94% | 100.1%
3. Case 3: dilution of three solutions containing one analyte
w1 | u(w1) | w2 | u(w2) | w3 | u(w3) | m0 | m1 | m2 | m3 | m4 | um0 | um1 | um2 | um3 | um4 | W1 | u(W1) | W2 | u(W2) | W3 | u(W3)
µg/g | µg/g | µg/g | µg/g | µg/g | µg/g | g | g | g | g | g | g | g | g | g | g | µg/g | µg/g | µg/g | µg/g | µg/g | µg/g
1001 | 2 | 1000.5 | 1.7951 | 755.6021 | 0.95765 | 16.4315 | 16.9341 | 18.5648 | 22.2451 | 46.549 | 0.1 | 0.2 | 0.3 | 0.4 | 1 | 16.7047 | 7.4529 | 54.17 | 12.1136 | 92.3356 | 12.9176
[15 empty] | Manual | 16.7047 | 7.4529 | 54.17 | 12.1136 | 92.3356 | 12.9176
[15 empty] | Verify | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE
[15 empty] | NIST unc | 16.7 | 7.43 | 54.2 | 12.1 | 92.3 | 12.9
[15 empty] | Verify | 100% | 100.3% | 99.94% | 100.1% | 100% | 100.1%
Macro for the calculation of external calibration results (linear model) by Gauss method (ResultEXLN_GS) and by Monte Carlo (ResultEXLN_MC)
Validation can be found:
- In this document
- DOC_MANAGER\Archive\Software\EX_linear
Source code can be found in Module 4 (ALT + F11)
1. Module for calculation of external calibration results - linear model OLS (result, uncertainty)
Data type | x | ux | y | uy | Note
CAL | 0.12929 | 0.00012929 | 0.14955 | 0.0014955 | Experimental data
CAL | 0.19295 | 0.00019295 | 0.22479 | 0.0022479 | Experimental data
CAL | 0.25458 | 0.00025458 | 0.29615 | 0.0029615 | Experimental data
CAL | 0.37386 | 0.00037386 | 0.43266 | 0.0043266 | Experimental data
CAL | 0.50964 | 0.00050964 | 0.59195 | 0.0059195 | Experimental data
PRIM ST | 1000 | 0 |  |  | Certified primary standard
SAMPLE | 0.26245 | 0.0026572 | 0.30466 | 0.0030466 | Calculation by ResultEXLN_GS
SAMPLE | 0.26251 | 0.0028794 | 0.30466 | 0.0030466 | Calculation by ResultEXLN_MT
SAMPLE | 0.26245 | 0.0026572 | 0.30466 | 0.0030466 | Manual calculation
SAMPLE | 0.2625 | 0.00266 | 0.30466 | 0.0030466 | NIST uncertainty machine - https://uncertainty.nist.gov/ (2024-04-09)
LINEST(y,x,TRUE,TRUE)
| a1 | a0
Value | 1.1603 | 0.00013791
Uncertainty | 0.0034153 | 0.0010996
| 0.99997 | 0.0010345
| 115418.4712 | 3
| 0.12353 | 3.2107e-06
2. Calculation of the derivatives
dx/dyS | 0.86186
dx/a0 | -0.86186
dx/a1 | -0.22619
3. Correlation coefficient
r(a0, a1) | -0.90717
4. When the uncertainty on the primary standard is not zero (ResultEXLN_GS)
Data type | x | ux | Note
SAMPLE | 0.26245 | 0.0027086 | Calculation by ResultEXLN_GS
SAMPLE | 0.26245 | 0.0027086 | Manual calculation
PRIM ST | 1000 | 2 | Certified primary standard
Validation | TRUE | TRUE
5. Comparing ResultEXLN_GS with other methods
Ratio | Ratio | ResultEXLN_GS
Value | uncertainty | vs
0.99978 | 0.92283 | ResultEXLN_MT
1 | 1 | Manual calculation
0.99981 | 0.99896 | NIST uncertainty machine - https://uncertainty.nist.gov/ (2024-04-09)
Macro for the calculation of IDMS result by fitting a rational function to the calibration plot. Gauss method (ResultIDMS_GS) and Monte Carlo (ResultIDMS_MC)
Validation can be found:
- In this document
- DOC_MANAGER\Archive\Software\IDMS_fitting
Source code can be found in Module 5 (ALT + F11)
0. Number of Monte Carlo simulations: |  |  |  | 10
1. Case when gravimetric correction for mA and mB is taken into consideration
|  |  |  | x-axis |  | y-axis
Standard | wA* | mA* | mB | wA*(mA*/mB) | u(wA*(mA*/mB)) | rA*B | u(rA*B)
C0 | 0 | 0.4649 | 0.1149 | 0 | 1e-14 | 0.035894 | 0.00035894
C1 | 0.48665 | 0.4837 | 0.1219 | 1.931 | 0.0038621 | 0.17152 | 0.0017152
C2 | 1.3556 | 0.4761 | 0.123 | 5.247 | 0.010494 | 0.33635 | 0.0033635
C3 | 2.217 | 0.4817 | 0.1237 | 8.6332 | 0.017266 | 0.48813 | 0.0048813
C4 | 3.1127 | 0.4711 | 0.1249 | 11.7405 | 0.023481 | 0.62778 | 0.0062778
C5 | 3.9299 | 0.4881 | 0.1248 | 15.3702 | 0.03074 | 0.75748 | 0.0075748
C6 | 4.8552 | 0.4756 | 0.1245 | 18.5473 | 0.037095 | 0.87926 | 0.0087926
[7 empty] | Gauss's method (GUM's Linear Approximation) |  |  |  | Monte Carlo method |  |  | Ratio wA | Ratio u(wA)
Sample | mA | mB | m = mB/mA | u(m) | rAB | u(rAB) | wA | u(wA) | uR(wA) | R^2 | wA | u(wA) | uR(wA) | GC/MT | GC/MT
Sample A_1 | 0.4646 | 0.1152 | 0.24796 | 0.00049591 | 0.72803 | 0.0072803 | 3.5663 | 0.056807 | 1.593% | 0.99918 | 3.5961 | 0.052213 | 1.452% | 99.17% | 108.8%
Sample A_2 | 0.4768 | 0.1234 | 0.25881 | 0.00051762 | 0.70879 | 0.0070879 | 3.5903 | 0.057297 | 1.596% | 0.99918 | 3.6104 | 0.065606 | 1.817% | 99.44% | 87.33%
Sample A_3 | 0.4826 | 0.1251 | 0.25922 | 0.00051844 | 0.71056 | 0.0071056 | 3.6081 | 0.057561 | 1.595% | 0.99918 | 3.6099 | 0.089288 | 2.473% | 99.95% | 64.47%
Sample A_4 | 0.4787 | 0.1249 | 0.26091 | 0.00052183 | 0.69569 | 0.0069569 | 3.53 | 0.056536 | 1.602% | 0.99918 | 3.5322 | 0.052062 | 1.474% | 99.94% | 108.6%
Sample A_5 | 0.4819 | 0.1248 | 0.25897 | 0.00051795 | 0.70727 | 0.0070727 | 3.5823 | 0.057187 | 1.596% | 0.99918 | 3.6201 | 0.063668 | 1.759% | 98.95% | 89.82%
The uncertainty on the primary standard wA* should be added at the end for both Gauss method and Monte Carlo
2. Case when gravimetric correction for mA and mB is not taken into consideration
| x-axis |  |  |  | y-axis
Standard | wA* | u(wA*) | mA* | mB | rA*B | u(rA*B)
C0 | 0 | 1e-14 | n/a | n/a | 0.035894 | 0.00035894
C1 | 0.48665 | 1e-14 | n/a | n/a | 0.17152 | 0.0017152
C2 | 1.3556 | 1e-14 | n/a | n/a | 0.33635 | 0.0033635
C3 | 2.217 | 1e-14 | n/a | n/a | 0.48813 | 0.0048813
C4 | 3.1127 | 1e-14 | n/a | n/a | 0.62778 | 0.0062778
C5 | 3.9299 | 1e-14 | n/a | n/a | 0.75748 | 0.0075748
C6 | 4.8552 | 1e-14 | n/a | n/a | 0.87926 | 0.0087926
[7 empty] | Gauss's method (GUM's Linear Approximation) |  |  |  | Monte Carlo method |  |  | Ratio wA | Ratio u(wA)
Sample | mA | mB | m = mB/mA | u(m) | rAB | u(rAB) | wA | u(wA) | uR(wA) | R^2 | wA | u(wA) | uR(wA) | GC/MT | GC/MT
Sample A_1 | n/a | n/a | 1 | 1e-14 | 0.72803 | 0.0072803 | 3.7388 | 0.059831 | 1.6% | 0.99914 | 3.7887 | 0.038977 | 1.029% | 98.68% | 153.5%
Sample A_2 | n/a | n/a | 1 | 1e-14 | 0.70879 | 0.0070879 | 3.6051 | 0.057811 | 1.604% | 0.99914 | 3.607 | 0.09546 | 2.647% | 99.95% | 60.56%
Sample A_3 | n/a | n/a | 1 | 1e-14 | 0.71056 | 0.0071056 | 3.6173 | 0.057986 | 1.603% | 0.99914 | 3.619 | 0.051514 | 1.423% | 99.95% | 112.6%
Sample A_4 | n/a | n/a | 1 | 1e-14 | 0.69569 | 0.0069569 | 3.5153 | 0.056585 | 1.61% | 0.99914 | 3.5026 | 0.0695 | 1.984% | 100.4% | 81.42%
Sample A_5 | n/a | n/a | 1 | 1e-14 | 0.70727 | 0.0070727 | 3.5947 | 0.057664 | 1.604% | 0.99914 | 3.5898 | 0.066275 | 1.846% | 100.1% | 87.01%
The uncertainty on the primary standard wA* should be added at the end for both Gauss method and Monte Carlo
Macro for the calculation of external calibration results using a quadratic model (ResultEXQD)
Validation can be found:
- In this document
Source code can be found in Module 6 (ALT + F11)
Equation: y = a0 + a1x + a2x2
Equation: a2x2 + a1x + a0 - y = 0
Solution: x1 = {-a1 + [a12 - 4a2 (a0 - y)]1/2}/(2a2)
Solution: x2 = {-a1 - [a12 - 4a2 (a0 - y)]1/2}/(2a2)
Linest with quadratic model
Type | x | y |  | a2 | a1 | a0
CAL | 0.039871 | 0.046203 |  | -0.026063 | 1.1041 | 0.0014866
CAL | 0.059502 | 0.065746 |  | 0.37834 | 0.076205 | 0.0032874
CAL | 0.078505 | 0.088284 |  | 0.99974 | 0.0011779 | #N/A
CAL | 0.11529 | 0.12898 |  | 3797.5321 | 2 | #N/A
CAL | 0.15716 | 0.17415 |  | 0.010537 | 2.7747e-06 | #N/A
Type | y | x1 | x2 | Note
SAMPLE | 0.085842 | 0.076541 | 42.2862 | Using ResultEXQD
SAMPLE | 0.085842 | 0.076541 | 42.2862 | Manual Calculation |  |  | TRUE | TRUE
Function unID4MS(CalMatrix As Object, SampleVector As Object) As Variant
The function unID4MS provides the analytical results wA with uncertainty u[wA] of a quadruple isotope dilution quantitation experiment.
Combined standard uncertainty of all results is estimated in accordance with the recommendations of ISO/GUM for uncorrelated input quantities.
The arguments of unID4MS function are two matrices:
1- The first matrix (CalMatrix) describes the properties of the calibration blends A*B-1 A*B-2 and A*B-3. The form of this matrix is given here in cells C5:J7.
2- The second matrix (SampleVector) describes the properties of the sample blend AB. The form of this vector is given in cells C10:H10.
In this case the correct syntax is: unID4MS(C5:J7,C10:H10). Function unID4MS returns an array (row vector) which is entered over 17 cells (I10:Y10) by pressing ctrl+shift+enter. In cell I10 is returned the value of mass fraction (wA) while on cell J10 is returned the standard uncertainty (u[wA]). On cells K10:Y10 the analysis of the variance component is presented.
N.B. the model of quadruple isotope dilution is based on two assumptions:
1- The isotopic composition of the analyte in the sample A and in the primary standard A* is identical.
2- The measured isotope amount ratios r are directly proportional to the actual amount ratios R (R = K∙r). This assumption may or may not be valid for a particular measuring device.
The impact of this two assumptions on the measurement results is not taken under consideration by ID4MS formulation and needs to be estimated separately.
Function DenLiq(Density As Object) As Variant
The function DenLiq provides the analytical results ρsample with uncertainty u[ρsample] of a density measurement experiment.
Combined standard uncertainty of all results is estimated in accordance with the recommendations of ISO/GUM for uncorrelated input quantities. The argument of DenLiq function is a row vector (ρH2O, u[ρH2O], ρair, u[ρair], m0, u[m0], m1, u[m1], m2, u[m2]). In this case the correct syntax is: DenLiq(B5:K5). Function DenLiq returns an array (row vector) which is entered over 7 cells (L5:R5) by pressing ctrl+shift+enter. In cell L5 is returned the value of liquid sample density (ρsample) while on cell M5 is returned the standard uncertainty (u[ρsample]). On cells N5:R5 the analysis of the variance component is presented.