Size-exclusion chromatography with post-column nucleic acid staining and fluorescence detection for ribonucleic acid analysis

Yannick Hoevelmann1,2, Annkathrin Ratter1,2, Fabian Merkel1, Sebastian Hickert1, Evgeniia Konishcheva1,3

1 Merck KGaA, Darmstadt, Germany

2 Both authors contributed equally to this work

3 Corresponding author: evgeniia.konishcheva@merckgroup.com

Supplementary information

Supplementary Fig. S1 A: SEC traces and molar mass of ssRNA ladder, B: SEC traces and molar mass of dsRNA ladder. Only UV response is shown. MALS, VS, and RI responses are omitted for simplicity reasons. The horizontal black lines correspond to the molar mass (left Y-axis) across the peaks, which was determined from MALS and RI signals: where I is the intensity of scattred light, M is molar mass, c is a concentration from RI and dn/dc=0.17 mL/g for RNA. Molar mass for ssRNA is shown as a continuous line as the species coelute, whereas most dsRNA components are baseline separated, and therefore molar mass can be shown for individual peaks. Mp corresponds to the molar mass at peak maximum, and exp. corresponds to the expected molar mass of the corresponding RNA species. The expected molar mass was calculated from the number of nucleotides (for ssRNA) or base pairs (for dsRNA) provided by a supplier.

Supplementary Fig. S2 SEC traces of dsRNA ladder: Overlay of UV, LS (in this case light scattering response at one out of 8 angles), VS, and RI detector responses

Supplementary Fig. S3 SEC traces (UV response only) of ssRNA (black) and dsRNA (blue) ladders. A: comparison of molar mass depicted on the left Y-axis (horizontal black line corresponds to the molar mass of ssRNA, and the blue lines correspond to the molar mass of dsRNA); B: comparison of intrinsic viscosities depicted on the left Y-axis (horizontal black line corresponds to the intrinsic viscosity of ssRNA, and the blue lines correspond to the intrinsic viscosity dsRNA). Some molar mass and viscosity data show scattering due to relatively low detector response because of low sample concentration. This Figure demonstrates that dsRNA (rod-like conformation) of a comparable nucleotide length elutes before the corresponding ssRNA (random coil conformation) and shows higher viscosity. This illustrates a large impact of conformation and aspect ratio on the elution behavior of RNA in SEC.

Supplementary Fig. S4 Schematic design of the used instrument platform for analysis of nucleic acids based on size-exclusion chromatography with UV detection and post-column nucleic acid staining prior to fluorescence detection.

Supplementary Fig. S5 Nucleic acid dyes investigated for their potential regarding analysis of RNA by size-exclusion chromatographic separation and post-column staining with fluorescence detection

Supplementary Fig. S6 Overlay of SEC-pcs-FLD chromatograms of a 1922 nt-long mRNA (left) and the dsRNA ladder (21 - 500 bp, 7 fragments, right) stained with YOYO-1 after size-exclusion chromatographic separation prior to fluorescence detection at two different emission wavelengths.

Supplementary equation SE1 Einstein-Simna equation for calculation of hydrodynamic radius Rh from intrinsic viscosity [η].

Supplementary Table S1 Excitation and emission wavelengths used for the individual nucleic acid dyes regarding fluorescence detection of RNA after SEC separation and post-column nucleic acid staining

|  | Excitation wavelength [λ] | Emission wavelength [λ] |
| --- | --- | --- |
| SYBR Green I | 485 | 530 |
| SYBR Gold | 495 | 537 |
| Thiazole Orange | 512 | 533 |
| Acridine Orange | 485 | 530/650 |
| YOYO-1 iodide | 491 | 509a/600 |

a	wavelength used regarding displayed SEC-pcs-FLD chromatograms and FLD/UV peak area ratios unless stated otherwise

Supplementary Table S2 mRNA concentrations of the individual calibration solutions

|  | mRNA concentration [µg/ml] |
| --- | --- |
| Calibration solution #1 | 1.0 |
| Calibration solution #2 | 2.5 |
| Calibration solution #3 | 7.5 |
| Calibration solution #4 | 10 |
| Calibration solution #5 | 25 |
| Calibration solution #6 | 75 |
| Calibration solution #7 | 100 |

Supplementary Table S3 The Rg/Rh ratio provides information about the conformation of macromolecules, where Rg is the radius of gyration determined from static light scattering (SLS or in case of chromatography MALS) measurements, and Rh is the hydrodynamic radius determined from dynamic light scattering (DLS) or online viscometry (VS) measurements [1].

Supplementary Table S4 Determined S/N ratios and derived LODs and LOQs based on analysis of the lowest-concentrated mRNA calibration sample using SEC-UV and SEC-pcs-FLD. YOYO-1 is not included, as the calibration samples were not analyzed with this dye due to a limited number of experiments performed owing to the high associated costs; Acridine Orange is not depicted either, as no FLD signal was observed for the mRNA + Acridine Orange due to the dye in flow path causing high background fluorescence saturating the detector

|  | mRNA concentration [µg/ml] (nga) | Nucleic acid dye | S/N ratio (RSD) | LODb |  | LOQb |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | [µg/ml] | [ng]a | [µg/ml] | [ng]a |
| SEC-pcs-FLD | 1 (10) | SYBR Green I | 266.9 (-) | 0.01 | 0.1 | 0.04 | 0.4 |
|  |  | SYBR Gold | 1063.7 (-) | 0.003 | 0.03 | 0.009 | 0.09 |
|  |  | Thiazole Orange | 87.0 (-) | 0.03 | 0.3 | 0.11 | 1.1 |
| SEC-UV |  | - | 6.8c (57%) | 0.44 | 4.4 | 1.48 | 14.8 |

a	amount on column considering 10 µl injection volume
b	considering S/N of 3 for LOD and S/N 10 for LOQ 
c	averaged value based on n = 3 individual samples analyzed on three separate days

Supplementary Table S5 Linearity and precision results obtained for the individual dyes applied for post-column staining of a 1922 nt-long mRNA and subsequent fluorescence detection after size-exclusion chromatographic separation. YOYO-1 is not included, as the calibration samples were not analyzed with this dye due to a limited number of experiments performed owing to the high associated costs; Acridine Orange is not depicted either, as no FLD signal was observed for the mRNA + Acridine Orange due to the dye in flow path causing high background fluorescence saturating the detector

|  | Coefficient of determination (R2) of calibration curvea | Precisionb |
| --- | --- | --- |
| SYBR Green I | 0.9983 | 3.1 |
| SYBR Gold | 0.9985 | 1.7 |
| Thiazole Orange | 0.9975 | 2.0 |

a	covering a concentration range of 1 - 100 µg/ml based on seven calibration solutions and linear regression without any weighting

b	relative standard deviation of peak area based on sixfold analysis of the mRNA sample (75 µg/ml) within one day

Supplementary Table S6 Calculated FLD/UV peak area ratios for the mRNA, ssRNA ladder and dsRNA ladder to assess dye selectivity towards ssRNA and dsRNA species. Acridine Orange is not depicted, as no FLD signal was observed for the mRNA + Acridine Orange due to the dye in flow path causing high background fluorescence saturating the detector

|  | mean FLD/UV peak area ratio mRNAa [-] (RSD) | mean FLD/UV peak area ratio ssRNA ladderb [-] (RSD) | mean FLD/UV peak area ratio dsRNA ladderb [-] (RSD) |
| --- | --- | --- | --- |
| SYBR Green I | 46.9 (4.0%) | 34.4 (9.2%) | 197.7 (30.1%) |
| SYBR Gold | 152.4 (3.8%) | 123.3 (9.6%) | 134.2 (10.1%) |
| Thiazole Orange | 18.8 (2.6%) | 14.4 (7.7%) | 15.3 (16.7%) |
| YOYO-1 iodide | 6.8 (6.5%) | -c | 39.1 (27.6%) |

a	mean value based on sixfold analysis of the mRNA calibration sample (75 µg/ml, n = 6) also used for assessing reproducibility of approach; regarding YOYO-1, displayed mean value is based on threefold analysis (n = 3) of the calibration sample due to a limited number of experiments performed owing to the high associated cost

b	corresponds to the mean FLD/UV ratio across all individual ladder fragments (n = 2 measurements, 7 fragments per ladder)

c	ssRNA ladder was not analyzed using YOYO-1 due to a limited number of experiments performed owing to the high associated costs

Supplementary Table S7 Calculated FLD peak area ratios regarding two different emission wavelengths for the mRNA and dsRNA ladder to assess selectivity of YOYO-1 towards ssRNA and dsRNA species.

|  | mean FLD peak area ratio (λem = 509 nm / 600 nm) mRNAa [-] (RSD) | mean FLD peak area ratio (λem = 509 nm / 600 nm) dsRNA ladderb [-] (RSD in %) |
| --- | --- | --- |
| YOYO-1 iodide | 1.3 (2.9%) | 5.8 (27.8%) |

a	mean value based on threefold analysis of a mRNA calibration sample (75 µg/ml, n = 3)

b	corresponds to the mean FLD peak ratio across all individual ladder fragments (n = 1 measurement, 7 fragments per ladder)

References

- Burchard, W. Adv. Polym. Sci. 1983,48, 1–124.