# Supplementary materials

The law of mass action allows to derive the following relationship between the fraction of bound complex, X=cABcA, total, for a given KD with cA, total being the total molar target concentration, cB, total the total ligand concentration and cAB the concentration of the complex.

X=cB,total+cA,total+KD-cB,total+cA,total+KD2-4(cA,total+cB,total)2cA,total

Supplementary equation (i)

Assuming the measured Fnorm values correspond to a linear superposition of the signals for unbound (Fnorm(0)) and bound (Fnorm(1)) fractions respectively, Fnorm values were fitted using supplementary equation (i) and (ii).

FnormX=Fnorm(0)+X∙Fnorm(1)-Fnorm(0)

Supplementary equation (ii)

| dye | buffer | KD [µM] | σ [µM] | Fluorescence change upon binding |
| --- | --- | --- | --- | --- |
| No dye (LabelFree) | sodium acetate+ a) | 4.1 | 1.1 | NO |
| Blue-NHS (Dylight 488) | sodium acetate+ a) | 10.0 | 1.8 | YES |
| Blue-NHS (Dylight 488) | PBS+ | 17.8 | 2.2 | YES |
| NT-RED-NHS | PBS+ | 97.6 | 18.4 | NO |
| NT-RED-NHS 2nd gen. | PBS+ | 25.5-43.2 b) | ~5 | NO |
| NT-RED-NHS 2nd gen. | TRIS+ | 47.9 | 2.7 | NO |

Supplementary table 1: list of labeling strategies and buffers tested for the lysozyme–NAG3 interaction prior to the 40-instrument benchmark. a) Sodium acetate + buffer: 100mM sodium acetate pH5.0, 0.1% Tween-20. b) Different labeling replicates with different labeling efficiencies.

b)

a)

Supplementary figure 1: a) Fnorm values derived in MOAA for 25 nM RED-NHS 2nd generation dye of each individual instrument for 20, 40 and 60% MST power respectively. b) Values from a) plotted as function of instrument manufacturing year.

Supplementary figure 2: KD versus Fnorm values of the unbound state (Fnorm(0)) derived in MOAA for Lysozyme–NAG3 interaction.

Supplementary figure 3: Exemplary ITC measurement for the lysozyme–NAG3 interaction in 20 mM Tris pH 7.8, 150 mM NaCl, 0.005% Tween-20 at 25°C. 29 µM Lysozyme in the cell and 450 µM NAG3 in the syringe measured with a MicroCal Auto-iTC200, analyzed with the MicroCal PEAQ-ITC analysis software.

Supplementary figure 4: Distribution of extracted KD values for the Lysozyme–NAG3 interaction regarding the S/N ratio of the replicates for all measurements. The horizontal line corresponds to the average KD of 54.8µM while the dashed lines correspond to the S/N dependent relative standard deviation of triplicates as estimated in equation (8).

Supplementary figure 5: Exemplary TRIC binding curve for the lysozyme–NB interaction in PBS+, data from instrument N20. For high ligand concentration non-ideal behavior can be seen with the upper baseline deviation from the expected 1:1 binding model. Graph exported from MOAA.

Supplementary figure 6: Exemplary binding curve from the fluorescence analysis for the lysozyme–NB interaction in PBS+, data from instrument N20. Curve nicely resembles a 1:1 binding model. Graph exported from MOAA.