Supplementary Materials and Methods

SPECT/CT Imaging Details

Karolinska University Hospital

System: Discovery 670 Pro, General Electric Medical Systems

Acquisition: 120 projections, 20 sec dwell time, Medium Energy General Purpose, Step and Shoot

Reconstruction: OSEM 6 iterations / 10 subsets, Butterworth (0.4 cutoff frequency, order 5) with attenuation correction, scatter correction and point spread function deconvolution (Xeleris version 4)

Princess Maxima Center

System: Symbia Intevo 16, Siemens Healthineers

Acquisition: 120 projections, 20 sec dwell time, Medium Energy

Reconstruction: OSCGM 20 iterations, 8 subsets, Gauss 10 mm, with attenuation correction and scatter correction