Supplemental Table 1 – False positive and negative results for zero calcium score
	Category
	Reason
	Percent

	False Positive 
AI detected calcium, but CACS zero with ground truth (n=92, 8.9%)
	Non-coronary calcium (bone, aortic, mitral annulus)
	45%

	
	Noise 
	39%

	
	Ground truth in question
	11%

	
	Artefacts 
	5%

	False Negative 
AI did not detect calcium, but CACS >0 with ground truth (n=40, 4.9%)
	Misclassification* of coronary calcium
	45%

	
	Partial misclassification* of a lesion
	33%

	
	Noise (mainly affecting distal RCA)
	20%

	
	Ground truth in question
	3%


Abbreviations: AI = artificial intelligence, CACS = coronary artery calcium score, RCA = right coronary artery
* As the algorithm classifies individual voxels, it is possible for 2 connected voxels to have different classifications. This is by design so that lesions that span the aorta/left main coronary artery and left anterior descending artery/circumflex artery boundaries can be classified.
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