Appendix 1

Analytical methods

Oxygen isotopes

Zircon oxygen isotope analyses were done using a Cameca IMS 1270 at the University of Edinburgh. The instrumental configuration involved a 133Cs+ primary beam, 10kV acceleration, 2-6nA current, and a 200 µm final aperture. The rest of the parameters are given in Table SM1. Oxygen isotopes 18O- and 16O- were detected using mulitcolector Faraday Cups (H”2 and  L”2). Each analysis was preceded by 60 sec of pre-sputtering. The data were collected within 10 cycles, with a counting time of 5s for each. An electron gun was used for charge neutralisation for all analyses. To correct for instrumental bias, all data were normalised to an internal standard, zircon 91500 (δ18O = 9.86 ± 0.11 ‰, Weidenbeck et al., 2004). Unknowns in blocks of 10 to 15, and internal standards in blocks of 6 to 8, were analysed over2 days. Reproducibility for δ18O value in the 91500 zircon standard zircon ranged from 0.38‰ to 0.64‰ (2 SD). Temora II zircon was anaylsed as the unknown to check the data quality and gave the value of 7.78 ± 0.53 ‰ (n=12, 1SD), which is within error with the published value of 8.2 ‰ (Black et al., 2004).

Black, L.P., Kamo, S.L., Allen, C.M., Davis, D.W., Aleinikoff, J.N., Valley, J.W., Mundil, R., Korsch, R.J., Williams, I.W., Foudoulis, C., 2004. Improved 206Pb/238U microprobe geochronology by the monitoring of a trace-element-related matrix effect; SHRIMP, ID-TIMS, ELA-ICP-MS and oxygen isotope documentation for series of zircon standards. Chemical Geology 205, 115–140.

Wiedenbeck et al. (2004) Further characterization of the 91500 zircon crystal. Geostandards and Geoanalytical Research, 28, 9-39.

| Table SM1 Instrumental parameters for O isotope analyses |  |
| --- | --- |
| Instrumental Configuration: |  |
| Secondary Beam | Negative secondary Ions. Mass Resolution 2,600MRP |
| Transfer lens optics | ~ 60 µm |
| Field Aperture | 4000 µm |
| Energy Slit Width | 40 eV |
| Entrance Slit Width | 80 µm |
| Interference of 16O+H2 on 18O | Mass resolution at 1580 |
| Interference of 17O+H on 18O | Mass resolution at 2309 |
| Automatic Secondary beam alignment | using TRx/y and Entrance slit alignment |

Trace elements

Trace elements in zircon were analysed using Laser ablation-Inductively Coupled Plasma Mass Spectrometer (LA-ICPMS) at the School Of Earth and Environmental Sciences at the University of Portsmouth. Detailed parameters of the analytical session are given in Table SM2. Data were normalized to the glass reference material NIST 610 (Pearce et al., 1997). Every 16 analyses of unknowns were bracketed by two analyses of NIST 610. Zirconium concentration from microprobe was used as internal standard. Data quality (accuracy and reproducibility) was monitored by analyzing two secondary standards as unknowns; the titanite KHAN (Heaman 2009) and the silicate glass MPDing T1-G (Jochum et al. 2006). For the glass MPIding, the measured and published values are in agreement within 10%, with the exception of Nb, Ta and Pb (Appendix 4b). For the titanite KHAN, the agreement is similar with larger variations for Nb, Ta, Hf and Y (Appendix 4b), which are probably due to compositional variability of the mineral (Heaman 2009)  . The reproducibility based on the repeatable analyses of secondary standards was usually below 5 % (2 S.E.) for KHAN and below 15 % (2 S.E.) for MPDing.

| Table SM2 Instrumental parameters during trace element analyses by LA-ICPMS |  |
| --- | --- |
| Laser ablation system |  |
| Make, Model and type | ASI/RESOlution ArF Excimer laser |
| Ablation cell and volume | The dual-volume S155 cell has a stage movement of 155 x 105mm and can accommodate up to twenty 25mm (1″) rounds, or nine 25 x 50mm (1 x 2″) thin sections. However the ablation occurs within a ~15cc funnel. |
| Laser wavelength (nm) | 193 nm |
| Pulse width (ns) | 5 ns |
| Fluence (J cm-2) | 4.5 J cm-2 |
| Repetition rate (Hz) | 5 Hz for unknowns and 10 Hz for NIST 610 |
| Ablation duration (s) | 40 s |
| Spot diameter (mm) nominal/actual | 25-30-40µm |
| Sampling mode / pattern | Static spot ablation |
| Carrier gas | Ar, He and N2 |
| Cell carrier gas flow (l min-1) | He=300 ml.min-1 N2=0.3 ml.min-1 Ar=0.80-0.86 l.min-1 |
| ICP-MS Instrument |  |
| Make, Model and type | Agilent 7500cs |
| Sample introduction | Ablation aerosol |
| RF power (W) | 1300 W |
| Masses measured | 29, 31, 43, 49, 51, 52, 85, 88, 89, 90, 93, 137, 139, 140, 141, 146, 147, 151, 157, 159, 163, 165, 167, 169, 173, 175, 177, 178, 181, 202, 208, 232, 238 |
| Integration time per peak/dwell times (ms); quadrupole settling time between mass jumps | 10 ms for each element |
| Total integration time per output data point (s) | sampling time ~ 0.40 s per mass, acquisition time ~ 74.5s |
| Data Processing |  |
| Gas blank | 25 s |
| Calibration strategy | Glass NIST 610 used as primary reference material, MPIding glass, titanite KHAN, and/or apatite Durango was used as secondary standard for validation |
| Reference Material info | galss NIST 610 (Pearce et al. 1997) |
|  | titanite KHAN (Heaman 2009) |
|  | glass MPIding (Jochum et al. 2006) |
| Data processing package used / Correction for LIEF | SILLS software |

Heaman, L. M. (2009), The application of U-Pb geochronology to mafic, ultramafic and alkaline rocks: An evaluation of three mineral standards, Chem. Geol., 261, 43–52.

Jochum K.P., B. Stoll, K. Herwig, M. Willbold, A.W. Hofmann, M. Amini, S. Aarburg, W. Abouchami, E. Hellebrand, B. Mocek, I. Raczek, A. Stracke, O. Alard, C. Bouman, S. Becker, M. Dücking, H. Brätz, R. Klemd, D.d. Bruin, D. Canil, D. Cornell, C.-J.d. Hoog, C. Dalpe, L. Danyushevsky, A. Eisenhauer, Y. Gao, J.E. Snow, N. Groschopf, D. Günther, C. Latkoczy, M. Guillong, E.H. Hauri, H.E. Höfer, Y. Lahaye, K. Horz, D.E. Jacob, S.A. Kasemann, A.J.R. Kent, T. Ludwig, T. Zack, P.R.D. Mason, A. Meixner, M. Rosner, K. Misawa, B.P. Nash, J. Pfänder, W.R. Premo, W.D. Sun, M. Tiepolo, R. Vannucci, T. Vennemann, D. Wayne, J.D. Woodhead (2006)  MPI-DING reference glasses for in situ microanalysis: new reference values for element concentrations and isotope ratios.   Geochem. Geophys. Geosyst., 7 (Q02008) (2006)

Pearce N.J.G., W.T. Perkins, J.A. Westgate, M.P. Gorton, S.E. Jackson, C.R. Neal, S.P. Chenery, (1997) A compilation of new and published major and trace element data for NIST SRM 610 and NIST SRM 612 glass reference materials. Geostand. Geoanal. Res., 21 (1) (1997), pp. 115–144.

Hafnium isotopes

The analyses were performed using a Thermo-Scientific Neptune multi-collector ICP-MS (Inductively Coupled Plasma–Mass Spectrometer) coupled to a New Wave Research UP193HE Deep-UV (193 nm) ArF Excimer laser ablation sampling system housed at the University of Bristol (Bristol Isotope Group). The beam size was 60 μm and measurements were made using single spots. The data were collected in April 2009 and the acquisition conditions and procedure was the same as that for the session described in Pietranik et al. (2011). Plešovice and 91500 standards were analysed during the session and the data were reduced in the same way as the unknowns. The obtained  176Hf/177Hf  for Plešovice was 0.282478 (12) (n=25), which is in excellent agreement with the  solution-based ratio of 0.282482 (13) reported by Sláma et al. (2008).  The average 176Hf/177Hf  for 91500 was 0.282299 (12) (n=18), which also agrees well with the solution based ratio 0.282302 (18) Woodhead et al. (2004). Altogether the Hf isotopes in reference materials, with the range of 176Yb/177Hf,  are accurately reproduced by the Bristol Isotope Group using the same correction scheme as used in this paper (Fisher et al. 2011). All data are normalized to a JMC 475 value of 176Hf/177Hf = 0.282160. εHf values were calculated using CHUR values of Bouvier et al. (2008) and 176Lu decay value of Söderlund et al (2004).

Bouvier A., J.D. Vervoort, P.J. Patchett (2008) The Lu–Hf and Sm–Nd isotopic composition of CHUR: constraints from unequilibrated chondrites and implications for the bulk composition of terrestrial planets. Earth and Planetray Science Letters, 273 (2008), pp. 48–57

Fisher C.M., J.M. Hanchar, S.D. Samson, B. Dhuime, J. Blichert-Toft, J.D. Vervoort, R. Lam (2011) Synthetic zircon doped with hafnium and rare earth elements: a reference material for in situ hafnium isotope analysis. Chemical Geology, 286 (2011), pp. 32–47

Söderlund U., J.P. Patchett, J.D. Vervoort, C.E. Isachsen (2004) The 176Lu decay constant determined by Lu–Hf and U–Pb isotope systematics of Precambrian mafi c intrusions. Earth and Planetary Science Letters, 219 (2004), pp. 311–324

Sláma J., J. Košler, D.J. Condon, J.L. Crowley, A. Gerdes, J.M. Hanchar, M.S.A. Horstwood, G.A. Morris, L. Nasdala, N. Norberg, U. Schaltegger, B. Schoene, M.N. Tubrett, M.J. Whitehouse (2008) Plešovice zircon — a new natural reference material for U–Pb and Hf isotopic microanalysis. Chemical Geology, 249 (2008), pp. 1–35.

Woodhead J.D., J.M. Hergt, M. Shelley, S. Eggins, R. Kemp (2004) Zircon Hf-isotope analysis with an excimer laser, depth profiling, ablation of complex geometries, and concomitant age estimation. Chemical Geology, 209 (2004), pp. 121–135

Electron microprobe

The analyses were performed using an electron microprobe (EPMA) CAMECA SX100 at the University of Warsaw. The analytical conditions were: 60nA beam current, 15kV voltage and focused beam. The data were used to characterize the general Hf in zircon distribution from core to rim (Fig. 3) and P concentration was used in Fig. 5. To check on quality of these measurements the data were compared to those obtained by LA-ICPMS. As shown in the Figure below, the general agreement between Hf and P analyses by both methods is acceptable.

Fig SM1: Comparison of LAICPMS and microprobe data. Dashed line shows the equal values from both methods.