## MetaData
LA-ICP-MS U-Th-Pb data report
Laboratory & Sample Preparation
Laboratory name | FIERCE, Frankfurt Isotope & Element Research Center Goethe Univesität, Frankfurt am Main
Sample type/mineral | Zircon
Sample preparation | 25 mm polished resin mounts
Imaging | Petrographic microscope & 2400 dpi digital scan
Laser ablation system
Make, Model & type | RESOlution ArF excimer laser (COMpex Pro 102)
Ablation cell | Two-volume ablation cell (Laurin Technic S155)
Laser wavelength | 193 nm
Pulse width | 20 ns
Fluence | ~ 2 J/cm2
Repetition rate | 6 Hz
Pre-ablation | 4 pulses (same parameters as main ablation)
Ablation duration | 18 s
Ablation rate | ~ 0.6 μm/s
Spot shape & size | 23 and 15 µm (diameter)
Sampling mode | Static spot ablation
Gasses | Sample cell: He. Funnel: He + Ar. Tubbing: He + Ar + N
Gas flows | He (350 ml/min), Ar (1050 ml/min), N (12 ml/min).
ICP-MS Instrument
Make, Model & type | ThermoScientific ElementXr sector field ICP-MS
Sample introduction | Ablation aerosol
RF power | 1275 W
Detection system | Secondary electron multiplier (with conversion dynode at -8kV). Simultaneous analogue and counting (pulse) modes of detection (conversion factors calculated per mass and applied offline). Magnetic field fixed. Detection by peak jumping with electrostatic analyzer.
Masses measured | 202, 204, 206, 207, 208, 232, 235, 238
Samples per peak/integration type | 10 for all masses/average
Total time per run | 145 ms
Number of runs/total time | 300 / 43.3 s
Acquisition mode | Trigger from laser (20 s after pre-ablation), background: 18 s, ablation: 18 s
Dead time | 19 ns
Data Processing
Gas blank | 20 s on-peak zero subtracted.
Calibration strategy | - Primary RM: GJ1 - Validation RM: Plesovice, BB-16, 91500
Data processing / LIEF correction | In-house VBA spreadsheet program (Gerdes and Zeh, 2006, 2009). Intercept method for LIEF correction.
Mass discrimination | ²⁰⁷Pb/²⁰⁶Pb (-0.15%) and ²⁰⁶Pb/²³⁸U (3.8%) normalised to primary RM.
Common-Pb correction | No common-Pb correction applied to the data.
Uncertainty level & propagation | Uncertainties are quoted at 2s absolute and are propagated by quadratic addition of the within run precision, counting statistics, background, common Pb correction (if applicable) and the excess of scatter (calculated from the primary RM). In addition, an excess of variance of 0.5 % (1s), calculated from the offset RM, was added quadratically to the ²⁰⁶Pb/²³⁸U ratios. Systematic uncertainties are reported as an expanded uncertainty, considering long term reproducibility (0.8%, 2s) and decay constant uncertainties.
Quality control / Validation | Plesovice: 339.2 ± 1.7 Ma (2s, MSWDC+E = 1.3, n = 7/7) BB-16: 563.2 ± 2.7 Ma (2s, MSWDC+E = 1.3, n = 7/7) 91500 : 1059.0 ± 5.1 Ma (2s, MSWDC+E = 0.43, n = 7/7) (All ages are concordia ages)
Other information

## DataTable
grain | name / | 207Pba | Ub | Pbb | Thb | 206Pbcc | 206Pbd | ±2s | 207Pbd | ±2s | 207Pbd | ±2s | rhoe | 206Pb | ±2s | 207Pb | ±2s | 207Pb | ±2s | concf
| sample | (cps) | (µg/g) | (µg/g) | U | (%) | 238U | (%) | 235U | (%) | 206Pb | (%) |  | 238U | (Ma) | 235U | (Ma) | 206Pb | (Ma) | (%)
U006 | ST7-z1 | 29600.0289 | 10.4124 | 30.7546 | 19.5945 | b.d. | 0.50683 | 1.452 | 12.9337 | 2.0284 | 0.18516 | 1.4164 | 0.71584 | 2643.0853 | 31.5151 | 2674.8306 | 19.204 | 2698.917 | 24.3487 | 97.9313
U007 |  | 23987.1833 | 8.4501 | 8.9787 | 4.5171 | b.d. | 0.50814 | 1.4847 | 12.9713 | 2.1579 | 0.18522 | 1.5659 | 0.68805 | 2648.6976 | 32.2793 | 2677.5732 | 20.4236 | 2699.451 | 26.728 | 98.1199
U008 |  | 22237.28 | 8.0055 | 6.4559 | 2.5113 | 0.010757 | 0.49668 | 1.5506 | 12.487 | 2.1862 | 0.18242 | 1.5411 | 0.70929 | 2599.5253 | 33.2018 | 2641.7516 | 20.6307 | 2674.2576 | 26.3826 | 97.2055
U009 |  | 20664.8975 | 7.4158 | 5.5008 | 2.0668 | 0.028693 | 0.48563 | 1.5616 | 12.3375 | 2.294 | 0.18434 | 1.6805 | 0.68072 | 2551.7595 | 32.9351 | 2630.4315 | 21.6209 | 2691.5387 | 28.5803 | 94.8067
U010 | z2 | 773052.9841 | 233.2569 | 142.2185 | 0.33087 | 0.045778 | 0.54865 | 1.3963 | 15.2801 | 1.629 | 0.20208 | 0.839 | 0.85716 | 2819.537 | 31.924 | 2832.8587 | 15.6434 | 2842.3512 | 15.3382 | 99.1973
U011 | z3 | 586980.3481 | 196.642 | 117.0247 | 0.47796 | b.d. | 0.524 | 1.3179 | 13.2154 | 1.5145 | 0.183 | 0.74639 | 0.87013 | 2716.105 | 29.2461 | 2695.1558 | 14.4136 | 2679.4832 | 14.0719 | 101.3667
U012 | z4 | 674769.2257 | 173.1388 | 110.3047 | 0.16658 | 0.011565 | 0.58563 | 1.3552 | 17.3001 | 1.5658 | 0.21435 | 0.78437 | 0.86549 | 2971.6809 | 32.305 | 2951.6227 | 15.1637 | 2937.9797 | 14.5133 | 101.1471
U013 | z5 | 1188553.1108 | 410.5346 | 221.2338 | 0.22537 | 0.0031131 | 0.49685 | 1.3237 | 12.9123 | 1.5126 | 0.18857 | 0.7319 | 0.87513 | 2600.2345 | 28.358 | 2673.2732 | 14.3699 | 2728.9953 | 13.8429 | 95.2818
U014 | z6 | 808549.1555 | 280.5426 | 156.318 | 0.40416 | b.d. | 0.49688 | 1.3388 | 12.7601 | 1.5455 | 0.18634 | 0.77203 | 0.86629 | 2600.3818 | 28.6824 | 2662.1033 | 14.6641 | 2709.3387 | 14.431 | 95.9785
U015 | z7 | 702708.3804 | 242.7006 | 139.2573 | 0.44275 | 0.013741 | 0.50543 | 1.4173 | 12.8728 | 1.6053 | 0.1848 | 0.75394 | 0.88286 | 2637.0961 | 30.7069 | 2670.3893 | 15.234 | 2695.6993 | 14.1742 | 97.826
U016 | z8 | 286717.6123 | 99.2999 | 61.4408 | 0.95987 | 0.47838 | 0.47708 | 1.4236 | 12.5073 | 1.7003 | 0.19023 | 0.9296 | 0.8373 | 2514.5294 | 29.6723 | 2643.2763 | 16.0868 | 2743.3686 | 16.7423 | 91.6585
U017 | z9 | 1077840.8304 | 372.9176 | 197.4667 | 0.13567 | 0.00036539 | 0.49846 | 1.6369 | 12.9568 | 1.7979 | 0.18861 | 0.74366 | 0.91045 | 2607.1954 | 35.1299 | 2676.5144 | 17.0453 | 2729.3071 | 14.0117 | 95.5259
U018 | z10 | 388665.66 | 123.772 | 81.237 | 1.1415 | b.d. | 0.50843 | 1.3751 | 14.1535 | 1.6339 | 0.20199 | 0.88243 | 0.84161 | 2649.9047 | 29.9111 | 2760.0452 | 15.6084 | 2841.6113 | 15.9728 | 93.2536
U019 | z11 | 451253.9738 | 144.9302 | 79.7257 | 0.25424 | 0.1391 | 0.49818 | 1.6636 | 13.4154 | 1.8999 | 0.1954 | 0.91765 | 0.87563 | 2605.9487 | 35.6885 | 2709.3442 | 18.0477 | 2787.3854 | 16.5302 | 93.4908
U040 | ST8-z1 | 191100.1973 | 65.5241 | 144.103 | 13.6125 | b.d. | 0.50315 | 1.3296 | 13.0024 | 1.5856 | 0.18751 | 0.86394 | 0.83853 | 2627.3373 | 28.7252 | 2679.8266 | 15.0612 | 2719.6699 | 15.7782 | 96.605
U046 |  | 144673.1436 | 50.1329 | 102.7473 | 12.3914 | b.d. | 0.50525 | 1.3608 | 12.9701 | 1.6173 | 0.18626 | 0.87403 | 0.84139 | 2636.3232 | 29.4777 | 2677.4829 | 15.3544 | 2708.7138 | 15.936 | 97.3275
U047 |  | 128652.6629 | 44.7798 | 88.2526 | 11.8812 | b.d. | 0.49584 | 1.3551 | 12.7058 | 1.633 | 0.18593 | 0.9113 | 0.82981 | 2595.91 | 28.9893 | 2658.0907 | 15.4773 | 2705.7569 | 16.4968 | 95.9403
U048 | z2 | 726439.6642 | 280.0525 | 137.9878 | 0.23423 | b.d. | 0.45454 | 1.3397 | 11.7114 | 1.537 | 0.18695 | 0.7533 | 0.87166 | 2415.4134 | 27.0198 | 2581.6094 | 14.4854 | 2714.7833 | 14.1566 | 88.9726
U049 | z5 | 37545.9561 | 157.3835 | 31.7711 | 0.0066152 | 0.59578 | 0.2114 | 2.3307 | 2.5087 | 2.9066 | 0.086106 | 1.7367 | 0.80186 | 1236.2833 | 26.2278 | 1274.553 | 21.1195 | 1339.689 | 33.973 | 92.2814
U050 | z6 | 47327.5544 | 38.8173 | 78.6269 | 11.4378 | 0.088825 | 0.51489 | 1.4948 | 12.8673 | 2.1903 | 0.18133 | 1.6009 | 0.68249 | 2677.4735 | 32.7841 | 2669.9863 | 20.7156 | 2664.3203 | 27.3614 | 100.4937
U051 | z7 | 86387.7914 | 32.1712 | 62.932 | 12.6137 | b.d. | 0.49393 | 1.416 | 12.5259 | 1.7586 | 0.18401 | 1.0429 | 0.80518 | 2587.6408 | 30.2104 | 2644.6726 | 16.6334 | 2688.6015 | 18.5158 | 96.2449
U052 | z8 | 282798.1945 | 239.5416 | 149.1403 | 0.84955 | 0.00047634 | 0.50995 | 1.8673 | 13.1845 | 2.2713 | 0.1876 | 1.2931 | 0.82213 | 2656.4035 | 40.6782 | 2692.9464 | 21.5147 | 2720.49 | 22.3676 | 97.6443
U073 | z9 | 338746.6953 | 124.614 | 87.8681 | 1.5778 | 0.00054269 | 0.49636 | 1.3326 | 12.241 | 1.5458 | 0.17894 | 0.78352 | 0.86203 | 2598.1184 | 28.5288 | 2623.0595 | 14.6201 | 2642.3652 | 14.6336 | 98.3255
U074 |  | 237783.4329 | 86.1435 | 101.0287 | 5.4275 | b.d. | 0.49949 | 1.3471 | 12.3847 | 1.5634 | 0.17991 | 0.79338 | 0.86166 | 2611.6088 | 28.9603 | 2634.0142 | 14.7969 | 2651.2753 | 14.7742 | 98.5039
U075 |  | 280929.8127 | 104.1431 | 63.0819 | 0.80244 | b.d. | 0.50037 | 1.3284 | 12.3202 | 1.5524 | 0.17866 | 0.80331 | 0.85571 | 2615.4065 | 28.5942 | 2629.1156 | 14.689 | 2639.6873 | 14.9279 | 99.0802
U076 | z10 | 1237037.6396 | 408.4147 | 243.062 | 0.41761 | b.d. | 0.52938 | 2.031 | 13.7017 | 2.1671 | 0.1878 | 0.7558 | 0.93721 | 2738.8357 | 45.343 | 2729.3121 | 20.5913 | 2722.2723 | 14.1897 | 100.6084
U077 | z11 | 134415.5878 | 109.1364 | 65.4179 | 0.73909 | 0.1348 | 0.49281 | 1.4889 | 12.3407 | 2.0586 | 0.1817 | 1.4217 | 0.72324 | 2582.816 | 31.7144 | 2630.6745 | 19.4183 | 2667.7065 | 24.4887 | 96.8179
U078 | z13 | 925.4245 | 3.7238 | 0.19917 | 0.099512 | 0.4498 | 0.049232 | 8.102 | 0.61199 | 13.1962 | 0.090197 | 10.4162 | 0.61396 | 309.8034 | 24.5073 | 484.8116 | 50.8708 | 1428.8352 | 198.9123 | 21.6822
U079 | z14 | 688495.1488 | 203.4808 | 118.6647 | 0.22347 | b.d. | 0.53011 | 1.3549 | 15.1973 | 1.6134 | 0.20801 | 0.87598 | 0.83976 | 2741.9236 | 30.2945 | 2827.6819 | 15.4903 | 2889.4029 | 15.85 | 94.8959
U080 | z16 | 394055.3289 | 136.8716 | 77.3788 | 0.39982 | b.d. | 0.5037 | 1.317 | 13.0165 | 1.5377 | 0.18751 | 0.79381 | 0.85645 | 2629.6637 | 28.4729 | 2680.8518 | 14.6137 | 2719.6857 | 14.7436 | 96.69
U096 | ST3-z23 | 152726.7945 | 54.2978 | 37.4311 | 1.3601 | b.d. | 0.50251 | 1.3405 | 12.6921 | 1.6021 | 0.18327 | 0.87738 | 0.83671 | 2624.5938 | 28.9345 | 2657.0765 | 15.187 | 2681.8997 | 16.0053 | 97.8632
U097 |  | 411110.9862 | 144.7916 | 102.7805 | 1.5688 | 0.34259 | 0.49993 | 1.3287 | 12.5479 | 1.5631 | 0.18212 | 0.82327 | 0.85006 | 2613.511 | 28.5838 | 2646.3271 | 14.8099 | 2671.5143 | 15.205 | 97.8288
U098 | z15 | 188103.8282 | 65.1245 | 35.7174 | 0.33139 | 0.061354 | 0.49407 | 1.3768 | 12.2336 | 1.6285 | 0.17966 | 0.86958 | 0.84549 | 2588.2708 | 29.3836 | 2622.4891 | 15.3892 | 2649.0115 | 15.912 | 97.707
U099 |  | 205025.5465 | 76.0738 | 43.3005 | 0.38741 | 0.084884 | 0.51357 | 1.3494 | 12.6726 | 1.6183 | 0.17904 | 0.89338 | 0.83382 | 2671.8683 | 29.5503 | 2655.6229 | 15.3371 | 2643.2714 | 16.2755 | 101.0819
U100 | z27 | 451810.8654 | 150.8059 | 95.3259 | 0.64509 | 0.033667 | 0.53574 | 1.342 | 14.3537 | 1.5905 | 0.1944 | 0.85369 | 0.84376 | 2765.595 | 30.2162 | 2773.37 | 15.2156 | 2779.0319 | 15.5888 | 99.5165
U106 | z31b | 1139872.7395 | 411.5293 | 232.4634 | 0.15735 | 0.049631 | 0.53176 | 1.5042 | 13.7554 | 1.8212 | 0.18769 | 1.0267 | 0.82595 | 2748.8643 | 33.6948 | 2733.0126 | 17.3385 | 2721.32 | 18.2333 | 101.0122
U107 | z31a
U108 | z41 | 694771.5656 | 230.2393 | 122.8053 | 0.083364 | 0.03254 | 0.5053 | 1.3203 | 13.3231 | 1.5361 | 0.19132 | 0.78521 | 0.85949 | 2636.5186 | 28.6049 | 2702.8229 | 14.6247 | 2752.7779 | 14.6019 | 95.7767
U109 | z19 | 628576.0387 | 468.0828 | 276.0503 | 0.20401 | 0.054749 | 0.54615 | 1.5323 | 14.8948 | 1.7757 | 0.19789 | 0.89727 | 0.86293 | 2809.1464 | 34.9236 | 2808.5426 | 17.003 | 2808.109 | 16.2131 | 100.0369
U020 | GJ1 15um | 21735.0205 | 265.5516 | 23.961 | 0.024277 | 0.01157 | 0.097656 | 1.4009 | 0.79857 | 2.1656 | 0.059335 | 1.6514 | 0.6469 | 600.6549 | 8.0411 | 596.0215 | 9.7717 | 578.4277 | 36.1755 | 103.8427
U026 |  | 21892.7895 | 268.9466 | 24.2405 | 0.02347 | 0.026736 | 0.097553 | 1.397 | 0.79196 | 2.0683 | 0.058906 | 1.5252 | 0.67545 | 600.0499 | 8.011 | 592.2815 | 9.2902 | 562.632 | 33.5417 | 106.6505
U027 |  | 22319.1763 | 269.9987 | 24.479 | 0.023678 | 0.034461 | 0.098077 | 1.3818 | 0.80229 | 2.1716 | 0.059355 | 1.6753 | 0.63629 | 603.1308 | 7.9623 | 598.1194 | 9.8239 | 579.1636 | 36.6849 | 104.1382
U028 |  | 21702.5614 | 265.8501 | 24.279 | 0.024136 | 0.029726 | 0.098895 | 1.4058 | 0.79832 | 2.146 | 0.058573 | 1.6214 | 0.65509 | 607.9285 | 8.1624 | 595.8794 | 9.6818 | 550.2753 | 35.6927 | 110.4772
U029 |  | 22456.0152 | 272.9116 | 24.709 | 0.023911 | 0.033819 | 0.097953 | 1.3634 | 0.79761 | 2.0838 | 0.059084 | 1.5759 | 0.65427 | 602.399 | 7.8474 | 595.4791 | 9.3968 | 569.2066 | 34.599 | 105.8313
U030 |  | 22080.2297 | 267.2708 | 24.2452 | 0.023859 | 0.030428 | 0.09812 | 1.3872 | 0.80436 | 2.2169 | 0.059483 | 1.7293 | 0.62572 | 603.3799 | 7.9966 | 599.2864 | 10.043 | 583.828 | 37.8207 | 103.3489
U031 |  | 23264.6065 | 277.9978 | 25.2732 | 0.023601 | 0.030874 | 0.098268 | 1.3729 | 0.81272 | 2.086 | 0.06001 | 1.5705 | 0.65814 | 604.2534 | 7.9252 | 603.9784 | 9.505 | 602.9468 | 34.2953 | 100.2167
U032 |  | 22725.8452 | 275.087 | 24.9107 | 0.023978 | 0.022087 | 0.097963 | 1.3822 | 0.80697 | 2.1215 | 0.059771 | 1.6094 | 0.65155 | 602.46 | 7.9567 | 600.7554 | 9.6287 | 594.3275 | 35.1779 | 101.3684
U033 |  | 22946.1642 | 271.122 | 24.7181 | 0.02393 | 0.026455 | 0.098531 | 1.3501 | 0.82638 | 2.1027 | 0.060856 | 1.612 | 0.64207 | 605.7948 | 7.8129 | 611.6023 | 9.6693 | 633.1685 | 35.0148 | 95.6767
U034 |  | 22192.8678 | 265.7114 | 24.2172 | 0.024006 | 0.015936 | 0.098576 | 1.3621 | 0.81621 | 2.0332 | 0.060079 | 1.5095 | 0.66992 | 606.0606 | 7.8856 | 605.9321 | 9.287 | 605.4514 | 32.9748 | 100.1006
U086 | BB16 15um | 22724.4105 | 296.0496 | 26.9043 | 0.26886 | 0.020083 | 0.091185 | 1.393 | 0.74132 | 2.2992 | 0.05899 | 1.8292 | 0.60586 | 562.5401 | 7.5099 | 563.1781 | 9.9461 | 565.7562 | 40.0924 | 99.4315
U087 |  | 21800.3953 | 289.1909 | 26.2987 | 0.2712 | 0.020779 | 0.091214 | 1.3826 | 0.73108 | 2.2956 | 0.058156 | 1.8325 | 0.60228 | 562.7144 | 7.4562 | 557.185 | 9.8513 | 534.6585 | 40.3727 | 105.2474
U088 |  | 20747.2813 | 273.5143 | 24.7388 | 0.26743 | 0.030852 | 0.090668 | 1.4096 | 0.73402 | 2.2324 | 0.058742 | 1.7311 | 0.63144 | 559.485 | 7.5601 | 558.9087 | 9.6028 | 556.5627 | 38.0288 | 100.5251
U089 |  | 20386.6617 | 269.2948 | 24.3424 | 0.26705 | 0.0095713 | 0.090761 | 1.3893 | 0.72902 | 2.1621 | 0.058282 | 1.6567 | 0.64256 | 560.0372 | 7.4581 | 555.9787 | 9.2642 | 539.3946 | 36.5209 | 103.827
U090 |  | 20543.4592 | 269.4053 | 24.4728 | 0.27011 | 0.016164 | 0.091154 | 1.3896 | 0.7305 | 2.3199 | 0.058149 | 1.8576 | 0.59902 | 562.3564 | 7.4896 | 556.8458 | 9.9508 | 534.3812 | 40.9195 | 105.2351
U091 |  | 21667.2671 | 286.67 | 25.992 | 0.27129 | 0.015449 | 0.09112 | 1.3687 | 0.72774 | 1.983 | 0.05795 | 1.4348 | 0.69026 | 562.1586 | 7.3747 | 555.2239 | 8.4892 | 526.8878 | 31.7761 | 106.6942
U092 |  | 19528.0611 | 259.0674 | 23.2917 | 0.26238 | 0.017018 | 0.090338 | 1.3944 | 0.72355 | 2.1995 | 0.058116 | 1.701 | 0.63397 | 557.5345 | 7.4535 | 552.7624 | 9.383 | 533.1504 | 37.5214 | 104.5736
U093 |  | 21630.9129 | 284.5734 | 25.6986 | 0.26455 | 0.023883 | 0.090741 | 1.4151 | 0.73654 | 2.0826 | 0.058896 | 1.528 | 0.67948 | 559.9156 | 7.5949 | 560.3834 | 8.9771 | 562.2837 | 33.604 | 99.5788
U094 |  | 22250.6274 | 294.6301 | 26.5633 | 0.26834 | 0.0099456 | 0.090505 | 1.3934 | 0.73034 | 2.0744 | 0.058552 | 1.5367 | 0.67171 | 558.5255 | 7.4608 | 556.7506 | 8.8982 | 549.4997 | 33.8633 | 101.6426
U095 |  | 23352.6178 | 301.1784 | 27.6235 | 0.27047 | 0.017759 | 0.091974 | 1.3823 | 0.74958 | 1.9709 | 0.059135 | 1.4049 | 0.70136 | 567.2005 | 7.5116 | 567.9794 | 8.5826 | 571.0989 | 30.9047 | 99.3174
RM0 (n=14) g |  | 53088.4552 | 290 | 26.2908 | 0.0248 | b.d. | 0.098099 | 0.91037 | 0.81255 | 1.3368 | 0.060101 | 0.98591 | 0.47994 | 603.2641 | 5.2551 | 603.8785 | 6.097 | 606.1709 | 21.3347 | 99.5492
RM1 (n=7) g |  | 77103.0053 | 850.7828 | 43.2001 | 0.10235 | b.d. | 0.054004 | 0.68982 | 0.39699 | 1.483 | 0.05334 | 1.6704 | 0.78229 | 339.0574 | 2.2678 | 339.4519 | 4.2714 | 342.1002 | 37.8606 | 99.3789
RM2 (n=7) g |  | 48909.8021 | 293.1112 | 27.0551 | 0.3574 | b.d. | 0.091327 | 1.965 | 0.74117 | 1.3371 | 0.058891 | 0.88264 | 0.77115 | 563.3717 | 10.5962 | 563.0878 | 5.7755 | 561.9928 | 19.1355 | 100.2814
RM3 (n=7) g |  | 12376.2839 | 30.1102 | 5.4313 | 0.31261 | b.d. | 0.17829 | 0.79172 | 1.8447 | 0.98518 | 0.07508 | 0.96353 | 0.58673 | 1057.5957 | 7.7409 | 1061.5579 | 6.6509 | 1069.7098 | 19.2833 | 98.8754
RM0
A001 | GJ1 | 52492.3944 | 289.144 | 26.2778 | 0.024695 | 0.0075732 | 0.09839 | 0.38288 | 0.8096 | 0.75468 | 0.05971 | 0.65035 | 0.50734 | 604.9587 | 2.2344 | 602.2355 | 3.4527 | 591.998 | 14.8267 | 102.1893
A002 | GJ1 | 53254.3516 | 287.9922 | 26.1591 | 0.024897 | 0.0060706 | 0.09827 | 0.44655 | 0.8177 | 0.84447 | 0.06038 | 0.71674 | 0.5288 | 604.2895 | 2.5961 | 606.7808 | 3.8794 | 616.0953 | 16.1471 | 98.0838
A021 | GJ1 | 52452.2991 | 282.7472 | 25.5216 | 0.024678 | 0.0090348 | 0.09762 | 0.3664 | 0.8145 | 0.99601 | 0.06054 | 0.92617 | 0.36787 | 600.4481 | 2.1251 | 604.975 | 4.5583 | 621.9704 | 20.5038 | 96.5397
A022 | GJ1 | 53022.9373 | 292.5238 | 26.3251 | 0.024921 | 0.0068315 | 0.09739 | 0.35222 | 0.8039 | 0.816 | 0.0599 | 0.73607 | 0.43164 | 599.0678 | 2.0403 | 599.0319 | 3.7148 | 598.8963 | 16.5877 | 100.0286
A041 | GJ1 | 54164.8492 | 293.5066 | 26.7243 | 0.024897 | 0.0058751 | 0.09855 | 0.41202 | 0.8123 | 0.8341 | 0.05981 | 0.72524 | 0.49397 | 605.9113 | 2.4047 | 603.7375 | 3.8182 | 595.5849 | 16.3701 | 101.7338
A042 | GJ1 | 53824.7491 | 295.1215 | 26.7466 | 0.024657 | 0.010158 | 0.09805 | 0.3782 | 0.8131 | 0.8031 | 0.06017 | 0.70847 | 0.47093 | 602.962 | 2.2008 | 604.1924 | 3.68 | 608.8109 | 15.9922 | 99.0393
A061 | GJ1 | 53605.5944 | 288.713 | 26.2531 | 0.024858 | 0.0082863 | 0.09836 | 0.33418 | 0.8197 | 0.83976 | 0.06047 | 0.7704 | 0.39795 | 604.7915 | 1.9561 | 607.894 | 3.8632 | 619.475 | 17.2524 | 97.6297
A062 | GJ1 | 52801.607 | 286.4155 | 26.1751 | 0.024955 | 0.0094211 | 0.09888 | 0.40782 | 0.8208 | 0.79934 | 0.06023 | 0.68748 | 0.5102 | 607.8708 | 2.388 | 608.4821 | 3.682 | 610.759 | 15.5539 | 99.5271
A081 | GJ1 | 52933.5278 | 290.8272 | 26.2341 | 0.024598 | 0.0024137 | 0.09765 | 0.33037 | 0.8058 | 0.67483 | 0.05987 | 0.58843 | 0.48956 | 600.6378 | 1.9217 | 600.0743 | 3.0846 | 597.9456 | 13.5461 | 100.4502
A082 | GJ1 | 53655.7895 | 298.0407 | 26.9936 | 0.024827 | 0.010493 | 0.09803 | 0.33149 | 0.8049 | 0.66615 | 0.05957 | 0.57782 | 0.49761 | 602.8679 | 1.9349 | 599.5755 | 3.0438 | 587.1438 | 13.3485 | 102.6781
A101 | GJ1 | 54637.9976 | 301.1177 | 27.1123 | 0.024751 | 0.0067203 | 0.09741 | 0.39973 | 0.8089 | 0.75502 | 0.06025 | 0.64052 | 0.52944 | 599.2118 | 2.3097 | 601.8298 | 3.4525 | 611.7041 | 14.5857 | 97.9578
A102 | GJ1 | 52183.5511 | 282.7485 | 25.6968 | 0.024814 | 0.0075297 | 0.09831 | 0.46355 | 0.8179 | 0.79512 | 0.06037 | 0.64602 | 0.58299 | 604.5111 | 2.6943 | 606.8954 | 3.6559 | 615.8076 | 14.6905 | 98.1656
A120 | GJ1 | 51834.9706 | 283.07 | 25.6462 | 0.024806 | 0.008506 | 0.09802 | 0.36131 | 0.8121 | 0.87803 | 0.06011 | 0.80024 | 0.4115 | 602.8076 | 2.1041 | 603.6168 | 4.0164 | 606.6567 | 17.907 | 99.3655
A121 | GJ1 | 52373.7535 | 288.0317 | 26.2055 | 0.024847 | 0.0088634 | 0.09846 | 0.42506 | 0.8145 | 0.85123 | 0.06003 | 0.73751 | 0.49935 | 605.3612 | 2.4773 | 604.9784 | 3.9016 | 603.5446 | 16.6068 | 100.301
|  | 53088.4552 | 290 | 26.2908 | 0.0248 | b.d. | 0.098099 | 0.91037 | 0.81255 | 1.3368 | 0.060101 | 0.98591 | 0.47994 | 603.2641 | 5.2551 | 603.8785 | 6.097 | 606.1709 | 21.3347 | 99.5492
RM1
B003 | Plešovice | 86228.8268 | 937.3056 | 47.6801 | 0.11214 | 0.005878 | 0.05388 | 1.367 | 0.3989 | 1.7267 | 0.05372 | 1.0549 | 0.7917 | 338.3146 | 4.5092 | 340.8455 | 5.0048 | 358.1415 | 24.2149 | 94.464
B023 | Plešovice | 93628.8532 | 1021.9744 | 51.7687 | 0.10916 | 0.01345 | 0.05366 | 1.4149 | 0.3989 | 1.7356 | 0.05394 | 1.0052 | 0.81523 | 336.9569 | 4.6487 | 340.8502 | 5.0306 | 367.4867 | 23.077 | 91.6923
B043 | Plešovice | 57883.3654 | 609.9055 | 30.9568 | 0.097084 | 0.0021401 | 0.05407 | 1.4647 | 0.395 | 1.9906 | 0.05301 | 1.3481 | 0.73579 | 339.4574 | 4.8467 | 338.0042 | 5.7276 | 328.0196 | 30.9039 | 103.4869
B063 | Plešovice | 68990.9573 | 779.3324 | 39.5168 | 0.10019 | 0.00078917 | 0.05406 | 1.3752 | 0.3933 | 1.723 | 0.05279 | 1.038 | 0.79816 | 339.4083 | 4.5505 | 336.7905 | 4.944 | 318.7552 | 23.9974 | 106.4793
B083 | Plešovice | 70512.4137 | 786.3129 | 39.9099 | 0.097861 | 0.0012277 | 0.05413 | 1.3315 | 0.3986 | 1.7021 | 0.05344 | 1.0604 | 0.78224 | 339.814 | 4.4111 | 340.6466 | 4.9312 | 346.3321 | 24.3847 | 98.1179
B103 | Plešovice | 89086.3917 | 990.2446 | 50.4755 | 0.10201 | 0.0031029 | 0.05423 | 1.3732 | 0.4006 | 1.7503 | 0.05359 | 1.0853 | 0.78453 | 340.4334 | 4.557 | 342.0485 | 5.0881 | 353.0376 | 24.9124 | 96.4298
B122 | Plešovice | 73390.2289 | 830.4039 | 42.0932 | 0.098016 | 0.015801 | 0.054 | 1.3574 | 0.3936 | 1.7666 | 0.05289 | 1.1307 | 0.76837 | 339.0172 | 4.4867 | 336.9778 | 5.0713 | 322.9286 | 26.0524 | 104.9821
|  | 77103.0053 | 850.7828 | 43.2001 | 0.10235 | b.d. | 0.054004 | 0.68982 | 0.39699 | 1.483 | 0.05334 | 1.6704 | 0.78229 | 339.0574 | 2.2678 | 339.4519 | 4.2714 | 342.1002 | 37.8606 | 99.3789
RM2
C004 | BB-16 | 48480.4601 | 289.1916 | 26.353 | 0.34629 | b.d. | 0.09034 | 1.3428 | 0.736 | 1.7473 | 0.05912 | 1.118 | 0.76851 | 557.531 | 7.1781 | 560.0716 | 7.5314 | 570.4087 | 24.7511 | 97.7424
C024 | BB-16 | 50963.7761 | 304.3463 | 27.9854 | 0.36759 | 0.00068695 | 0.09062 | 1.3636 | 0.7408 | 1.7585 | 0.05932 | 1.1103 | 0.77545 | 559.2254 | 7.3104 | 562.8865 | 7.6081 | 577.7119 | 24.5594 | 96.8
C044 | BB-16 | 49109.4048 | 292.2026 | 26.9937 | 0.35233 | b.d. | 0.09152 | 1.3487 | 0.7431 | 1.72 | 0.05892 | 1.0674 | 0.78413 | 564.5066 | 7.2958 | 564.2403 | 7.4553 | 563.167 | 23.6976 | 100.2379
C064 | BB-16 | 51001.3111 | 309.6709 | 28.3742 | 0.3671 | b.d. | 0.09044 | 1.3313 | 0.7341 | 1.7222 | 0.0589 | 1.0925 | 0.77301 | 558.1194 | 7.1237 | 558.9556 | 7.4124 | 562.3616 | 24.2383 | 99.2457
C084 | BB-16 | 42372.4178 | 255.0644 | 23.509 | 0.33201 | 0.00082799 | 0.09184 | 1.3433 | 0.7427 | 1.7748 | 0.05868 | 1.1599 | 0.75688 | 566.408 | 7.29 | 563.9623 | 7.6893 | 554.109 | 25.7146 | 102.2196
C104 | BB-16 | 51144.3338 | 305.7261 | 28.5235 | 0.38107 | 0.00028995 | 0.09176 | 1.3454 | 0.7424 | 1.7608 | 0.05871 | 1.1359 | 0.7641 | 565.9248 | 7.2958 | 563.8044 | 7.6275 | 555.2575 | 25.1948 | 101.9211
C123 | BB-16 | 49296.9111 | 295.5764 | 27.6469 | 0.35539 | b.d. | 0.09277 | 1.3392 | 0.7491 | 1.7259 | 0.05859 | 1.0887 | 0.77596 | 571.8864 | 7.3354 | 567.6938 | 7.5153 | 550.934 | 24.1988 | 103.8031
|  | 48909.8021 | 293.1112 | 27.0551 | 0.3574 | b.d. | 0.091327 | 1.965 | 0.74117 | 1.3371 | 0.058891 | 0.88264 | 0.77115 | 563.3717 | 10.5962 | 563.0878 | 5.7755 | 561.9928 | 19.1355 | 100.2814
RM3
D005 | 91500 | 12346.7955 | 30.2568 | 5.4683 | 0.31122 | 0.044059 | 0.1789 | 1.4189 | 1.838 | 2.5608 | 0.07456 | 2.1317 | 0.55409 | 1060.8531 | 13.8906 | 1059.2241 | 16.8557 | 1055.8689 | 43.216 | 100.472
D025 | 91500 | 12205.5242 | 29.5107 | 5.3606 | 0.31522 | 0.043523 | 0.1794 | 1.4403 | 1.854 | 2.3169 | 0.07501 | 1.8148 | 0.62165 | 1063.6691 | 14.1344 | 1065.0231 | 15.301 | 1067.7977 | 36.8208 | 99.6134
D045 | 91500 | 12614.6471 | 30.0662 | 5.4032 | 0.3111 | 0.042597 | 0.1777 | 1.4399 | 1.839 | 2.4981 | 0.07507 | 2.0414 | 0.57639 | 1054.6033 | 14.0193 | 1059.417 | 16.4458 | 1069.3473 | 41.3285 | 98.6212
D065 | 91500 | 12658.6675 | 31.1905 | 5.6098 | 0.31268 | 0.038677 | 0.1779 | 1.5416 | 1.836 | 2.7258 | 0.0749 | 2.2479 | 0.56557 | 1055.286 | 15.0173 | 1058.4041 | 17.9319 | 1064.8401 | 45.4813 | 99.1028
D085 | 91500 | 12314.1032 | 29.9742 | 5.412 | 0.31244 | 0.014043 | 0.1787 | 1.4506 | 1.856 | 2.2694 | 0.07536 | 1.7452 | 0.63921 | 1060.1094 | 14.1916 | 1065.7255 | 14.9934 | 1077.2398 | 35.3875 | 98.4098
D105 | 91500 | 12296.8208 | 30.1565 | 5.4301 | 0.31268 | 0.017994 | 0.1778 | 1.4522 | 1.837 | 2.505 | 0.07496 | 2.0412 | 0.57971 | 1054.7587 | 14.1408 | 1058.6407 | 16.4843 | 1066.6531 | 41.34 | 98.8849
D124 | 91500 | 12197.4292 | 29.6167 | 5.3351 | 0.31295 | 0.0539 | 0.1776 | 1.4213 | 1.853 | 2.4914 | 0.0757 | 2.0461 | 0.57051 | 1053.8903 | 13.8306 | 1064.4711 | 16.4459 | 1086.2219 | 41.3226 | 97.0235
|  | 12376.2839 | 30.1102 | 5.4313 | 0.31261 | b.d. | 0.17829 | 0.79172 | 1.8447 | 0.98518 | 0.07508 | 0.96353 | 0.58673 | 1057.5957 | 7.7409 | 1061.5579 | 6.6509 | 1069.7098 | 19.2833 | 98.8754