## Slide 1
In vitro measurements of ultrafiltration precision in hemofiltration and haemodialysis devices used in infants, Part 2*
AIM: To determine whether CARPEDIEM and PrisMax have sufficient precision of ultrafiltration control to treat small babies
Crosier et al. 2025
CONCLUSION: PrisMax and CARPEDIEM control fluid ultrafiltration better than Prismaflex and Aquarius, but the novel volumetric circuit of the NIDUS infant haemodialysis device is the most precise for small babies.
DESIGN & OUTCOMES:
NIDUS CVVHD tested in 2022
CARPEDIEM
CVVH & CVVHD
PrisMax CVVH,
CVVHD & CVVHDF
All 3 devices tested in vitro:
- NIDUS tested and reported in 2022.
- CARPEDIEM & PrisMax tested now.
Treated a ‘baby substitute’ saline bag.
Quantity of fluid removed measured by weighing and compared with the device’s reported UF volume.
CARPEDIEM & PrisMax reasonably precise.
NIDUS fluid control more precise.
(*Part 2: Comparison of PrisMax and CARPEDIEM with 2022 data on NIDUS, when Prismaflex and Aquarius shown to be imprecise)