## Slide 1
Clinical performance and microstructural assessment of hemodialysis catheters for children on maintenance dialysis
AIM: to review the causes of CVC dysfunction and inspect the extent and nature of occlusions when catheters are removed
DESIGN & OUTCOMES:
Bruno et al. 2026
CONCLUSION: CVC dysfunctions are frequent and do not correlate with luminal diameter. Fibrin-rich thrombi are a common cause of incomplete occlusion. These findings have implications for future CVC design to optimize vascular access in children on maintenance dialysis.
Single-center collection of explanted CVC
Multicenter retrospective data collection on CVC performance
n = 16 CVCs explanted and collected
n = 154 patients aged 0-18 years on HD via CVC
288 catheters of 6 models of different sizes
54.2% of the lines had at least one dysfunction
Median CVC survival was 215.5 days, with 133 CVC replacements in 72 children
Larger CVCs (≥14Fr vs. 10Fr) did not show better access survival, although this may be influenced by design differences
Micro-CT revealed occlusions in all explanted CVCs, with histological confirmation of thrombi in 7 CVCs
Thrombi were rich in fibrin and predominantly located near side holes and tip