Table S1

(a) Operating conditions for the LA-ICP-SF-MS equipment for U-Th-Pb isotope analysis

| Laboratory & Sample Preparation |  |
| --- | --- |
| Laboratory name | Institute for Applied Geosciences, KIT (Karlsruhe Institute of Technology), Germany |
| Sample type/mineral | Magmatic zircons |
| Sample preparation | Conventional mineral separation, 1 inch resin mount, polish to finish |
| Imaging | VEGA TESCAN with Oxford detector, BSE mode 15 KV, <10nA (at KIT) |
| Laser ablation system |  |
| Make, Model & type | Teledyne Photon Machines, Analyte Exite+ (Excimer) |
| Ablation cell | Two-volume ablation cell (HELEX 100, EQC COMP), ANUAustralia) |
| Laser wavelength | 193 nm |
| Pulse width | < 5 ns |
| Fluence | 2.7 J/cm-2 |
| Repetition rate | 8 Hz |
| Spot size | 33 µm |
| Sampling mode / pattern | single spot |
| Carrier gas | He(cell) 0.30 l/min, He (cup) 0.15 l/min, 0.91 l/min Ar, N2 (9 ml) |
| Background collection | 15 seconds |
| Ablation duration | 15 seconds |
| Pre-ablation | 3 pulses |
| Wash-out delay | 15 seconds |
| Cell carrier gas flow (He) | 0.45 l/min (total) |
| ICP-MS Instrument |  |
| Make, Model & type | Thermo-Scientific ELEMENT XR (sector field) |
| Sample introduction | via conventional tubing |
| RF power | 1210 W |
| Sampler, skimmer cones | Ni-Jet-cone |
| Extraction lenses | X - type |
| Make-up gas flow (Ar) | 0.91 l/min |
| Detection system | single collector secondary electron multiplier |
| Data acquisition protocol | Time-resolved analysis |
| Scanning mode | Peak hopping, four point per peak |
| Detector mode | Pulse counting mode |
| Masses measured | 202Hg,204(Hg + Pb), 206Pb, 207Pb, 208Pb, 232Th, 238U |
| Integration time per peak | 10 milliseconds |
| Sensitivity / Efficiency | 54000 (cts/ppm) (238U, NIST612, spot-35μm, 10μm/s, 10 Hz, 5 J/cm2) oxide formation rate: U/UO <0.1 , Th/U = 0.95 |
| Data Processing |  |
| Gas blank | 15 seconds on-peak |
| Calibration strategy | BB1 zircon standard used as primary reference material, Plešovice, KA1 (Kaap Valley tonalite) as secondary reference material (quality control) |
| Reference Material info | BB (Santos et al., 2017): concordia age = 562 ± 0.6 Ma (LA-ICP-MS) Plešovice (Sláma et al., 2008) = 337.13 ± 0.37 Ma (TIMS) KA1 (Schoene et al. (2006): 207/206 = 3227.2 ± 0.2 Ma (CA-ID-TIMS). |
| Data processing package used | in-house EXCEL spreadsheet (Gerdes and Zeh, 2006, 2009) |
| Quality control / Validation | BB (session 1-4): concordia age = 561.4 ± 1.2 Ma (n = 86), MSWD(C+E)=0.79, Prob.(C+E)=0.98 Plešovice (session1-4): concordia age = 337.35 ± 0.75 Ma (n = 85) MSWD(C+E)=0.81, Prob.(C+E)=0.97 KA1(session 1-4): 207/206 u.i. age = 3230.0± 8.1 Ma (n = 56) MSWD=0.66, Prob.=0.98 |

(b) Operating conditions for the LA-MC-ICP-MS equipment for Lu-Hf isotope analysis

| Laboratory |  |
| --- | --- |
| Laboratory name | FIERCE (Frankfurt Isotope and Element Research Centre), Goethe University Frankfurt, Germany |
| Laser ablation system |  |
| Make, Model & type | ComPexPro 102F, Coherent (Excimer) |
| Ablation cell | Two-volume ablation cell (Laurin Technic, Australia) |
| Laser wavelength | 193 nm |
| Pulse width | 20 ns |
| Fluence | 6 J/cm-2 |
| Repetition rate | 6.0 Hz |
| Spot size | Circular, 33 µm (diameter) |
| Sampling mode / pattern | Single spot |
| Carrier gas | 0.89 l min-1Ar + “squid” smoothing device directly after ablation cell |
| Background collection | 30 seconds |
| Ablation duration | 39 seconds |
| Wash-out delay | ca. 30 seconds |
| Cell carrier gas flow | He (300 ml/min), Ar (1100 ml/min), N2 (5-7 ml/min) |
| ICP-MS Instrument |  |
| Make, Model & type | Thermo-Finnigan NEPTUNE Plus - MC ICP-MS |
| Sample introduction | Ablation aerosol |
| RF power | 1300 W |
| Sampler, skimmer cones | Ni-Jet |
| Extraction lenses | X-cone |
| Make-up gas flow (Ar) | 0.89 l/min |
| Detection system | Static mode with faraday collectors, 10-11 Ω amplifiers; Raw data correction for gain and baseline. |
| Data acquisition protocol | See Gerdes and Zeh (2009) |
| Scanning mode |  |
| Detector mode | Multi collector, static |
| Masses measured | 172Yb,173Yb,175Lu,176Hf-Yb-Lu,177Hf,178Hf, 179Hf, |
| Dwell time | 0.51s for all masses (static mode) |
| Sensitivity / Efficiency | 120 mV/pg Hf |
| Number of runs/total time | 70 / 36 s |
| Acquisition mode | Trigger from laser (3s after start of ablation), background aquired every 30 mins, ablation: 39 s |
| Data Processing |  |
| Gas blank | 38 s on-peak zero subtracted. |
| Calibration strategy | Reference zircon GJ1 and Temora 2 were measured along with unknowns |
| Reference Material info | Woodhead and Hergt (2005); Morel et al. (2008) |
| Data processing package used | Excel spreadsheet (Gerdes and Zeh, 2006, 2009) |
| Quality control / Validation | Results of standard measurements: GJ-1 176Hf/177Hf = 0.282000 ± 0.000026 (n = 22) Temora 1: 176Hf/177Hf = 0.282650 ± 0.000031 (n = 22) |

References

Gerdes, A., Zeh, A., 2006. Combined U–Pb and Hf isotope LA-(MC-)ICP-MS analyses of detrital zircons: Comparison with SHRIMP and new constraints for the provenance and age of an Armorican metasediment in Central Germany. Earth Planet. Sci. Lett. 249, 47–61.

Gerdes, A., Zeh, A., 2009. Zircon formation versus zircon alteration – new insights from combined U–Pb and Lu–Hf in-situ LA-ICP-MS analyses, and consequences for the interpretation of Archean zircon from the Central Zone of the Limpopo Belt. Chem. Geol. 261, 230–243.

Morel, M.L.A., Nebel, O., Nebel-Jacobsen, Y. J., Miller, J. S., Vroon, P. Z., 2008. Hafnium isotope characterization of the GJ-1 zircon reference material by solution and laser-ablation MC-ICPMS. Chem. Geol. 255, 231–235.

Santos, M.M., Lana, C., Scholz, R., Buick, I., Schmitz, M.D., Kamo, S.L., Gerdes, A., Corfu, F., Tapster, S., Lancaster, P., Storey, C.D., Basei, M.A.S., Tohver, E., Alkmim, A., Nalini, H., Krambrock, K., Fantini, C., Wiedenbeck, M., 2017. A new appraisal of Sri Lankan BB zircon as a reference material for LA–ICP–MS U–Pb geochronology and Lu–Hf isotope tracing. Geostand. Geoanalyt. Res. 41, 335–358.

Sláma, J., Kŏsler, J., Condon, D.J., Crowley, J.L., Gerdes, A., Hanchar, J.M., Horstwood, M.S.A., Morris, G.A., Nasdala, L., Norberg, N., Schaltegger, U., Schoene, B., Tubrett, M.N., Whitehouse, M.J., 2008. Plešovice zircon — A new natural reference material for U-Pb and Hf isotopic microanalysis. Chem. Geol. 249 (1-2), 1–35.

Schoene, B., Crowley, J. L., Condon, D. J., Schmitz, M. D., Bowring, S. A., 2006. Reassessing the uranium decy constants for geochronology using ID-TIMS U-Pb data. Geochimica et Cosmochimica Acta 70, 426-445.

Stern, R. A., Bodorkos, S., Kamo, S. L., Hickman, A. H., Corfu, F., 2009. Measurement of SIMS instrumental mass fractionation of Pb isotopes during zircon dating. Geostand. Geoanalyt. Res. 33,145–168.

Woodhead, J.D., Hergt, J.M., 2005. A preliminary appraisal of seven natural zircon reference materials for in situ Hf isotope determination. Geostand. Geoanalyt. Res. 29, 183–195.