New insights into Cadomian basin evolution and stratigraphic affiliation of sedimentary units of Saxo-Thuringia, Germany: Part 2 – detrital zircon U–Pb ages

by Victoria Kühnemann, Guido Meinhold, Armin Zeh, Sebastian Weber

ESM4: Operating conditions for the LA-ICP-MS equipment for U-Th-Pb isotope analysis

- Measuring session Nov_2023

| Laboratory & Sample Preparation |  |
| --- | --- |
| Laboratory name | Institute for Applied Geosciences, KIT (Karlsruhe Institute of Technology), Germany |
| Sample type/mineral | detrital zircons |
| Sample preparation | Conventional mineral separation, 1 inch resin mount, polish to finish |
| Imaging | VEGA TESCAN with Oxford detector, COMPO SEM mode 15 KV (at Institut für Geologie, TU Bergakademie Freiberg) |
| Laser ablation system |  |
| Make, Model & type | Teledyne Photon Machines, Analyte Excite+ (Excimer) |
| Ablation cell | Two-volume ablation cell (HELEX 100, EQC COMP), ANUAustralia) |
| Laser wavelength | 193 nm |
| Pulse width | < 5 ns |
| Fluence | 2.0 J/cm-2 |
| Repetition rate | 10 Hz |
| Spot size | 20 µm |
| Sampling mode / pattern | single spot |
| Carrier gas | He(cell) 0.30 l/min, He (cup) 0.20 l/min, 0.91 l/min Ar, N2 (12 ml) |
| Background collection | 15 seconds |
| Ablation duration | 15 seconds |
| Pre-ablation | 3 pulses |
| Wash-out delay | 15 seconds |
| Cell carrier gas flow (He) | 0.50 l/min (total) |
| ICP-MS Instrument |  |
| Make, Model & type | Thermo-Scientific ELEMENT XR (sector field) |
| Sample introduction | via conventional tubing |
| RF power | 1225 W |
| Sampler, skimmer cones | Ni-X-cone |
| Extraction lenses | X - type |
| Make-up gas flow (Ar) | 0.91 l/min |
| Detection system | single collector secondary electron multiplier |
| Data acquisition protocol | Time-resolved analysis |
| Scanning mode | Peak hopping, four point per peak |
| Detector mode | Pulse counting mode |
| Masses measured | 202Hg,204(Hg + Pb), 206Pb, 207Pb, 208Pb, 232Th, 238U |
| Integration time per peak | 10 milliseconds |
| Sensitivity / Efficiency | 77000 (cts/ppm) (238U, NIST612, spot-35 μm, 10μm/s, 10 Hz, 5 J/cm2) oxide formation rate: U/UO <0.07, Th/U = 0.90 |
| Data Processing |  |
| Gas blank | 15 seconds on-peak |
| Calibration strategy | BB1 zircon standard used as primary reference material, Plešovice, KA1 (Kaap Valley tonalite) as secondary reference material (quality control) |
| Reference Material info | BB (Santos et al., 2017): concordia age = 562 ± 0.6 Ma (LA-ICP-MS) Plešovice (Sláma et al., 2008) = 337.13 ± 0.37 Ma (TIMS) KA1 (Schoene et al. (2006): 207/206 = 3227.2 ± 0.2 Ma (CA-ID-TIMS). |
| Data processing package used | in-house EXCEL spreadsheet (Gerdes and Zeh, 2006, 2009) |
| Quality control / Validation | BB (session 1): concordia age = 562.1 ± 1.6 Ma (n = 28) Plešovice (session1): concordia age = 337.4 ± 1.1 Ma (n = 28) KA1(session 1): 207/206 age = 3227.0± 7.9 Ma (n = 19) BB (session 2) Concordia age = 562.3 ± 2.0 Ma (n = 22) Plešovice (session 2): concordia age = 338.0 ± 1.4 Ma (n = 22) KA1(session 2): 207/206 age = 3231.0 ± 9.0 Ma (n = 16) |

- Measuring session Feb_2024

| Laboratory & Sample Preparation |  |
| --- | --- |
| Laboratory name | Institute for Applied Geosciences, KIT (Karlsruhe Institute of Technology), Germany |
| Sample type/mineral | detrital zircons |
| Sample preparation | Conventional mineral separation, 1 inch resin mount, polish to finish |
| Imaging | VEGA TESCAN with Oxford detector, COMPO SEM mode 15 KV, (at Institut für Geologie, TU Bergakademie Freiberg) |
| Laser ablation system |  |
| Make, Model & type | Teledyne Photon Machines, Analyte Excite+ (Excimer) |
| Ablation cell | Two-volume ablation cell (HELEX 100, EQC COMP), ANUAustralia) |
| Laser wavelength | 193 nm |
| Pulse width | < 5 ns |
| Fluence | 2.2 J/cm-2 |
| Repetition rate | 10 Hz |
| Spot size | 25 µm |
| Sampling mode / pattern | single spot |
| Carrier gas | He(cell) 0.35 l/min, He (cup) 0.20 l/min, 0.94 l/min Ar, N2 (10 ml) |
| Background collection | 15 seconds |
| Ablation duration | 15 seconds |
| Pre-ablation | 5 pulses |
| Wash-out delay | 15 seconds |
| Cell carrier gas flow (He) | 0.55 l/min (total) |
| ICP-MS Instrument |  |
| Make, Model & type | Thermo-Scientific ELEMENT XR (sector field) |
| Sample introduction | via conventional tubing |
| RF power | 1220 W |
| Sampler, skimmer cones | Ni-X-cone |
| Extraction lenses | X - type |
| Make-up gas flow (Ar) | 0.94 l/min |
| Detection system | single collector secondary electron multiplier |
| Data acquisition protocol | Time-resolved analysis |
| Scanning mode | Peak hopping, four point per peak |
| Detector mode | Pulse counting mode |
| Masses measured | 202Hg,204(Hg + Pb), 206Pb, 207Pb, 208Pb, 232Th, 238U |
| Integration time per peak | 10 milliseconds |
| Sensitivity / Efficiency | 27000 (cts/ppm) (238U, NIST612, spot-20 μm, 10μm/s, 10 Hz, 5 J/cm2) oxide formation rate: U/UO <0.07, Th/U = 0.90 |
| Data Processing |  |
| Gas blank | 15 seconds on-peak |
| Calibration strategy | BB1 zircon standard used as primary reference material, Plešovice, KA1 (Kaap Valley tonalite) as secondary reference material (quality control) |
| Reference Material info | BB (Santos et al., 2017): concordia age = 562 ± 0.6 Ma (LA-ICP-MS) Plešovice (Sláma et al., 2008) = 337.13 ± 0.37 Ma (TIMS) KA1 (Schoene et al. (2006): 207/206 = 3227.2 ± 0.2 Ma (CA-ID-TIMS). |
| Data processing package used | in-house EXCEL spreadsheet (Gerdes and Zeh, 2006, 2009) |
| Quality control / Validation | BB (session 1): concordia age = 562.0 ± 1.6 Ma (n = 28) Plešovice (session1): concordia age = 338.43 ± 0.98 Ma (n = 28) KA1(session 1): 207/206 age = 3225.8 ± 7.5 Ma (n = 19) BB (session 2): concordia age = 562.3 ± 1.6 Ma (n = 28) Plešovice (session 2): concordia age = 337.76 ± 0.94 Ma (n = 28) KA1(session 2): 207/206 age = 3230.1 ± 7.8 Ma (n = 19) |

- Measuring session Nov_2024

| Laboratory & Sample Preparation |  |
| --- | --- |
| Laboratory name | Institute for Applied Geosciences, KIT (Karlsruhe Institute of Technology), Germany |
| Sample type/mineral | detrital zircons |
| Sample preparation | Conventional mineral separation, 1 inch resin mount, polish to finish |
| Imaging | VEGA TESCAN with Oxford detector, BSE mode 15 kV, <10nA (at KIT) |
| Laser ablation system |  |
| Make, Model & type | Teledyne Photon Machines, Analyte Excite+ (Excimer) |
| Ablation cell | Two-volume ablation cell (HELEX 100, EQC COMP), ANUAustralia) |
| Laser wavelength | 193 nm |
| Pulse width | < 5 ns |
| Fluence | 2.1 J/cm-2 |
| Repetition rate | 10 Hz |
| Spot size | 20 µm |
| Sampling mode / pattern | single spot |
| Carrier gas | He(cell) 0.30 l/min, He (cup) 0.20 l/min, 0.96 l/min Ar, N2 (11.5 ml) |
| Background collection | 15 seconds |
| Ablation duration | 15 seconds |
| Pre-ablation | 3 pulses |
| Wash-out delay | 15 seconds |
| Cell carrier gas flow (He) | 0.50 l/min (total) |
| ICP-MS Instrument |  |
| Make, Model & type | Thermo-Scientific ELEMENT XR (sector field) |
| Sample introduction | via conventional tubing |
| RF power | 1250 W |
| Sampler, skimmer cones | Ni-X-cone |
| Extraction lenses | X - type |
| Make-up gas flow (Ar) | 0.91 l/min |
| Detection system | single collector secondary electron multiplier |
| Data acquisition protocol | Time-resolved analysis |
| Scanning mode | Peak hopping, four point per peak |
| Detector mode | Pulse counting mode |
| Masses measured | 202Hg,204(Hg + Pb), 206Pb, 207Pb, 208Pb, 232Th, 238U |
| Integration time per peak | 10 milliseconds |
| Sensitivity / Efficiency | 36850 (cts/ppm) (238U, NIST612, spot-20 μm, 10μm/s, 10 Hz, 5 J/cm2) oxide formation rate: U/UO <0.1, Th/U = 0.90 |
| Data Processing |  |
| Gas blank | 15 seconds on-peak |
| Calibration strategy | BB1 zircon standard used as primary reference material, Plešovice, KA1 (Kaap Valley tonalite) as secondary reference material (quality control) |
| Reference Material info | BB (Santos et al., 2017): concordia age = 562 ± 0.6 Ma (LA-ICP-MS) Plešovice (Sláma et al., 2008) = 337.13 ± 0.37 Ma (TIMS) KA1 (Schoene et al. (2006): 207/206 = 3227.2 ± 0.2 Ma (CA-ID-TIMS). |
| Data processing package used | in-house EXCEL spreadsheet (Gerdes and Zeh, 2006, 2009) |
| Quality control / Validation | BB (session 1): concordia age = 562.2 ± 1.5 Ma (n = 28) Plešovice (session1): concordia age = 337.4 ± 1.2 Ma (n = 28) KA1(session 1): 207/206 age = 3229.5± 8.4 Ma (n = 18) |

References

Gerdes A, Zeh A (2006) Combined U–Pb and Hf isotope LA-(MC-)ICP-MS analyses of detrital zircons: Comparison with SHRIMP and new constraints for the provenance and age of an Armorican metasediment in Central Germany. Earth Planet Sci Lett 249:47–61

Gerdes A, Zeh A (2009) Zircon formation versus zircon alteration – new insights from combined U–Pb and Lu–Hf in-situ LA-ICP-MS analyses, and consequences for the interpretation of Archean zircon from the Central Zone of the Limpopo Belt. Chem Geol 261:230–243

Santos MM, Lana C, Scholz R, Buick I, Schmitz MD, Kamo SL, Gerdes A, Corfu F, Tapster S, Lancaster P, Storey CD, Basei MAS, Tohver E, Alkmim A, Nalini H, Krambrock K, Fantini C, Wiedenbeck M (2017) A new appraisal of Sri Lankan BB zircon as a reference material for LA–ICP–MS U–Pb geochronology and Lu–Hf isotope tracing. Geostand Geoanalyt Res 41:335–358

Sláma J, Kŏsler J, Condon DJ, Crowley JL, Gerdes A, Hanchar JM, Horstwood MSA, Morris GA, Nasdala L, Norberg N, Schaltegger U, Schoene B, Tubrett MN, Whitehouse MJ (2008) Plešovice zircon — A new natural reference material for U-Pb and Hf isotopic microanalysis. Chem Geol 249:1–35

Schoene B, Crowley JL, Condon DJ, Schmitz MD, Bowring SA (2006) Reassessing the uranium decay constants for geochronology using ID-TIMS U-Pb data. Geochim et Cosmochim Acta 70:426–445