Supplementary Material:

User Questionnaire For Evaluation of Hybrid Exoscope Scenario

A During Surgery

- How well does the position of the microscope allow the main surgeon free working space to handle surgical instruments and perform surgery in good view of the monitor?
□ excellent	□ good       □ acceptable      □ not acceptable

- How reliably can the main surgeon have eye contact with the entire surgical team?
□ excellent	□ good       □ acceptable      □ not acceptable

- How well does the position of the microscope provide the surgical assistant free working space and adequate visualization to assist surgery?
□ excellent	□ good       □ acceptable      □ not acceptable

- How well does the position of the microscope provide the sterile nurse free working space and adequate visualization to assist surgery?
□ excellent	□ good       □ acceptable      □ not acceptable

- How safely can the range and pace of the robotic microscope arm be safely adjusted in and around the surgical field by the main surgeon?

□ excellent	□ good       □ acceptable      □ not acceptable

- How well does the digital approach allow for additional freedom for the main surgeon including utilizing previously restricted microscope orientations?

□ excellent	□ good       □ acceptable      □ not acceptable

- How well could the main surgeonbenefit from utilization of autofocus functionality?

□ excellent	□ good       □ acceptable      □ not acceptable

- How well does the magnification functionality allow for a precise and comfortable assessment of the surgical field by the main surgeon?

□ excellent	□ good       □ acceptable      □ not acceptable

- How well does the digital visualization provide sharp tissue-tissue/tissue-pathology contrast to the main surgeon?

□ excellent	□ good       □ acceptable      □ not acceptable

- How acceptable is the frequency of alternate viewing directions (between site and monitor or additional screens) by the main surgeon when using the digital monitor image?

□ excellent	□ good       □ acceptable      □ not acceptable

B After Surgery

- How well can the surgical assistant share the monitor image with the main surgeon to assist the surgery?

□ excellent	□ good       □ acceptable      □ not acceptable

- How well can the surgical assistant utilize the system monitor at the microscope stand to assist surgery?

□ excellent	□ good       □ acceptable      □ not acceptable

- How reliably does the OR setup provide suitable position and space for the anesthesiologist?

□ excellent	□ good       □ acceptable      □ not acceptable

- How easily can the microscope and video monitor be integrated into the standard OR setup?

□ excellent	□ good       □ acceptable      □ not acceptable

- How effortlessly can the microscope and video monitor be positioned according to patient positioning?

□ excellent	□ good       □ acceptable      □ not acceptable

- How well can the sterile working zones of main surgeon, assistant surgeon and sterile nurse be integrated with microscope and monitor positioning?

□ excellent	□ good       □ acceptable      □ not acceptable

- How suitable are the estimated technical and workflow team training requirements regarding microscope and monitor for surgical routine?

□ excellent	□ good       □ acceptable      □ not acceptable

- How well does the position of the microscope allow the main surgeon to adjust the microscope to any needed position or orientation to observe the surgical site?
□ excellent	□ good       □ acceptable      □ not acceptable

- How precisely and efficiently can the main surgeon complete the surgical task?
□ excellent	□ good       □ acceptable      □ not acceptable

- How realiably does the digital visualization of the monitor image allow for precise hand-eye coordination of the main surgeon?
□ excellent	□ good       □ acceptable      □ not acceptable

- How realistic is the 3D impression provided by the digital visualization for the main surgeon?
□ excellent	□ good       □ acceptable      □ not acceptable

- How comfortable is the ergonomic position of the main surgeon and how reliably does it allow for working at high concentration on the surgical procedure?
□ excellent	□ good       □ acceptable      □ not acceptable

- How reliably does the digital visualization provide a real-time iamge without noticeable latency for the main surgeon? 
□ excellent	□ good       □ acceptable      □ not acceptable

- How well does the digital visualization allow the main surgeon to perform surgical tasks without noticeable fatigue or other negative impact, such as e.g. nausea?

□ excellent	□ good       □ acceptable      □ not acceptable

- How well can the main surgeon concentrate on the surgical task without distraction or disturbance from the OR setting and activities of the surgical team?

□ excellent	□ good       □ acceptable      □ not acceptable

- How good is the experience of the main surgeon regarding movement hesitations and/or corrections when using the digital monitor image?

□ excellent	□ good       □ acceptable      □ not acceptable

- How good is the experience of the main surgeon regarding restrictions in range and speed of instrument handling when using the digital monitor image?

□ excellent	□ good       □ acceptable      □ not acceptable

- How precisely and effortlessly can the main surgeon utilize the foot control panel to adjust the microscopic setup, including positioning and focusing? 
□ excellent	□ good       □ acceptable      □ not acceptable

- How would you as the main surgeon rate the necessity to visually control your performance with the naked eye when using the digital monitor image?

□ excellent	□ good       □ acceptable      □ not acceptable

- How much can using Hybrid Digital Visualization be recommended for this surgical approach/test scenario?

□ excellent	□ good       □ acceptable      □ not acceptable

Supplementary Material

Table 3: User Questionnaire For Evaluation of Hybrid Exoscope Scenario: Overview of Results

| Item No. | Short Description | Median Rating Results |
| --- | --- | --- |
| 1 | working space S1 | 1 |
| 2 | eye contact | 1 |
| 3 | working space S2 | 2 |
| 4 | working space N1 | 1 |
| 5 | robotic safety | 2 |
| 6 | microscope orientations S1 | 1,5 |
| 7 | autofocus S1 | 2 |
| 8 | magnification S1 | 1,5 |
| 9 | image contrast S1 | 1,5 |
| 10 | viewing directions S1 | 1,5 |
| 11 | use of 55” monitor S2 | 2 |
| 12 | use of small monitor S2 | 4 |
| 13 | working space A1 | 1,5 |
| 14 | integration into OR setup | 2 |
| 15 | patient position | 2 |
| 16 | sterility | 2 |
| 17 | OR workflow | 2 |
| 18 | microscope adjustments S1 | 1 |
| 19 | task completion S1 | 1 |
| 20 | hand-eye coordination S1 | 1 |
| 21 | 3D image S1 | 2 |
| 22 | ergonomics S1 | 1 |
| 23 | transmission delay S1 | 1 |
| 24 | fatigue S1 | 1 |
| 25 | distraction S1 | 2 |
| 26 | hesitations/corrections S1 | 2 |
| 27 | instrumentation S1 | 2 |
| 28 | FCP S1 | 2 |
| 29 | visual control S1 | 1,5 |
| 30 | overall satisfaction S1 | 1,5 |

abbreviations: A1 = anesthesiologist, FCP = foot control panel, N1 = sterile nurse, OR = operating room, S1 = main surgeon, S2 = assistant surgeon

Supplementary Material

Table 4: User Questionnaire For Evaluation of Hybrid Exoscope Scenario: Free Text Comments

| Item No. | advantages of hybrid exoscope | disadvantages of hybrid exoscope |
| --- | --- | --- |
| 1 | - | - |
| 2 | - improved team communication | - |
| 3 | - | - 90° position and view are problematic - monitor view restricted by obstacles - due to discrepancy of working and viewing angle - increased impact on hand-eye coordination - face-to-face setups currently not supported in 3D |
| 4 | - | - |
| 5 | - | - frequent collisions with head position |
| 6 | - improved ergonomics - in general comparable to ocular-based surgery | - brainstem approaches and other deep-seated lesions potentially difficult - only if using long instruments |
| 7 | - to date, only infrequently used | - |
| 8 | - | - loss of sharpness in high magnification |
| 9 | - | - ocular vision is superior |
| 10 | - parallel use of ocular and monitor image if required or convenient | - |
| 11 | - | - assistant has to be placed opposite to monitor - very limited view |
| 12 | - | - |
| 13 | - | - |
| 14 | - more variable settings | - monitor placement should be more flexible - ceiling-mounted solution preferable |
| 15 | - more variable angulations |  |
| 16 |  | - depending on distance of camera to site - risk of head collision |
| 17 | - | - |
| 18 | - adjustments can be performed more smoothly | - |
| 19 | - instrument handling facilitated by using foot control panel for camera positioning - broader area of view in monitor-based surgery | - except for detailed microneurosurgical tasks - depending on access route |
| 20 | - | - manipulating suture easier when using ocular |
| 21 | - | - restricted depth perception - field of view smaller than in ocular-based surgery - in small structures, ocular vision superior - working off-focus more cumbersome than in ocular-based surgery |
| 22 | - | - |
| 23 | - | - |
| 24 | - (comparable to ocular-based surgery) | - |
| 25 | - monitor size sufficient | - distraction caused by team members - (monitor position should be close to main surgeon) |
| 26 | - | - initial adjustment/training period required |
| 27 | - | - alternate views between monitor and surgical site |
| 28 | - easy to learn - hand-free control of visualization facilitates surgery - good grip and stability | - shoeless interaction recommended - too many buttons - too unprecise - handles preferable - additional training/learning curve required - potential conflicts with additional foot panels - waterproof cover required |
| 29 | - | - |
| 30 | - hybrid tool - suitable for routine cases - improved ergonomics - shared view for whole team | - restricted space in OR for circulating nurse |