## Analytical conditions - details
SupplTable S1. Spectrometer settings for carbonate EPMA analysis
Element | X-ray | Crystal | L-value Peak (mm) | Lower back. (mm) | Upper back. (mm)
Si | Ka | TAP | 77.782 | 6 | 5
Ca | Ka | PETJ | 107.507 | 5 | 6
Fe | Ka | LIFH | 134.718 | 5 | 5
Mn | Ka | LIFH | 146.24 | 5 | 5
Mg | Ka | TAP | 107.497 | 6 | 6.5
Sr | La | TAP | 74.664 | 0.4 | 1.2
Ba | La | PETL | 88.894 | 2 | 3
SupplTable S2. Spectrometer settings for silicates EPMA analysis
Element | X-ray | Crystal | L-value Peak (mm) | Lower back. (mm) | Upper back. (mm)
Na | Ka | TAP | 129.348 | 5 | 5
Si | Ka | TAP | 77.788 | 6 | 6
K | Ka | PETJ | 119.768 | 5 | 5.5
Ca | Ka | PETJ | 107.531 | 5 | 6
Fe | Ka | LIFH | 134.718 | 5 | 5
Ti | Ka | LIFH | 191.188 | 5 | 5
Ni | Ka | LIFH | 115.31 | 5 | 5
F | Ka | TAP | 199 | 5 | 5
Al | Ka | TAP | 90.677 | 5.5 | 5
Mg | Ka | TAP | 107.507 | 5 | 6
Cr | Ka | LIFL | 159.251 | 5 | 5
Mn | Ka | LIFL | 146.203 | 5 | 5
SupplTable S3. Detector settings, error, detection limit, standards for carbonate EPMA analysis
| Detector settings | [7 empty] | Error, Detection limits, Standards
Element | Counting Peak (sec) | Counting Back. (sec) | Peak Search | Gain | High.V | Base.L | Window.W | Mode | Average Error (1σ) (%) | Average DL (ppm oxide) | Standards
Si | 10 | 5 | 1 | 32 | 1694 | 0.7 | 0 (V) | Int | > 100 | 131.42 | Olivine_SPI
Ca | 10 | 5 | 1 | 64 | 1677 | 0.7 | 0 (V) | Int | 0.23 | 89.43 | Calcite_SPI
Fe | 10 | 5 | 1 | 32 | 1678 | 0.7 | 0 (V) | Int | > 100 | 57.77 | Olivine_SPI
Mn | 10 | 5 | 1 | 32 | 1678 | 0.7 | 0 (V) | Int | > 100 | 40.72 | Rhodonite_SPI
Mg | 10 | 5 | 1 | 32 | 1684 | 0.7 | 9.3 (V) | Dif | 4.14 | 33.68 | Dolomite_SPI
Sr | 30 | 20 | 1 | 32 | 1684 | 0.7 | 9.3 (V) | Dif | > 100 | 53.49 | Celestite_SPI
Ba | 30 | 20 | 1 | 16 | 1700 | 0.5 | 9.3 (V) | Dif | > 100 | 91.55 | Benitoite_SPI
SupplTable S4. Detector settings, error, detection limit, standards for silicate EPMA analysis
| Detector settings | [7 empty] | Error, Detection limits, Standards
Element | Counting Peak (sec) | Counting Back. (sec) | Peak Search | Gain | High.V | Base.L | Window.W | Mode | Average Error (1σ) (%) | Average DL (ppm oxide) | Standards
Na | 10 | 5.0 (s) | 1 | 32 | 1694 | 0.7 | 0 (V) | Int | 5.6 | 185 | Jadeite_SPI
Si | 10 | 5.0 (s) | 1 | 32 | 1694 | 0.7 | 0 (V) | Int | 0.34 | 212.75 | Cr-diopsoide_SPI
K | 10 | 5.0 (s) | 1 | 64 | 1677 | 0.7 | 0 (V) | Int | 0.95 | 104.5 | Orthoclase_SPI
Ca | 10 | 5.0 (s) | 1 | 64 | 1677 | 0.7 | 0 (V) | Int | 17.55 | 123 | Plagioclase_SPI
Fe | 10 | 5.0 (s) | 1 | 32 | 1678 | 0.7 | 0 (V) | Int | 19.61 | 207.5 | Olivine_SPI
Ti | 10 | 5.0 (s) | 1 | 32 | 1678 | 0.7 | 0 (V) | Int | 137.55 | 199.5 | Rutile_SPI
Ni | 10 | 5.0 (s) | 1 | 32 | 1678 | 0.7 | 0 (V) | Int | 100 | 253.25 | Pentlandite_SPI
F | 10 | 5.0 (s) | 1 | 32 | 1684 | 0.7 | 9.3 (V) | Dif | 40.63 | 914 | Apatite_SPI
Al | 10 | 5.0 (s) | 1 | 32 | 1684 | 0.7 | 9.3 (V) | Dif | 0.36 | 151.25 | Plagioclase_SPI
Mg | 10 | 5.0 (s) | 1 | 32 | 1684 | 0.7 | 9.3 (V) | Dif | 2.05 | 134.75 | Cr-diopsoide_SPI
Cr | 10 | 5.0 (s) | 1 | 32 | 1700 | 0.5 | 9.3 (V) | Dif | 136.03 | 150 | Chromite_SPI
Mn | 10 | 5.0 (s) | 1 | 32 | 1700 | 0.5 | 9.3 (V) | Dif | 193.89 | 154.75 | Rhodonite_SPI