Supplementary Table 3. Mean CT attenuation pairwise differences (presented as absolute values) between all combinations of the four image reconstruction types.

|  | ASiR-V – DLIR-L | ASiR-V –DLIR-M | ASiR-V –DLIR-H | DLIR-L –DLIR-M | DLIR-L –DLIR-H | DLIR-M –DLIR-H | P value |
| --- | --- | --- | --- | --- | --- | --- | --- |
| CT ATTENUATION (HU) |  |  |  |  |  |  |  |
| Thalamic GM | 0.86±0.61bcd | 0.85±0.61bcd | 0.80±0.61bcd | 0.33±0.24 | 0.36±0.33 | 0.27±0.23 | < 0.001 |
| PLIC WM | 0.79±0.69abd | 0.85±0.72abcd | 1.03±0.77bcd | 0.28±0.25c | 0.44±0.29 | 0.33±0.21 | < 0.001 |
| M5 GM | 0.73±0.71abd | 0.73±0.72abd | 1.00±0.82bcd | 0.35±0.25c | 0.54±0.38d | 0.39±0.33 | < 0.001 |
| CSO WM | 0.52±0.45abd | 0.57±0.48abd | 0.75±0.53bcd | 0.27±0.25 | 0.40±0.31d | 0.29±0.25 | < 0.001 |

CT ATTENUATION (HU) mean absolute CT attenuation difference in Hounsfield units, GM gray matter, WM white matter, PLIC posterior limb of the internal capsule, M5 M5 cortex region (lateral MCA territory) according to Alberta Stroke Program Early CT Score – ASPECTS, CSO centrum semiovale, ASiR-V adaptive statistical iterative reconstruction-Veo, DLIR-L deep learning-based image reconstruction low strength level, DLIR-M deep learning-based image reconstruction medium strength level, DLIR-H deep learning-based image reconstruction high strength level

Post hoc pairwise multiple comparisons procedure with the Dunn-Bonferroni test showed a statistically significant (P <0.05) difference between means when compared with pairwise difference group ‘ASiR-V–DLIR-M’ (*), ‘ASiR-V–DLIR-H’ (a), ‘DLIR-L–DLIR-M’ (b), ‘DLIR-L–DLIR-H’ (c), and ‘DLIR-M–DLIR-H’ (d)