Online supplemental- survey 1

| Question | Answers |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Patient preparation |  |  |  |  |  |  |  |  |
| 2.1 Fasting before scanning? | Yes 3 | No 3 | Don’t’ know 0 | No recommendation 2 |  |  | Other 1 |  |
| 2.2 Restriction on liquids before scanning? | Yes 3 | No 1 | Don’t’ know 1 | No recommendation 1 |  |  | Other 3 |  |
| 2.3 Any other patient preparation procedure? | Yes 0 | No 4 | Don’t’ know 0 | No recommendation 2 |  |  | Other 3 |  |
| Other topics to be included on patient preparation? | Open answers |  |  |  |  |  |  |  |
| Hardware |  |  |  |  |  |  |  |  |
| 3.1 Field strength? | 1.5T 0 | 3T 2 | Both 6 | Other 1 |  |  |  |  |
| 3.2 Transmitter coil? | Body 9 |  |  | No recommendation 1 |  |  | Other 0 |  |
| 3.3 Receiver coil? | Body 8 |  |  | No recommendation 1 |  |  | Other 0 |  |
| Other topics to be included on hardware? | Open answers |  |  |  |  |  |  |  |
| T1 mapping scheme |  |  |  |  |  |  |  |  |
| 4.1 T1 mapping scheme? | Classic IR 1 |  | LL-variant 5 | VFA 1 | SASHA 1 |  | No recommendation 1 |  |
| Classic IR |  |  |  |  |  |  |  |  |
| 5.1 Inversion pulse type? | FOCI 1 | Hyperbolic secant 1 |  | No recommendation 4 |  |  | Do not know 3 |  |
| 5.2 Number of inversion times to use ? [choose a single option – state inversion times and field strength used in ‘other’] | 5 3 | 11 1 | 13 1 | No recommendation 2 |  |  | Do not know 1 |  |
| 5.3 Breathing scheme? | BH 1 | RT 1 | FB 1 | No recommendation 3 |  |  | Do not know 1 | Other 2 |
| 5.4 Image readout type? | 2D-multisclice 2 |  | 3D 2 | No recommendation 4 |  |  | Do not know 1 |  |
| 5.5 Slice orientation? | Cor/Cor-oblique 5 | Sag 0 | Tra 0 | No recommendation 2 |  |  | Do not know 1 | Other 1 |
| 5.6 Readout sequence for 2D? | SE-EPI 2 | GRE-EPI 1 | bSSFP 0 | No recommendation 4 |  |  | Do not know 1 |  |
| 5.7 Number of slices in 2D? | 2-4 1 | 4-6 2 | >10 1 | No recommendation 3 |  |  | Do not know 2 |  |
| 5.8 Slice thickness for 2D? | 5mm 3 | 6mm 2 | 8mm 1 | No recommendation 3 |  |  | Do not know 2 |  |
| 5.9 Slice gap in 2D multi-slice? | 0.5-1mm 1 |  | No gap 2 | No recommendation 4 |  |  | Do not know 2 |  |
| 5.10 Readout sequence for 3D? | 3D-RARE 1 | 3D-GRASE 1 |  | No recommendation 5 |  |  | Do not know 3 |  |
| 5.11 Use segmented 3D sequences? | Yes 1 | No 0 | Don’t’ know 3 | No recommendation 5 |  |  |  |  |
| 5.12 Number of slices in 3D? | 5-10 slices 0 | 10-15 0 | 15-20 slices 2 | No recommendation 4 |  |  | Don’t’ know 3 |  |
| 5.13 Slice thickness for 3D? | 3mm 1 | 4mm 1 | Other 0 | No recommendation 4 |  |  | Don’t’ know 3 |  |
| 5.14 Required FOV of image? | 244x244 1 | 300x300 1 | 400x 400 2 | No recommendation 4 |  |  | Don’t’ know 1 |  |
| 5.15 In-plane resolution? | 2mm 1 | 3mm 3 | Other 0 | No recommendation 4 |  |  | Don’t’ know 1 |  |
| 5.16 Partial Fourier phase-encoding? | None 1 | 80-90% 1 | Other 0 | No recommendation 6 |  |  | Don’t’ know 1 |  |
| 5.17 Partial Fourier slice (partition)-encoding in 3D? | 60-70% 0 | 70-80% 0 | 80-90% 0 | No recommendation 6 |  |  | Don’t’ know 3 |  |
| 5.18 Use parallel imaging? [choose a single option] | Yes 3 | No 0 | Don’t’ know 1 | No recommendation 5 |  |  |  |  |
| 5.19 Parallel imaging acceleration factor? | 1.5 1 | 2 2 | 3 0 | No recommendation 5 |  |  | Don’t’ know 1 |  |
| 5.20 TR (between inversion pulses)? | 4-5s 1 | 6-7s 3 | 8-9s 0 | No recommendation 4 |  |  | Don’t’ know 1 |  |
| Other topics relevant to the use of a Classic Inversion Recovery Scheme? | Open answers |  |  |  |  |  |  |  |
| MOLLI variant |  |  |  |  |  |  |  |  |
| 6.1 Inversion pulse type? | FOCI 0 | Hyperbolic secant 3 |  | No recommendation 2 |  |  | Do not know 4 |  |
| 6.2 Look Locker variant to use | 5(3)3 3 | 5(1)1(1)1 1 | 16-8-4 1 | No recommendation 2 |  |  | Do not know 3 |  |
| 6.3 Do you ECG trigger or use physiological simulation? | ECG 2 | Simulaton 60 bpm 3 | Other 1 | No recommendation 1 |  |  | Do not know 3 |  |
| 6.4 What MOLLI flip angle should be used? | 12° 1 | 35° 1 | Other 0 | No recommendation 2 |  |  | Do not know 3 |  |
| 6.5 Slice orientation? | Cor/Cor-oblique 7 | Sag 0 | Tra 0 | No recommendation 1 |  |  | Do not know 1 |  |
| 6.6 Required FOV of image? | 300x300 1 | 400x400 4 | Other 0 | No recommendation 3 |  |  | Do not know 1 |  |
| 6.7 In-plane resolution? | 1mm 1 | 2mm 3 | 3mm 2 | No recommendation 2 |  |  | Do not know 1 |  |
| 6.8 Breathing scheme? | BH 4 | FB 0 | Other 0 | No recommendation 2 |  |  | Do not know 3 |  |
| 6.9 Is Minimization of off-resonance effects performed? | Main field shim and center frequency 0 |  |  | No recommendation 3 |  |  | Do not know 5 |  |
| 6.10 Is the MOLLI scheme repeated for a number of slices? | 1 1 | 2 1 | 3-5 2 | No recommendation 3 |  |  | Do not know 4 |  |
| VFA variant |  |  |  |  |  |  |  |  |
| 7.1 How many flip angles do you use? | 2 2 | 5 1 | Other 0 | No recommendation 3 |  |  | Do not know 3 |  |
| 7.2 Slice orientation? | Cor/Cor-oblique 4 | Sag 0 | Tra 0 | No recommendation 2 |  |  | Do not know 3 |  |
| 7.3 Required FOV of image? | 380x380 1 | 400x400 1 |  | No recommendation 4 |  |  | Do not know 3 |  |
| 7.4 In-plane resolution? | 1mm 1 | 2mm 1 | 3mm 1 | No recommendation 2 |  |  | Do not know 4 |  |
| 7.5 Breathing scheme? | BH 3 | FB 0 | Other 0 | No recommendation 2 |  |  | Do not know 4 |  |
| Other topics relevant to the use of a VFA variant? | Open answers |  |  |  |  |  |  |  |
| B0 & B1 map sequence |  |  |  |  |  |  |  |  |
| 8.1 Do you collect a separate B0 and B1 map sequence? | Yes 3 | No 4 | Other 1 | No recommendation 0 |  |  | Do not know 1 |  |
| 8.2 Which B0 mapping scheme? | Dual-echo GRE 2 | 6-echo Dixon 1 |  | No recommendation 3 |  |  | Do not know 3 |  |
| 8.2 Which B1 mapping scheme? | DREAM 1 | Dual TR-AFI 1 | Other 0 | No recommendation 4 |  |  | Do not know 3 |  |
| 8.3 Slice orientation? | Cor/Cor-oblique 5 | Sag 0 | Tra 0 | No recommendation 1 |  |  | Do not know 3 |  |
| 8.4 Breathing scheme? | BH 3 | FB 0 | Other 1 | No recommendation 2 |  |  | Do not know 3 |  |
| 8.5 Motion compensation? | None 1 1 | Expiration respiratory triggering 2 | Navigators 2 | No recommendation 1 |  |  | Do not know 3 |  |
| Other topics relevant to the use of a B1 map sequence? | Open answers |  |  |  |  |  |  |  |
| Data pre-processing |  |  |  |  |  |  |  |  |
| 9.1 Motion correction? | None 3 | Rigid 1 | Affine 3 | Non-rigid 2 |  | No recommendation 0 | Do not know 0 |  |
| 9.2 Outlier detection and rejection? | Yes 3 | No 4 | No recommendation 1 |  |  |  | Do not know 1 |  |
| Other topics relevant to data preprocessing? | Open answers |  |  |  |  |  |  |  |
| 10 Is your T1 mapping protocol also used for quantification of ASL | Yes 4 | No 5 |  |  |  |  |  |  |
| Other topics relevant to collecting T1 mapping for quantification of ASL? | Open answers |  |  |  |  |  |  |  |
| Quantification of T1 values |  |  |  |  |  |  |  |  |
| 11.1 Do you use B1 mapping in T1 quantification? | Yes 1 | No 6 | Do not know 0 | No recommendation 1 |  |  | Other 1 |  |
| 11.2 Do you use an inversion factor correction in your T1 quantification? | Yes 1 | No 5 | No recommendation 3 |  |  |  | Do not know 0 |  |
| 11.3 Which method do you use for MOLLI T1 quantification? [choose a single option]? | Original fit 4 |  | Other 0 | No recommendation 2 |  |  | Do not know 3 |  |
| Other topics relevant to quantification of T1 values? | Open answers |  |  |  |  |  |  |  |
| Reporting T1 values |  |  |  |  |  |  |  |  |
| 12.1 Region of interest selection method (cortex, and medulla)? | Manual 5 | TLCO 1 | k-means 1 | Other 1 |  |  | No recommendation 1 |  |
| 12.2 Region of interest measurements? | Mean+SD 5 | Mode+FWHM 1 | Other 1 | No recommendation 1 |  |  | Do not know 1 |  |
| 12.3 Measurements for corticomedullary differentiation ? | C-M ratio 3 | C-M difference 1 | 100x (C-M)/C 1 | No recommendation 2 |  |  | Do not know 2 |  |
| 13.1 Please enter any comments on the questionnaire or suggestions (e.g. additional questions) | Open answers |  |  |  |  |  |  |  |