| Sample Name | T1 (s) |  |  | T2 (s) |  |
| --- | --- | --- | --- | --- | --- |
|  | NMR | MRI - In vivo protocol | MRI - Reference protocol | NMR | MRI - In vivo protocol |
| 0.025 mmol/L GdGl3-EDTA | 1.246 (0.0146) | 1.384 (0.0325) | 1.342 (0.0079) | 1.272 (0.0326) | 0.901 (0.0211) |
| 0.1% agarose mass concentration (%) | 2.241 (0.0096) | 2.392 (0.0673) | 2.229 (0.1041) | 0.604 (0.004) | 0.597 (0.0154) |
| 0.1 mmol/L CuSO4 | 1.695 (0.0188) | 1.743 (0.0493) | 1.731 (0.0862) | 1.901 (0.3131) | 1.007 (0.0346) |
| 0.1 mmol/L GdGl3-EDTA in 1.5% agarose mass concentration (%) | 0.436 (0.0042) | 0.438 (0.0071) | 0.438 (0.0038) | 0.068 (0.0007) | 0.056 (0.001) |
| 0.5% agarose mass concentration (%) | 1.846 (0.0101) | 2.027 (0.1016) | 1.947 (0.0798) | 0.244 (0.0012) | 0.183 (0.0015) |
| 0.75 mmol/L CuSO4 in 0.25% agarose mass concentration (%) | 0.476 (0.0029) | 0.517 (0.0076) | 0.482 (0.0152) | 0.245 (0.0017) | 0.207 (0.0018) |
| 1 mmol/L CuSO4 in 1% agarose mass concentration (%) | 0.315 (0.0044) | 0.341 (0.0059) | 0.326 (0.0041) | 0.094 (0.0044) | 0.088 (0.0025) |
| 2 mmol/L CuSO4 | 0.235 (0.0035) | 0.222 (0.0027) | 0.219 (0.0024) | 0.222 (0.0026) | 0.216 (0.0045) |
| 4 mmol/L NiCl2 in 1.2% agarose mass concentration (%) | 0.328 (0.0074) | 0.335 (0.0034) | 0.329 (0.0033) | 0.087 (0.0005) | 0.078 (0.0028) |
| Deionized water | 2.475 (0.0223) | 2.555 (0.1022) | 2.351 (0.1474) | 2.193 (0.0829) | 1.21 (0.0417) |

Table S.1 T1 and T2 mean and standard deviations for test samples measured for the NMR system, for the MRI system and the in vivo protocols. Also included for T1 are measurements for the MRI system reference protocol.

Fig. S.1 a) Example IR raw data for one subject. b) Example T1 map for the subject in (a), and c) shows the result of the skull stripping protocol. d) Example T2 map for a different volunteer, with e) the result of the skull stripping protocol.

| Subject | T1 (s) |  |  |  |  |  | T2 (s) |  |  |  |  | Voxels |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | WM |  | GM |  | CSF |  | WM | WM & GM | GM | CSF |  |  |
|  | Manual | Auto | Manual | Auto | Manual | Auto | Manual | Auto | Manual | Manual | Auto |  |
| 1 | 0.234 (0.01) | 0.284 (0.0339) | 0.409 (0.0155) | 0.389 (0.0411) | 1.386 (0.2561) | 1.207 (0.201) | 0.079 (0.0057) | 0.097 (0.0122) | 0.102 (0.0074) | 1.134 (0.0645) | 0.567 (0.2338) | 8782 |
| 2 | 0.259 (0.0071) | 0.27 (0.0317) | 0.38 (0.0049) | 0.374 (0.0394) | 1.107 (0.1338) | 1.089 (0.2205) | 0.082 (0.0027) | 0.097 (0.0132) | 0.116 (0.0049) | 1.104 (0.1251) | 0.504 (0.1973) | 8982 |
| 3 | 0.248 (0.0206) | 0.272 (0.0242) | 0.358 (0.0051) | 0.342 (0.0226) | 0.944 (0.1485) | 0.694 (0.1436) | 0.084 (0.0029) | 0.095 (0.0106) | 0.103 (0.0047) | 0.422 (0.0767) | 0.346 (0.1176) | 6917 |
| 4 | 0.243 (0.0179) | 0.293 (0.0387) | 0.395 (0.01) | 0.418 (0.0425) | 2.173 (0.4374) | 1.445 (0.4664) | 0.081 (0.0035) | 0.099 (0.0136) | 0.167 (0.0367) | 1.13 (0.1247) | 0.43 (0.1902) | 7453 |
| 5 | 0.266 (0.0073) | 0.298 (0.0351) | 0.434 (0.0262) | 0.406 (0.038) | 1.203 (0.139) | 2.054 (0.6738) | 0.085 (0.004) | 0.1 (0.0126) | 0.103 (0.0046) | 1.185 (0.0678) | 0.514 (0.2139) | 7631 |
| 6 | 0.296 (0.0132) | 0.326 (0.0534) | 0.351 (0.0127) | 0.738 (0.1946) | 4.43 (1.2622) | 4.716 (0.7551) | 0.082 (0.0051) | 0.097 (0.0154) | 0.085 (0.0011) | 1.655 (0.0804) | 0.779 (0.3398) | 7597 |
| 7 | 0.251 (0.008) | 0.298 (0.0316) | 0.371 (0.0074) | 0.39 (0.0349) | 1.174 (0.1456) | 1.027 (0.13) | 0.081 (0.0032) | 0.097 (0.0098) | 0.098 (0.0038) | 1.115 (0.0484) | 0.41 (0.2004) | 8387 |
| 8 | 0.25 (0.0077) | 0.312 (0.0557) | 0.389 (0.0287) | 0.588 (0.1052) | 2.81 (0.5314) | 1.938 (0.4432) | 0.077 (0.0074) | 0.096 (0.012) | 0.089 (0.0054) | 1.166 (0.1164) | 0.579 (0.2491) | 8676 |
| 9 | 0.237 (0.0086) | 0.28 (0.0356) | 0.384 (0.0121) | 0.399 (0.0435) | 1.831 (0.1938) | 1.317 (0.3497) | 0.08 (0.0019) | 0.099 (0.0142) | 0.106 (0.0023) | 1.294 (0.0317) | 0.643 (0.2597) | 8880 |
| 10 | 0.258 (0.0133) | 0.305 (0.0484) | 0.305 (0.0029) | 0.555 (0.1107) | 3.674 (1.326) | 3.057 (0.8628) | 0.084 (0.0059) | 0.098 (0.0145) | 0.082 (0.0023) | 1.516 (0.2218) | 0.753 (0.3107) | 8309 |
| All | 0.254 (0.0177) | 0.294 (0.0179) | 0.377 (0.0351) | 0.46 (0.1258) | 2.073 (1.2007) | 1.854 (1.2097) | 0.081 (0.0024) | 0.097 (0.0016) | 0.105 (0.0241) | 1.172 (0.3234) | 0.553 (0.1422) | 81614 |

Table S.2 Per-subject manual and automatic segmentation T1 and T2 results, with standard deviations shown in parentheses. The number of voxels per slice that were considered within the skull are given. Averages over all subjects are given, along with the total number of voxels measured. Note that for the T2 auto-segmentation results, WM and GM were indistinguishable.