Supplementary information

Deducing sample activity from the radioluminescence signal mapped using the goniometric scanner

The radioluminescence map produced with the goniometer-driven optical detector (scanner) is a 3-dimensional structure with X-axis representing the yaw, Y-axis representing the pitch, and Z-axis showing the signal strength (color-coded when depicted as an image). The total effect Inet is therefore quantified as the net volume under the signal surface which is the integral of count rates in each pixel of the source domain D:

Inet=D nnetx, ydxdy≈Dnnetx, y∆x∆y,#1

where nnet is the net count rate in an (x, y) pixel, and ∆x, ∆y are the scan steps in X- and Y-axis, respectively. Normally, a UV image has square pixels, with ∆x=∆y=R called the scan resolution. Equation (1) demonstrates that increasing the resolution of a scan improves the accuracy of the source activity calculation since the integral is estimated as a sum.

The algorithm for activity calculation from a radioluminescence map has the following steps:

- Marking the source domain and getting the total gross count rate. The raw UV image is visually analyzed, and the bounding shape is drawn around the source. The total gross count rate Σgross is calculated as the count rate sum across all the pixels within the source boundary, Psrc:

Σgross=i=1Psrcni.#2

- Marking the background domain and getting the average background count rate. The adjacent source regions are selected, from which the average background count rate nbg is calculated:

nbg=1Pbgj=1Pbgnj ,#3

where Pbg is the number of pixels in the background domain.

- The total noise in the source domain Σnoise is estimated from the average background domain count rate:

Σnoise=nbg Psrc ,#4

where Psrc is the number of pixels in the source domain.

- The total net count rate in the source domain Σnet is determined as

Σnet=Σgross-Σnoise .#5

- The total surface activity of the source As is defined with the following formula:

As=InetTs ,#6

where T is the combined transmission of the light propagation medium between the source and the scanner, and s is the scanner sensitivity at the same measurement conditions determined using the reference alpha source. With Inet approximation given by Eq. (1), the activity is estimated as

As≈ R2ΣnetTs ,#7

where R is the scan resolution in degrees. Since R and T are dimensionless coefficients in the formula, they can be condensed to a single factor k, shaping the final formula for the total surface activity calculation:

As≈kΣnets .#8

Much like the area under the curve, but in two dimensions.

Does not include the transmission factors of the scanner optics as they are already included in the sensitivity coefficient.