Supplemental:

Major bioanalytical LC-MS/MS methods used for in vivo and in vivo sample analysis for this work

A: LC-MS/UV method used for Hepatocyte and LM stability assays

HPLC: Shimadzu HPLC (including pumps, autosampler, and UV detector at 254 nm)

Analytical column: Waters Atlantis T3, 150 × 4.6 mm, 5-µm

Mobile Phase A: 0.1% (v/v) formic acid in water

Mobile phase B: Acetonitrile

Flow rate: 1.0 mL/min; Split: after UV, 20% to MS, and 80% to waste

Gradient: From 0 – 3 min, MPB 5 – 10%; from 3 – 8 min, MPB 10- 50%; from 8- 10 min, MPB 50 - 100%; then back to MPB 5% and re-equilibrating for 2 mins.

MS instrument and conditions:

MS: Finnigan TSQ7000

Software: Xcalibur v 1.2 (acquiring data), v 2.06 (processing data)

Ionisation: Electrospray (ESI), positive ion mode

Spray voltage: 4.5 kV

Heated capillary temperature: 350 ◦C

Sheath gas: N2, 50 psi

Auxiliary gas: N2, 10 units

Scan range: 120 to 600 amu

Scan rate: 1 s/scan

MS/MS conditions: collision gas – Argon, 2 mTorr; collision energy 20 – 25 eV; Scan range: m/z 10 to (M+5) amu (where M is the protonated molecular ion); scan rate 1 s/scan.

B: LC-MS/MS method used for Human Hepatocytes and Liver Microsomes, and Induction studies

Samples were analyzed by multiple reaction monitoring LC-MS/MS methods developed at the testing facility. Analysis was performed by electrospray ionization (ESI) source coupled with Sciex (or Waters) mass spectrometer. LC system used was a Shimadzu Nexera (or Waters Acquity UPLC). Sample analysis, integration and reporting were conducted according to corresponding laboratory SOPs.

C: LC-MS/MS method used for the transporter study

Agilent 1290 HPLC coupled with MS/MS (Agilent 6470A) was used for sample analysis in the transporter study. The LC-MS/MS conditions are summarized below:

D: LC-MS/MS method used for in vitro metabolism using HµRel® System

E: LC-MS/MS methods used for phenotyping study sample assays at WuXi (XBL) Lab

E-1: For SEP-363856:

UHPLC System:	Shimadzu CBM-20A system controller, LC-30AD pump, SIL-30AC autosampler

Data System: 	Analyst® 1.6.3

HPLC Column:		Waters Acquity UPLC BEH C18 2.1 x 50 mm

Mobile Phase A:	0.1% Formic Acid in Water

Mobile Phase B:	0.1% Formic Acid in Acetonitrile

Flow rate:	0.60 mL/min

Gradient:

| Time (min) | 0.0 | 0.5 | 2 | 2.5 | 2.51 | 3.5 |
| --- | --- | --- | --- | --- | --- | --- |
| MPB (%) | 5 | 5 | 95 | 95 | 5 | Stop |

Mass Spectrometer:	AB Sciex API 4000

Ionization Mode:	ESI Positive

Scan Mode:	Multiple Reaction Monitoring (MRM)

CAD:	6

CUR:	25

GS1 = GS2:	50

Ion Spray Voltage: 	5.0 kV

Source Temperature: 	500 °C

EP:	10 V

CXP:	10 V

Dwell time	100 ms

MS/MS parameters:

| Compound | Precursor ion (m/z) | Product ion (m/z) | DP (V) | CE (eV) |
| --- | --- | --- | --- | --- |
| SEP-363856 | 184.1 | 135.1 | 65 | 30 |
| SEP-363855-d3 | 187.2 | 135.1 | 65 | 30 |
| Labetalol | 329.2 | 162.1 | 50 | 50 |
| Tolbutamide | 271.1 | 155.3 | 50 | 25 |

E-2: For SEP-363854:

UPLC System:		Waters UPLC

Data System: 	Masslynx v.4.2

HPLC Column:		Waters XBridge Amide 100 x 2.1 mm, 3.5µm

Mobile Phase A:	0.1% Formic Acid in Water

Mobile Phase B:	0.1% Formic Acid in Acetonitrile

Flow rate:	0.70 mL/min

Gradient:

| Time (min) | 0.0 | 0.5 | 2 | 2.01 | 2.2 | 2.21 | 3.5 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| MPB (%) | 95 | 95 | 60 | 50 | 50 | 95 | 95 |

Mass Spectrometer:	Waters-TQS

Ionization Mode:	ESI Positive

Scan Mode:	MRM

Cone Voltage:	30 V

Desolvation gas:	800 L/hr

Cone:	150 L/hr

Ion Spray Voltage: 	5.0 kV

Temperature: 	500°C

EP:	10 V

Dwell time	100 ms

MS/MS parameters:

| Compound | Precursor ion (m/z) | Product ion (m/z) | CE (eV) |
| --- | --- | --- | --- |
| SEP-363854 | 170.1 | 135.1 | 20 |
| SEP-363055-d4 | 174.1 | 138.1 | 20 |

E-3: For SEP-383103:

HPLC System:               Shimadzu CBM-20A system controller, LC-30AD pump, SIL-30AC autosampler

Data System:                 Analyst® 1.6.3

HPLC Column:             Waters ACQUITY UPLC BEH Amide 50 x 2.1 mm

Mobile Phase A:            0.2% acetic acid in water, pH 4.0

Mobile Phase B:            Acetonitrile

Flow Rate:                     0.60 mL/min

Gradient:

| Time (min) | 0.0 | 0.5 | 2.0 | 2.3 | 2.31 | 3.5 |
| --- | --- | --- | --- | --- | --- | --- |
| MPB (%) | 60 | 60 | 40 | 40 | 60 | stop |

Mass Spectrometer:       AB Sciex API 6500

Ionization Mode:           ESI Negative

Scan Mode:                   MRM

CAD	6

CUR	25

GS1	50

GS2	50

Ion Spray Voltage:	-4.5 kV

Source Temperature:	500 °C

EP	-10 V

CXP	-15 V

Dwell time:	70 ms

MS/MS parameters:

| Compound | Precursor Ion (m/z) | Product Ion (m/z) | DP (V) | CE (eV) |
| --- | --- | --- | --- | --- |
| SEP-383103 | 183 | 137 | -35 | 11 |
| IS: SEP-376955-d4 | 187 | 140 | -35 | 11 |

F: LC-MS/MS method used for Metabolite SEP-383103 CYP induction and inhibition

LC-MS/MS method and conditions used for sample analysis in the SEP383103 CYP induction and inhibition study are summarized below:

G: LC-MS/MS method used for Metabolite SEP-383103 transporter study

LC-MS/MS method used for sample analysis in the SEP383103 transporter study are summarized below:

H: LC-MS/MS method used for mouse plasma and brain PK measurement

The concentrations of SEP-0363856 in mouse plasma and brain samples were determined using a liquid chromatography with tandem mass spectrometry (LC-MS/MS) based method. A protein precipitation method was used for sample preparation. For non-diluted plasma samples: An aliquot of 30 µL sample was mixed with 200 µL IS (Propranolol, 40 ng/mL in ACN). The mixture was vortexed for 5 min and centrifuged at 5800 rpm for 10 min. An aliquot of 1 µL supernatant was injected for LC-MS/MS analysis. For 10x diluted plasma samples: An aliquot of 6 µL plasma sample was mixed with 54 µL blank plasma to obtain the diluted samples, with a final sample dilution factor of 10. The extraction procedure for diluted samples was the same as those for non-diluted samples. For brain homogenate samples: The sample was homogenized with 3 volumes (v/w) of PBS. The dilution factor was 4. The samples preparation was the same as plasma samples. When calculating brain tissue concentration, a dilution factor of 4 was applied, and brain concentration is reported as “ng/g”. The instrument setup consisted of an AB Sciex API 5500 Mass Spectrometer equipped with a Waters UPLC Acquity. Separation of SEP-0363856 and the IS from potential interfering compounds in the study samples is achieved on an ACQUITY UPLC HSS T3 column by a reversed-phase gradient chromatographic method. SEP-0363856 and the IS were ionized under a positive ion spray mode and detected through the multiple-reaction monitoring (MRM) of a mass transition pair at m/z 184.2  135.1 for SEP-0363856 and 260.2 → 116.1 for IS, respectively. Using a plasma or brain homogenate sample volume of 30 μL, linearity was achieved in the SEP-0363856 concentration range of 1.00 to 3000 ng/mL for the IP/PO dosing group plasma samples, and 1.00-3000 ng/g for brain samples from the same PK groups.

I: LC-MS/MS method used for Rat plasma PK measurement

LC-MS/MS method used for rat plasma PK measurement was summarized below:

J: LC-MS/MS method used for Dog plasma PK measurement

LC-MS/MS method used for dog plasma PK measurement was summarized below:

K: LC-MS/MS method used for Monkey plasma PK measurement

LC-MS/MS method used for monkey plasma PK measurement was summarized below: