# Supplementary Materials

ReCIVA and CASPER breath collection, paired background sampling, TD-GC-MS analysis, and the feature extraction method are collectively known as the ‘OMNI’ method. More details into the OMNI settings are contained here.

| Breath sample collection settings |  |
| --- | --- |
| Parameter | Value |
| Breath collection software and version | Breath Sample Collect (BSC) (version 3.4.0) |
| Number of tubes per pump | 2 |
| Flow rate per tube (mL/min) | 250 |
| Sampled volume (mL) per tube | 1250 |
| Breath fraction sampled | Lower airways |
| Blank fraction sampled | Continuous |

Supplementary Table 1: ReCIVA Breath Sampler collection settings. The table shows the various settings used with the ReCIVA Breath Sampler device utilized as part of the OMNI method used during this study.

Supplementary Figure 1: A – an overview of the breath collection hardware used in the OMNI method in this study. The ReCIVA Breath Sampler was used with a mouthpiece and connected to the CASPER clean air supply. The subjects breathed normally into the device and inhaled air from the device originating from the CAPSER. B – A schematic of the ReCIVA Breath Sampler is shown. The ReCIVA has inbuilt CO2 and pressure sensors that can be used to specifically activate the pumps to draw specific breath factions (i.e. end-tidal breath). C – An example of the CO2 and pressure values across normal breathing. The sensors can trigger the pumps to draw breath into the adsorbent tubes to specifically collect VOCs from the end-tidal breath fraction.

| Internal Standards |  |  |  |
| --- | --- | --- | --- |
| Compound Name | InChI Key | Class | Subclass |
| Hexane-D14 | VLKZOEOYAKHREP-ZLKPZJALSA-N | Saturated hydrocarbons | Alkanes |
| 2-Butanone-1,1,1,3,3-D5 | ZWEHNKRNPOVVGH-PDWRLMEDSA-N | Organooxygen compounds | Carbonyl compounds |
| Heptane-D16 | IMNFDUFMRHMDMM-NEBSKJCTSA-N | Saturated hydrocarbons | Alkanes |
| 2-Pentanone-1,1,1,3,3-D5 | XNLICIUVMPYHGG-PVGOWFQYSA-N | Organooxygen compounds | Carbonyl compounds |
| Octane-D18 | TVMXDCGIABBOFY-VAZJTQEUSA-N | Saturated hydrocarbons | Alkanes |
| Decane-D22 | DIOQZVSQGTUSAI-DNYWURFXSA-N | Saturated hydrocarbons | Alkanes |
| Dodecane-D26 | SNRUBQQJIBEYMU-DWXHPOBZSA-N | Saturated hydrocarbons | Alkanes |
| Hexane-D14 | VLKZOEOYAKHREP-ZLKPZJALSA-N | Saturated hydrocarbons | Alkanes |

Supplementary Table 2: The identifiers (name and InChI key) and chemical class for each of the eight deuterated internal standard compounds liquid injected onto all samples, prior to analysis.

| TD-GC-MS method settings |  |
| --- | --- |
| TD |  |
| Tube desorption temperature (°C) | 240 |
| Tube desorption time (min) | 5 |
| Trap desorb flow (mL/min) | 50 |
| Trap low temperature (°C) | 5 |
| Trap heat rate (°C/s) | 40 |
| Trap high temperature (°C) | 270 |
| Trap desorption time (min) | 10 |
| GC |  |
| Column chemistry | TraceGOLDTM-624SilMS (6% cyanopropylphenyl-94% dimethylpolysiloxane), 0.25 mm (D) x 1.4 µm (Film thickness) x 30 m (L) |
| Oven temperature profile | Hold at 37°C (4 min), 5°C/min ramp to 150°C, 20°C/min ramp to 300°C and hold (10.9 min) |
| Carrier gas flow (mL/min) | 2 |
| Aux heater 1 & 2 temperature (°C) | 300 |
| MS |  |
| Scan type | Full |
| Scan range (m/z) | 30-450 |
| Ionization mode | Positive |
| Electron energy (eV) | 35 |
| Transfer line temperature (°C) | 300 |
| Ion source temperature (°C) | 280 |

Supplementary Table 3: TD-GC-MS method settings. The table shows the various settings used for Thermal Desorption (TD), Gas Chromatography (GC), and Mass Spectrometry (MS) in the OMNI method.

Supplementary Figure 2: A total ion chromatogram (TIC) from a breath sample. Retention time (min) is shown on the x-axis and peak area intensity on the y-axis.

| Peak detection |  |
| --- | --- |
| Mass tolerance (ppm) | 5 |
| Spectral S/N threshold | 3 |
| Peak S/N threshold | 3 |
| TIC threshold | 150,000 |
| Ion overlap window (%) | 98 |
| Group Compounds |  |
| RT tolerance (s) | 7 |
| Dot product threshold | 500 |
| Composition threshold (%) | 35 |

Supplementary Table 4: Feature extraction method parameters. The table shows the various settings used for peak detections and group compounds in the OMNI method.