ONLINE SUPPLEMENT

for

Detection of Apnea and Hypopnea Events Using a Wireless Abdomen-Worn Sensor with SpO₂ Integration

Thi Hang Dang1,2, Hyung-ki Min2, Nam-hwan Sung2, Seong-mun Kim, Heein Yoon1, Bien Franklin1,2*

(1) Department of Electrical Engineering, Ulsan National Institute of Science and Technology, Ulsan, South Korea.

(2) SB Solutions Inc., Ulsan National Institute of Science and Technology, Ulsan, South Korea.

(*) Correspondence:

Franklin Don Bien

Ulsan National Institute of Science and Technology (UNIST)

50, UNIST-gil, Ulsan 44919, South Korea

Tel. No. +82 10-5545-4800; E-mail: bien@unist.ac.kr

Evaluation of apnea and hypopnea detection performance using different feature groups

The evaluation aimed to compare the performance of apnea and hypopnea detection using different groups of features. We tested multiple feature combinations with MLP-Mixer models for apnea and hypopnea detection and assessed their effectiveness based on the area under the curve (AUCROC). Each model was trained using five-fold cross-validation with a batch size of 32, a window length of 500 samples, and two hidden layers of dimension 32. The Adam optimizer (learning rate: 0.0005) was used for 100,000 iterations. Tables A and B summarize the AUC results for various feature groups in the Soom and SoomOxy models. The Soom model achieved the best performance using cap, gap, pos, acc_length, and ratio_cap, while the SoomOxy model achieved the highest AUC when including these five features plus the three SpO₂-based features: normalized SpO₂, depth_oxy, and dur_oxy.

Table A Area under the curve comparison between Soom models for apnea and hypopnea detection using different feature groups

| Group | Features | Apnea | Hypopnea |
| --- | --- | --- | --- |
| 1 | cap, gap, pos, acc_length | 0.9161 | 0.7006 |
| 2 | cap, gap, pos, acc length, ratio_cap | 0.9213 | 0.7152 |

Table B Area under the curve comparison between SoomOxy models for apnea and hypopnea detection using different feature groups

| Group | Features | Apnea | Hypopnea |
| --- | --- | --- | --- |
| 1 | cap, gap, pos, acc_length, ratio_cap, normalized SpO2 | 0.9418 | 0.8461 |
| 2 | cap, gap, pos, acc_length, ratio_cap, normalized SpO2, depth_oxy | 0.9412 | 0.8671 |
| 3 | cap, gap, pos, acc_length, ratio_cap, normalized SpO2, depth_oxy, dur_oxy | 0.9447 | 0.8702 |