SUPPLEMENTARY MATERIALS
FIGURE 3 Receiver operating characteristic curve for pulse transit time–derived systolic blood pressure detecting nocturnal hypertension.
The curve plots sensitivity versus 1 – specificity for pulse transit time–derived systolic blood pressure in predicting a mean nighttime blood pressure ≥ 120/70 mm Hg (ambulatory blood pressure monitoring reference). The area under the curve is 0.62 (95% CI 0.48–0.77). Coordinates and Youden index values (optimum threshold approximately 104 mm Hg) are provided in the accompanying table.
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	Threshold (mmHg)
	Specificity
	Sensitivity
	Youden’s J Index

	-Inf
	0.00000000
	1.00
	0.00000000

	87.5
	0.02325581
	1.00
	0.02325581

	92.0
	0.06976744
	1.00
	0.06976744

	93.5
	0.11627907
	1.00
	0.11627907

	95.0
	0.13953488
	0.95
	0.08953488

	96.5
	0.16279070
	0.95
	0.11279070

	97.5
	0.23255814
	0.95
	0.18255814

	98.5
	0.25581395
	0.95
	0.20581395

	99.5
	0.25581395
	0.90
	0.15581395

	101.0
	0.25581395
	0.85
	0.10581395

	102.5
	0.32558140
	0.85
	0.17558140

	104.0
	0.39534884
	0.85
	0.24534884

	105.5
	0.39534884
	0.80
	0.19534884

	106.5
	0.39534884
	0.75
	0.14534884

	107.5
	0.44186047
	0.75
	0.19186047

	110.5
	0.53488372
	0.70
	0.23488372

	113.5
	0.53488372
	0.65
	0.18488372

	114.5
	0.58139535
	0.60
	0.18139535

	115.5
	0.58139535
	0.50
	0.08139535

	116.5
	0.72093023
	0.40
	0.12093023

	117.5
	0.74418605
	0.40
	0.14418605

	119.0
	0.76744186
	0.40
	0.16744186

	120.5
	0.79069767
	0.40
	0.19069767

	121.5
	0.81395349
	0.40
	0.21395349

	122.5
	0.83720930
	0.35
	0.18720930

	123.5
	0.83720930
	0.30
	0.13720930

	126.0
	0.83720930
	0.25
	0.08720930

	128.5
	0.83720930
	0.20
	0.03720930

	130.0
	0.83720930
	0.10
	-0.06279070

	132.5
	0.90697674
	0.10
	0.00697674

	135.0
	0.93023256
	0.10
	0.03023256

	136.5
	0.95348837
	0.10
	0.05348837

	138.5
	0.95348837
	0.05
	0.00348837

	141.5
	1.00000000
	0.05
	0.05000000

	Inf
	1.00000000
	0.00
	0.00000000





	FIGURE 4 Receiver operating characteristic curve for pulse transit time–derived diastolic blood pressure detecting nocturnal hypertension.
This curve evaluates pulse transit time–derived diastolic blood pressure against the same nocturnal-hypertension standard as in Figure A (a mean nighttime blood pressure ≥ 120/70 mm Hg, as determined by ambulatory blood pressure monitoring). The area under the curve is 0.58 (95% CI 0.42–0.74). The table lists all thresholds, sensitivities, specificities, and Youden index values (optimum threshold approximately 78.5 mm Hg).
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	Threshold (mmHg)
	Specificity
	Sensitivity
	Youden’s J Index

	-Inf
	0.00000000
	1.00
	0.00000000

	50.0
	0.02325581
	1.00
	0.02325581

	53.5
	0.04651163
	1.00
	0.04651163

	54.5
	0.06976744
	1.00
	0.06976744

	55.5
	0.06976744
	0.95
	0.01976744

	57.5
	0.09302326
	0.95
	0.04302326

	59.5
	0.11627907
	0.90
	0.01627907

	60.5
	0.16279070
	0.90
	0.06279070

	61.5
	0.16279070
	0.85
	0.01279070

	62.5
	0.20930233
	0.85
	0.05930233

	63.5
	0.23255814
	0.85
	0.08255814

	64.5
	0.23255814
	0.80
	0.03255814

	65.5
	0.27906977
	0.75
	0.02906977

	66.5
	0.30232558
	0.75
	0.05232558

	67.5
	0.32558140
	0.70
	0.02558140

	69.0
	0.39534884
	0.65
	0.04534884

	70.5
	0.48837209
	0.65
	0.13837209

	71.5
	0.48837209
	0.60
	0.08837209

	72.5
	0.58139535
	0.55
	0.13139535

	73.5
	0.58139535
	0.45
	0.03139535

	74.5
	0.65116279
	0.45
	0.10116279

	75.5
	0.67441860
	0.45
	0.12441860

	76.5
	0.69767442
	0.40
	0.09767442

	77.5
	0.74418605
	0.40
	0.14418605

	78.5
	0.81395349
	0.40
	0.21395349

	79.5
	0.83720930
	0.35
	0.18720930

	80.5
	0.88372093
	0.30
	0.18372093

	82.0
	0.88372093
	0.20
	0.08372093

	84.0
	0.88372093
	0.15
	0.03372093

	85.5
	0.93023256
	0.10
	0.03023256

	88.0
	0.95348837
	0.10
	0.05348837

	90.5
	0.97674419
	0.10
	0.07674419

	96.5
	1.00000000
	0.10
	0.10000000

	102.5
	1.00000000
	0.05
	0.05000000

	Inf
	1.00000000
	0.00
	0.00000000


FIGURE 5 Receiver operating characteristic curve for pulse transit time–derived systolic blood pressure detecting masked hypertension.
The curve assesses pulse transit time–derived systolic blood pressure performance in identifying masked hypertension, defined by abnormal daytime or nighttime ambulatory blood pressure monitoring in patients without a prior hypertension diagnosis. The area under the curve is 0.58 (95% CI 0.38–0.77). Threshold coordinates and Youden indices (optimum threshold approximately 104 mm Hg) accompany the plot.
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	Threshold (mmHg)
	Specificity
	Sensitivity
	Youden’s J Index

	-Inf
	0.00000000
	1.0000000
	0.00000000

	87.5
	0.03703704
	1.0000000
	0.03703704

	92.0
	0.11111111
	1.0000000
	0.11111111

	93.5
	0.18518519
	1.0000000
	0.18518519

	95.0
	0.22222222
	0.8888889
	0.11111111

	96.5
	0.25925926
	0.8888889
	0.14814815

	97.5
	0.33333333
	0.8888889
	0.22222222

	99.0
	0.37037037
	0.8888889
	0.25925926

	101.0
	0.37037037
	0.7777778
	0.14814815

	102.5
	0.44444444
	0.7777778
	0.22222222

	104.5
	0.48148148
	0.7777778
	0.25925926

	106.5
	0.48148148
	0.6666667
	0.14814815

	107.5
	0.55555556
	0.6666667
	0.22222223

	110.5
	0.66666667
	0.4444444
	0.11111111

	114.0
	0.66666667
	0.3333333
	0.00000000

	115.5
	0.66666667
	0.2222222
	-0.11111111

	118.5
	0.77777778
	0.1111111
	-0.11111111

	121.5
	0.81481481
	0.1111111
	-0.07407407

	123.0
	0.85185185
	0.1111111
	-0.03703704

	127.5
	0.85185185
	0.0000000
	-0.14814815

	132.5
	0.92592593
	0.0000000
	-0.07407407

	135.0
	0.96296296
	0.0000000
	-0.03703704

	Inf
	1.00000000
	0.0000000
	0.00000000





FIGURE 6 Receiver operating characteristic curve for pulse transit time–derived diastolic blood pressure detecting masked hypertension.
This curve presents pulse transit time–derived diastolic blood pressure accuracy for masked hypertension; the area under the curve is 0.51 (95% CI 0.29–0.73). A full table of thresholds, sensitivities, specificities, and Youden indices (optimum threshold approximately 64 mm Hg) accompany the plot.
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	Threshold (mmHg)
	Specificity
	Sensitivity
	Youden’s J Index

	-Inf
	0.00000000
	1.0000000
	0.00000000

	50.0
	0.03703704
	1.0000000
	0.03703704

	53.5
	0.07407407
	1.0000000
	0.07407407

	54.5
	0.11111111
	1.0000000
	0.11111111

	55.5
	0.11111111
	0.8888889
	0.00000000

	57.5
	0.14814815
	0.8888889
	0.03703704

	60.5
	0.18518519
	0.8888889
	0.07407407

	62.5
	0.22222222
	0.8888889
	0.11111111

	64.0
	0.25925926
	0.8888889
	0.14814815

	65.5
	0.25925926
	0.7777778
	0.03703704

	66.5
	0.29629630
	0.7777778
	0.07407407

	67.5
	0.29629630
	0.6666667
	-0.03703703

	69.0
	0.33333333
	0.6666667
	0.00000000

	70.5
	0.40740741
	0.6666667
	0.07407407

	71.5
	0.40740741
	0.5555556
	-0.03703703

	72.5
	0.51851852
	0.4444444
	-0.03703704

	73.5
	0.51851852
	0.3333333
	-0.14814815

	74.5
	0.59259259
	0.3333333
	-0.07407407

	75.5
	0.62962963
	0.3333333
	-0.03703704

	76.5
	0.66666667
	0.2222222
	-0.11111111

	77.5
	0.74074074
	0.2222222
	-0.03703704

	79.0
	0.85185185
	0.2222222
	0.07407407

	80.5
	0.88888889
	0.2222222
	0.11111111

	83.0
	0.88888889
	0.1111111
	0.00000000

	87.5
	0.96296296
	0.1111111
	0.07407407

	96.5
	1.00000000
	0.1111111
	0.11111111

	Inf
	1.00000000
	0.0000000
	0.00000000
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