Supplementary

Differential prevalence of PFAS, PCBs and pesticides of hunted game – Supporting Table 1

Alexandra Esther, Michelle Peter, Vera Ritz, Daniel Esther, Doreen Gabriel, Roman Trommler, Detlef Schenke, Klaus Polaczek, Kathrin Fisch, Christoph Müller

Content

Table S1: List of analysed pollutants

S2 Anticoagulant rodenticides

Table S2_a: Configuration of LC-MS/MS

Table S2_b: LC-MS/MS - MRM-conditions

Table S2_c1: FIRST ANALYSIS CAMPAIGN: Validation – Recovery

Table S2_c2: FIRST ANALYSIS CAMPAIGN: – Hunted liver samples

Table S2_d1: SECOND ANALYSIS CAMPAIGN: Validation – Recovery

Table S2 _d2: SECOND ANALYSIS CAMPAIGN: Validation – Hunted liver samples

S3 Herbicide, herbicide safener, insecticide, fungicide, nematicide, microbiocide, plant growth regulator

Table S3_a: Configuration of GC-MS/MS for the analysis of herbicide, herbicide safener, insecticide, fungicide, nematicide, microbiocide, plant growth regulator, PCB

S4 PFAS

Table S4_a: Configuration of LC-MS/MS for the analysis of PFAS

Table S4_b: LC-MS/MS - MRM-conditions for the analysis of PFAS

Table S1 List of analysed pollutants including name, pollutant type, retention time (RT), MRM transitions, collision energy (CE) and LOQ (limit of quantification). Quantifier transitions are marked in bold. OC: organochlorine, PCB: polychlorinated biphenyl. Detected pollutants are highlighted in gray; in italics determined by LC-MS/MS, more detailed information on the MS/MS settings are provided in Supplement S2; n.a. not applicable; n.d. not determined; * The rodenticides were analysed in two sequences at different retention times and LOQs (Supporting Table S2).

| Name | Pollutant type | RT [min] | MRM transitions | CE [eV] | LOQ [µg kg-1] |
| --- | --- | --- | --- | --- | --- |
| 2,4-D-ethyl ester | Herbicide | 7.49 | 247.9 185.0 185.0 114.9 175.0 111.0 | 10 25 10 | n.d. |
| 2-Phenylphenol | Microbiocide | 6.25 | 169.1 91.0 141.1 63.0 115.1 65.0 | 35 45 25 | n.d. |
| 8-Hydroxyquinoline | Fungicide, Microbiocide | 5.38 | 145.0 63.0 117.0 63.0 117.0 39.1 | 40 40 40 | n.d. |
| Acequinocyl | Insecticide | 16.77 | 342.9 188.8 341.9 187.9 187.9 131.0 | 20 15 20 | n.d. |
| Acetamiprid | Insecticide | 13.85 | 221.0 56.1 126.0 90.0 126.0 72.9 | 15 5 20 | n.d. |
| Acibenzolar-S-methyl | Fungicide | 9.30 | 182.0 167.1 182.0 153.1 182.0 135.0 | 10 10 15 | n.d. |
| Aclonifen | Herbicide | 12.39 | 264.1 194.2 194.1 167.1 194.1 139.1 | 15 20 25 | n.d. |
| Acrinathrin | Insecticide | 15.02 | 288.9 92.8 207.8 152.0 181.0 127.0 | 10 35 30 | n.d. |
| Aldrin | Insecticide (OC) | 9.94 | 262.9 192.9 262.9 190.9 254.9 220.0 | 35 35 20 | n.d. |
| Ametoctradin | Fungicide | 15.20 | 275.0 246.2 275.0 190.3 246.0 188.2 | 0 15 25 | 20 |
| Amisulbrom | Fungicide | 16.14 | 227.9 147.0 225.9 147.0 214.0 160.0 | 15 15 20 | 4 |
| Azoxystrobin | Fungicide | 18.30 | 344.1 182.9 344.1 171.9 344.1 155.8 | 25 40 40 | 4 |
| Beflubutamid | Herbicide | 10.67 | 192.9 145.1 192.9 95.0 176.1 79.1 | 15 35 25 | n.d. |
| Benalaxyl | Fungicide | 12.87 | 266.0 148.1 233.9 146.0 206.0 162.1 | 5 20 5 | n.d. |
| Bentazone | Herbicide | 10.11 | 225.0 181.9 198.0 92.0 182.0 90.0 | 5 30 15 | n.d. |
| Benthiavalicarb-isopropyl | Fungicide | 14.57 | 222.0 125.9 180.0 127.0 180.0 83.0 | 40 20 30 | n.d. |
| Bifenazate | Insecticide | 13.94 | 184.1 91.1 184.1 77.0 168.1 140.1 | 40 40 10 | 20 |
| Bifenox | Herbicide | 14.21 | 340.9 309.9 340.9 280.9 189.1 126.0 | 10 15 20 | n.d. |
| Bifenthrin | Insecticide | 13.83 | 181.0 115.1 166.0 139.1 166.0 115.1 | 45 35 35 | 4 |
| Boscalid | Fungicide | 16.50 | 140.0 112.0 140.0 76.0 111.9 76.0 | 10 25 15 | 1 |
| Brodifacoum | Rodenticide | *2.76 2.57 | 521.0 78.8 521.0 135.0 | -128 -46 | *2 1 |
| Bromadiolone | Rodenticide | *2.57 2.43 | 527.0 250.0 525.0 181.0 | -50 -46 | *2 1 |
| Bromoxynil | Herbicide | 7.41 | 276.8 88.0 274.7 167.9 274.7 88.0 | 30 15 30 | n.d. |
| Bromuconazole (2 isomers) | Fungicide | 13.85 14.29 | 295.0 172.9 293.0 172.9 173.0 109.0 | 10 10 30 | n.d. |
| Bupirimate | Fungicide | 11.80 | 315.8 207.9 208.0 68.9 193.0 109.0 | 10 30 15 | n.d. |
| Buprofezin | Insecticide | 11.74 | 304.9 175.0 249.1 193.0 171.1 115.0 | 10 10 10 | n.d. |
| Captan | Fungicide | 10.73 | 263.8 79.0 149.0 70.0 116.9 82.0 | 15 15 30 | n.d. |
| Carbetamide | Herbicide | 9.95 | 120.1 92.0 120.1 77.0 119.1 64.1 | 10 15 25 | 20 |
| Carboxin | Fungicide | 11.75 | 234.9 143.0 234.9 87.0 131.9 77.0 | 10 20 20 | n.d. |
| Carfentrazone-ethyl | Herbicide | 12.81 | 339.9 311.9 329.9 309.9 311.9 150.8 | 10 10 20 | n.d. |
| cis-Chlordane | Insecticide (OC) | 11.20 | 374.8 265.8 372.8 265.8 271.7 236.9 | 15 15 15 | 1 |
| trans-Chlordane | Insecticide (OC) | 10.94 | 374.8 265.8 372.8 265.8 271.7 236.9 | 15 15 15 | 1 |
| Chloridazon (Pyrazon) | Herbicide | 13.04 | 221.0 220.2 220.0 193.1 220.0 166.0 | 5 20 25 | n.d. |
| Chlorophacinone | Rodenticide | *2.34 2.12 | 373.0 201.0 | 32 | *10 2 |
| Chlorothalonil | Fungicide | 8.54 | 265.9 230.9 265.9 133.0 265.9 109.0 | 20 45 45 | n.d. |
| Chlorotoluron | Herbicide | 9.70 | 212.1 166.0 212.1 72.0 167.0 132.1 | 10 15 15 | n.d. |
| Chlorpropham | Herbicide, Plant growth regulator | 7.11 | 213.0 171.1 171.0 127.1 153.0 90.0 | 5 5 25 | n.d. |
| Chlorpyrifos | Insecticide | 9.86 | 313.8 257.8 196.9 107.0 196.9 98.0 | 15 40 30 | 1 |
| Chlorpyrifos-methyl | Insecticide | 9.14 | 287.9 92.9 285.9 93.0 124.9 47.0 | 20 25 15 | 1 |
| Clodinafop-propargyl | Plant growth regulator | 12.97 | 348.9 265.9 348.9 237.8 238.0 130.0 | 10 15 15 | n.d. |
| Clomazone | Herbicide | 7.98 | 205.1 107.1 127.0 101.0 125.0 89.0 | 20 20 15 | n.d. |
| Cloquintocet-mexyl | Herbicide safener | 14.00 | 220.0 191.9 163.0 128.0 163.0 101.0 | 10 15 30 | n.d. |
| Coumatetralyl | Rodenticide | *2.18 1.92 | 291.0 141.0 291.0 247.0 | -38 -32 | *1 0.2 |
| Cyflufenamid | Fungicide | 11.88 | 188.1 88.0 118.1 90.0 118.1 89.0 | 35 10 25 | 4 |
| Cyfluthrin (3 isomers) | Insecticide | 16.17 16.25 16.37 | 206.0 176.9 206.0 150.0 162.9 127.0 | 25 40 5 | n.d. |
| Cyhalofop-butyl | Herbicide | 14.68 | 357.1 229.1 256.2 120.1 229.2 109.1 | 15 10 15 | n.d. |
| Cyhalothrin (gamma and lambda isomer) | Insecticide | 14.79 14.60 | 208.0 181.0 208.0 152.0 197.0 161.1 | 5 25 5 | 4 |
| Cypermethrin (3 isomers) | Insecticide | 16.39 16.48 16.57 | 165.0 127.1 165.0 91.1 162.9 127.0 | 0 10 0 | 4 |
| Cyproconazole | Fungicide | 11.99 | 222.0 124.9 138.9 111.0 138.9 75.0 | 25 15 35 | n.d. |
| Cyprodinil | Fungicide | 10.39 | 225.2 224.3 224.2 131.1 210.0 93.0 | 10 15 20 | n.d. |
| Cyromazine | Insecticide | 7.97 | 165.9 109.0 151.0 82.0 109.0 68.0 | 20 30 20 | n.d. |
| Dazomet | Fungicide | 7.76 | 161.9 89.0 89.0 46.0 88.9 74.0 | 25 5 15 | n.d. |
| o,p‘-DDD | Insecticide (OC) (Breakdown) | 11.78 | 235.0 200.1 235.0 139.1 199.1 164.1 | 10 45 20 | n.d. |
| p,p‘-DDD | Insecticide (OC) (Breakdown) | 12.36 | 237.0 200.1 199.1 164.1 165.1 139.0 | 15 20 35 | n.d. |
| o,p‘-DDE | Insecticide (OC) (Breakdown) | 10.98 | 317.8 248.0 248.0 176.2 246.0 176.2 | 15 30 30 | n.d. |
| p,p‘-DDE | Insecticide (OC) (Breakdown) | 11.52 | 317.8 246.0 315.8 246.0 246.1 176.2 | 15 15 30 | 1 |
| o,p‘-DDT | Insecticide (OC) | 12.27 | 237.0 199.1 235.0 199.1 199.0 163.1 | 15 15 35 | n.d. |
| p,p‘-DDT | Insecticide (OC) | 12.94 | 237.0 165.2 235.0 199.2 235.0 165.2 | 20 15 20 | 2 |
| Deltamethrin | Insecticide | 18.02 | 252.9 174.0 252.9 93.1 251.0 172.0 | 0 15 0 | 4 |
| Desmedipham | Herbicide | 7.59 | 181.0 122.0 181.0 109.0 135.0 52.0 | 10 10 25 | n.d. |
| Diazinon | Insecticide | 8.29 | 276.0 137.1 199.1 135.1 179.1 137.1 | 25 10 20 | n.d. |
| Dicamba-methyl ester | Herbicide | 6.26 | 234.0 173.0 205.0 149.0 175.0 111.0 | 20 15 20 | n.d. |
| Diclofop-methyl | Herbicide | 13.26 | 339.9 252.9 280.8 119.9 253.0 162.1 | 10 10 15 | n.d. |
| Dieldrin | Insecticide (OC) | 11.62 | 277.0 241.0 262.9 193.0 262.9 191.0 | 5 35 35 | 2 |
| Diethofencarb | Fungicide | 9.76 | 225.0 96.0 207.0 179.1 207.0 151.0 | 30 5 15 | n.d. |
| Difenacoum | Rodenticide | *2.64 2.50 | 443.0 135.0 443.0 93.0 | -46 -80 | *1 0.2 |
| Difenoconazole (2 isomers) | Fungicide | 17.72 17.78 | 324.8 266.8 322.8 264.8 264.9 202.0 | 15 15 20 | 4 |
| Difethialone | Rodenticide | *2.79 2.64 | 539.0 80.8 539.0 151.0 | -92 -46 | *1 1 |
| Diflubenzuron | Insecticide | 5.01 | 141.0 113.0 141.0 63.0 113.0 63.0 | 40 40 40 | 4 |
| Diflufenican | Herbicide | 13.29 | 393.9 265.9 266.0 246.1 218.0 140.1 | 10 15 20 | n.d. |
| Dimethachlor | Herbicide | 8.99 | 209.9 134.1 196.9 148.2 134.1 79.1 | 10 10 20 | n.d. |
| Dimethenamide-P | Herbicide | 9.02 | 229.9 154.0 229.9 111.0 202.9 154.0 | 10 25 10 | n.d. |
| Dimethoate | Insecticide | 7.79 | 228.7 87.0 157.0 93.0 157.0 63.0 | 5 10 25 | n.d. |
| Dimethomorph (2 isomers) | Fungicide | 18.35 18.66 | 302.9 164.9 300.9 165.0 300.9 138.8 | 10 10 15 | 4 |
| Dimoxystrobin | Fungicide | 13.85 | 237.0 116.0 205.0 116.0 174.0 115.0 | 15 10 30 | n.d. |
| Diuron | Herbicide | 10.66 | 231.7 71.8 186.9 124.0 158.9 123.9 | 15 20 10 | n.d. |
| Dodemorph (2 isomers) | Fungicide | 10.24 10.55 | 281.0 154.0 238.1 55.1 154.0 112.1 | 10 20 10 | n.d. |
| Epoxiconazole | Fungicide | 13.52 | 192.0 138.1 192.0 111.0 138.0 75.0 | 10 25 25 | 1 |
| Ethofenprox | Insecticide | 16.78 | 183.0 168.0 163.0 135.1 163.0 107.1 | 10 10 20 | n.d. |
| Ethofumesate | Herbicide | 9.61 | 285.9 207.1 178.9 137.1 178.9 105.1 | 5 0 15 | n.d. |
| Ethoprophos | Insecticide | 7.02 | 199.9 97.0 157.9 97.0 157.9 81.0 | 20 15 15 | n.d. |
| Etoxazole | Insecticide | 14.07 | 329.9 315.0 299.9 284.9 299.9 269.9 | 20 10 20 | n.d. |
| Etridiazole | Fungicide | 5.85 | 211.1 183.0 211.1 140.0 185.0 142.0 | 10 25 15 | n.d. |
| Famoxadone | Fungicide | 18.44 | 329.9 329.0 329.9 223.9 223.9 196.2 | 10 10 10 | n.d. |
| Fenamiphos | Insecticide | 11.31 | 302.9 287.9 302.9 153.9 287.9 259.7 | 10 15 5 | n.d. |
| Fenazaquin | Insecticide | 14.19 | 160.0 145.2 160.0 117.1 146.0 118.1 | 5 20 10 | 2 |
| Fenbuconazole | Fungicide | 16.21 | 197.9 129.0 197.9 102.0 125.0 89.0 | 5 30 20 | n.d. |
| Fenhexamid | Fungicide | 12.97 | 301.0 97.0 179.0 115.0 177.1 113.0 | 15 15 15 | 2 |
| Fenoxaprop-P-ethyl | Herbicide | 15.33 | 360.8 287.8 287.8 118.8 287.8 90.9 | 10 10 20 | n.d. |
| Fenoxycarb | Insecticide | 13.86 | 256.1 187.2 186.2 109.0 186.2 77.1 | 10 15 20 | n.d. |
| Fenpropidin | Fungicide | 9.45 | 273.0 98.0 145.0 117.0 145.0 91.0 | 5 10 25 | 4 |
| Fenpropimorph | Fungicide | 9.81 | 128.1 110.1 128.1 86.1 128.1 70.1 | 5 10 10 | 4 |
| Fenpyroximate | Insecticide | 7.85 | 212.0 185.0 212.0 76.9 198.1 114.0 | 40 40 35 | n.d. |
| Fenvalerate (2 isomers) | Insecticide | 17.31 17.50 | 419.1 166.8 167.0 125.1 167.0 89.0 | 10 10 40 | 2 |
| Fipronil | Insecticide | 10.64 | 366.8 212.8 350.8 254.8 254.9 228.0 | 25 15 15 | n.d. |
| Fipronil sulfide | Insecticide (Breakdown) | 10.50 | 420.0 350.9 351.0 254.9 254.9 156.9 | 10 20 35 | n.d. |
| Fipronil sulfone | Insecticide (Breakdown) | 11.71 | 384.8 256.8 382.8 254.9 254.9 227.9 | 20 20 15 | 2 |
| Flocoumafen | Rodenticide | *2.70 2.56 | 541.0 382.0 541.0 161.0 | -36 -48 | *1 0.4 |
| Fluazifop-p-butyl | Herbicide | 11.97 | 382.9 282.0 281.9 238.0 254.0 146.1 | 10 15 15 | 1 |
| Fludioxonil | Fungicide | 11.51 | 248.0 182.1 248.0 154.1 248.0 127.1 | 10 20 30 | 1 |
| Flufenacet | Herbicide | 9.96 | 211.0 123.0 211.0 96.0 183.0 69.0 | 5 15 20 | n.d. |
| Flumetralin | Herbicide | 11.19 | 403.9 156.8 359.9 313.9 157.0 109.0 | 15 15 25 | n.d. |
| Flumioxazin | Herbicide | 17.43 | 354.0 325.9 354.0 175.8 287.0 258.7 | 5 15 15 | n.d. |
| Fluometuron | Herbicide | 6.98 | 232.0 72.0 213.0 167.9 187.0 109.0 | 15 10 20 | n.d. |
| Fluopyram | Fungicide | 10.57 | 395.9 223.1 222.9 196.0 222.9 187.1 | 5 10 10 | 1 |
| Fluorochloridone | Herbicide | 10.11 | 311.0 174.1 311.0 102.9 187.1 109.1 | 15 15 20 | n.d. |
| Flupyradifurone | Insecticide | 14.87 | 288.0 126.1 128.0 90.0 126.0 73.0 | 15 10 25 | 20 |
| Fluquinconazole | Fungicide | 15.85 | 342.0 107.8 340.0 298.0 340.0 107.8 | 40 15 40 | n.d. |
| Fluroxypyr-meptyl | Herbicide | 13.29 | 237.0 209.0 237.0 181.0 208.9 178.9 | 5 15 20 | n.d. |
| Flurtamone | Herbicide | 14.43 | 332.7 120.0 157.0 137.1 157.0 107.0 | 15 15 25 | n.d. |
| Flutolanil | Herbicide | 11.38 | 322.9 281.0 280.9 173.0 173.0 95.0 | 5 10 30 | n.d. |
| Flutriafol | Fungicide | 11.30 | 219.1 123.1 219.1 95.0 164.1 109.1 | 15 35 20 | n.d. |
| tau-Fluvalinate (2 isomers) | Insecticide | 17.48 17.52 | 252.0 200.0 250.0 200.1 250.0 198.1 | 15 15 40 | 4 |
| Fluxapyroxad | Fungicide | 14.57 | 321.1 152.9 222.0 152.9 222.0 125.9 | 35 15 40 | n.d. |
| Fosthiazate (2 isomers) | Nematicide | 10.27 10.31 | 199.0 102.0 195.0 60.0 165.9 106.0 | 5 20 10 | n.d. |
| Fuberidazole | Fungicide | 9.16 | 184.0 155.1 156.0 103.1 155.0 129.1 | 30 20 10 | n.d. |
| Haloxyfop-P-methyl | Herbicide | 10.93 | 375.1 316.0 375.1 91.1 288.0 180.0 | 10 35 25 | n.d. |
| alpha-HCH | Insecticide (OC) | 7.64 | 218.9 183.0 216.9 181.0 180.9 145.0 | 5 5 15 | n.d. |
| beta-HCH | Insecticide (OC) | 7.99 | 218.9 183.1 216.9 181.1 181.0 145.0 | 5 5 15 | 1 |
| gamma-HCH (Lindane) | Insecticide (OC) | 8.08 | 218.9 183.1 216.9 181.0 181.0 145.0 | 5 5 15 | 1 |
| delta-HCH | Insecticide (OC) | 8.51 | 217.0 181.1 183.1 147.1 181.1 145.1 | 5 15 15 | n.d. |
| epsilon-HCH | Insecticide (OC) | 8.69 | 254.0 180.9 218.9 182.9 182.9 109.0 | 10 5 30 | n.d. |
| Heptachlor | Insecticide (OC) | 9.34 | 273.7 238.9 273.7 236.9 271.7 236.9 | 15 15 15 | 1 |
| Heptachlor endo-epoxide | Insecticide (OC) (Breakdown) | 10.67 | 216.9 182.0 216.9 109.0 183.0 119.0 | 20 45 30 | n.d. |
| Heptachlor exo-epoxide | Insecticide (OC) (Breakdown) | 10.61 | 354.8 264.9 352.8 262.9 262.9 193.0 | 15 15 35 | n.d. |
| Hexachlorobenzene | Fungicide | 7.70 | 283.8 213.9 281.8 211.9 248.9 179.0 | 30 30 30 | n.d. |
| Imazalil | Fungicide | 11.48 | 216.8 175.0 174.9 147.0 172.9 109.0 | 5 15 30 | n.d. |
| Imidacloprid | Insecticide | 11.31 | 211.0 113.0 126.0 89.9 126.0 73.0 | 15 5 25 | 20 |
| Indoxacarb | Insecticide | 18.02 | 264.0 175.8 202.9 134.0 202.9 106.0 | 15 20 15 | n.d. |
| Ipconazole | Fungicide | 15.00 | 249.0 125.0 167.0 125.0 125.0 89.0 | 15 5 20 | n.d. |
| Iprovalicarb (2 isomers) | Fungicide | 11.61 11.78 | 158.0 98.0 134.1 93.0 116.0 98.1 | 10 15 5 | n.d. |
| Isopyrazam | Fungicide | 15.30 | 359.0 159.0 302.1 262.1 159.0 139.0 | 40 15 10 | n.d. |
| Isoxaben | Herbicide | 15.17 | 165.0 150.0 165.0 107.0 149.9 121.9 | 15 25 5 | n.d. |
| Kresoxim-methyl | Fungicide | 11.81 | 206.0 131.1 206.0 116.0 116.0 89.0 | 10 5 15 | n.d. |
| Lenacil | Herbicide | 12.95 | 233.9 153.1 153.1 110.1 153.1 82.1 | 5 20 20 | 2 |
| Lufenuron | Insecticide | 5.58 | 251.6 157.8 202.9 75.9 173.9 109.9 | 15 40 30 | n.d. |
| Malathion | Insecticide | 9.73 | 172.9 117.0 172.9 99.0 157.8 125.0 | 15 10 5 | n.d. |
| MCPA-methyl ester | Herbicide | 6.51 | 214.1 155.1 214.1 141.1 155.1 125.1 | 10 10 10 | n.d. |
| MCPB-methyl ester | Herbicide | 8.17 | 211.1 155.0 142.1 107.1 142.1 77.1 | 10 10 30 | n.d. |
| Mefenpyr-diethyl | Herbicide safener | 13.59 | 299.0 252.9 253.0 190.0 253.0 189.0 | 10 20 30 | n.d. |
| Mepanipyrim | Fungicide | 11.16 | 222.2 158.1 221.2 220.2 207.1 179.1 | 25 15 25 | n.d. |
| Metalaxyl | Fungicide | 9.33 | 234.0 146.1 220.0 160.1 206.1 162.1 | 20 10 5 | n.d. |
| Metamitron | Herbicide | 11.83 | 202.1 186.1 202.1 104.1 173.1 132.1 | 5 15 10 | n.d. |
| Metazachlor | Herbicide | 10.45 | 209.0 133.2 209.0 132.2 209.0 117.1 | 10 15 35 | 1 |
| Metconazole | Fungicide | 14.22 | 153.1 125.0 153.1 70.0 125.0 89.0 | 10 5 20 | n.d. |
| Methiocarb | Insecticide | 9.58 | 169.0 154.1 168.0 109.1 153.0 91.1 | 10 15 20 | n.d. |
| Metobromuron | Herbicide | 8.79 | 258.0 61.0 196.9 89.9 169.9 142.9 | 10 25 20 | n.d. |
| (S)-Metolachlor | Herbicide | 9.89 | 238.0 162.2 238.0 133.2 162.1 133.2 | 10 30 15 | n.d. |
| Metrafenone | Fungicide | 15.24 | 394.8 364.8 376.9 346.8 226.9 169.0 | 15 20 10 | 4 |
| Metribuzin | Herbicide | 9.00 | 198.0 82.0 198.0 55.0 182.0 114.9 | 15 30 10 | n.d. |
| Myclobutanil | Fungicide | 11.68 | 179.0 125.1 179.0 90.0 150.0 123.0 | 10 30 15 | 2 |
| Napropamide | Herbicide | 11.40 | 271.0 100.1 271.0 72.1 128.0 100.1 | 15 15 10 | n.d. |
| Oryzalin | Herbicide | 15.51 | 316.8 274.9 275.0 217.0 258.0 193.9 | 5 5 5 | n.d. |
| Oxadiazon | Herbicide | 11.63 | 301.8 175.0 257.8 112.0 174.9 112.0 | 15 30 15 | n.d. |
| Oxamyl | Insecticide, Nematicide | 6.30 | 162.0 114.9 145.0 71.9 145.0 60.9 | 10 20 10 | n.d. |
| Oxychlordane | Insecticide (OC) (Breakdown) | 10.53 | 386.7 262.7 236.9 142.9 184.9 121.0 | 15 25 15 | 2 |
| Oxyfluorfen | Herbicide | 11.71 | 299.9 222.8 252.0 196.0 252.0 146.0 | 15 20 30 | n.d. |
| Paclobutrazol | Plant growth regulator | 11.09 | 236.0 167.1 167.1 132.1 125.1 89.0 | 10 10 20 | n.d. |
| Parathion | Insecticide | 9.97 | 291.0 137.1 291.0 109.0 139.0 81.0 | 5 15 15 | n.d. |
| Parathion-methyl | Insecticide, Nematicide | 9.14 | 262.9 109.0 262.9 79.0 109.0 79.0 | 10 30 5 | n.d. |
| PCB 28 | PCB | 9.04 | 258.0 186.0 256.0 186.0 186.0 151.0 | 25 25 25 | 1 |
| PCB 52 | PCB | 9.61 | 291.9 221.9 289.9 219.9 255.0 220.0 | 25 25 10 | 1 |
| PCB 101 | PCB | 11.12 | 325.9 255.9 325.9 253.9 253.9 184.0 | 35 30 30 | 1 |
| PCB 138 | PCB | 13.12 | 361.9 289.9 359.9 289.9 287.9 217.9 | 30 30 40 | 1 |
| PCB 153 | PCB | 12.62 | 361.9 289.9 359.9 289.9 287.9 217.9 | 25 25 40 | 1 |
| PCB 180 | PCB | 14.30 | 395.8 325.8 393.8 358.8 393.8 323.8 | 30 15 30 | 1 |
| Penconazole | Fungicide | 10.54 | 250.0 194.1 250.0 157.1 159.0 123.0 | 15 25 20 | n.d. |
| Pendimethalin | Herbicide | 10.52 | 251.8 162.2 251.8 146.1 161.9 147.0 | 10 20 10 | n.d. |
| Pentachloro-nitrobenzene | Fungicide, Nematicide | 8.20 | 294.8 236.8 248.8 213.8 141.9 106.9 | 15 15 30 | 1 |
| Perfluorooctanoic acid (PFOA) | PFAS | 3.8 | 412.8 368.9 412.8 219.0 412.8 168.9 | -16 -20 -22 | 8 |
| Perfluoro-octaesulfonic acid (PFOS) | PFAS | 2.3 | 498.8 129.9 498.8 98.8 498.8 80.0 | -58 -58 -82 | 8 |
| Perfluro-hexanesulfonic acid (PFHxS) | PFAS | 2.7 | 398.7 118.9 398.7 98.9 398.7 79.9 | -46 -48 -78 | 8 |
| Permethrin (cis and trans isomer) | Insecticide | 15.51 15.63 | 165.0 127.0 162.9 127.0 162.9 91.0 | 0 0 10 | 4 |
| Phenmedipham | Herbicide | 7.08 | 167.0 135.0 167.0 122.0 122.0 94.0 | 15 15 15 | n.d. |
| Phosmet | Insecticide | 13.90 | 161.0 134.0 161.0 78.0 160.0 133.1 | 10 20 10 | n.d. |
| Phosmet-oxon | Insecticide (Breakdown) | 13.00 | 301.0 191.8 172.9 104.0 160.0 133.0 | 10 15 15 | n.d. |
| Picloram-methyl ester | Herbicide | 9.55 | 198.0 163.1 198.0 161.0 196.0 181.0 | 15 15 15 | n.d. |
| Picolinafen | Herbicide | 13.87 | 376.0 239.1 376.0 238.1 238.1 145.1 | 10 20 25 | 2 |
| Picoxystrobin | Fungicide | 11.29 | 334.9 172.9 302.8 156.9 145.0 102.1 | 10 15 25 | n.d. |
| Pirimicarb | Insecticide | 8.73 | 238.0 166.2 166.0 71.1 152.0 123.0 | 10 25 10 | 1 |
| Pirimiphos-methyl | Insecticide | 9.58 | 290.0 125.0 232.9 151.0 232.9 125.0 | 20 5 5 | n.d. |
| Prochloraz | Fungicide | 15.91 | 310.0 69.8 266.0 69.9 180.0 68.9 | 15 10 15 | n.d. |
| Propamocarb | Fungicide | 5.39 | 188.0 58.0 143.0 99.1 129.1 84.1 | 10 10 5 | n.d. |
| Propaquizafop | Herbicide | 19.92 | 298.8 254.8 162.9 135.8 162.9 99.9 | 25 10 20 | n.d. |
| Propiconazole (2 isomers) | Fungicide | 12.89 13.00 | 258.8 172.9 172.9 109.0 172.9 74.0 | 15 30 45 | 2 |
| Prosulfocarb | Herbicide | 9.37 | 251.0 218.3 251.0 128.2 251.0 100.1 | 10 5 5 | n.d. |
| Prothioconazole-desthio | Fungicide | 11.91 | 186.0 89.0 186.0 70.0 125.0 99.0 | 10 10 20 | n.d. |
| Pymetrozine | Insecticide | 11.51 | 132.0 105.0 132.0 78.0 113.0 98.0 | 10 20 5 | n.d. |
| Pyraclostrobin | Fungicide | 17.46 | 324.8 131.7 164.0 132.1 110.8 75.0 | 15 10 15 | n.d. |
| Pyraflufen-ethyl | Herbicide | 13.03 | 412.0 349.0 349.0 307.0 338.9 288.9 | 10 15 15 | n.d. |
| Pyridaben | Insecticide | 15.77 | 309.0 147.1 147.2 132.2 147.2 117.1 | 15 10 20 | n.d. |
| Pyridalyl | Insecticide | 16.73 | 204.0 148.0 164.0 146.0 146.0 126.0 | 25 15 10 | n.d. |
| Pyridate | Herbicide | 17.26 | 205.2 141.1 205.2 114.0 205.2 102.0 | 25 35 30 | n.d. |
| Pyrimethanil | Fungicide | 8.24 | 198.0 183.1 198.0 158.1 198.0 118.1 | 15 20 35 | 1 |
| Pyriproxyfen | Insecticide | 14.61 | 321.0 222.0 321.0 153.0 136.1 96.0 | 10 25 15 | n.d. |
| Quinoclamine | Herbicide | 9.76 | 209.0 172.1 207.0 172.1 172.0 89.0 | 10 20 20 | n.d. |
| Quinoxyfen | Fungicide | 12.85 | 306.8 237.0 271.9 237.1 237.0 208.0 | 20 10 30 | 1 |
| Spirodiclofen | Insecticide | 15.56 | 312.1 259.0 312.1 108.9 157.0 73.0 | 10 15 25 | 2 |
| Spiromesifen | Insecticide | 13.71 | 272.0 209.2 253.8 185.1 231.0 157.1 | 10 15 15 | n.d. |
| Spiroxamine (2 Isomers) | Fungicide | 9.08 9.53 | 198.0 126.1 126.0 84.0 100.0 58.1 | 5 5 10 | n.d. |
| Tebuconazole | Fungicide | 13.22 | 250.0 125.0 125.0 99.0 125.0 89.0 | 20 20 15 | 2 |
| Tebufenpyrad | Insecticide | 14.09 | 332.9 171.0 318.0 131.0 275.9 171.1 | 15 15 10 | n.d. |
| Tefluthrin | Insecticide | 8.41 | 199.0 161.1 197.0 161.1 177.1 127.1 | 5 5 15 | n.d. |
| Terbuthylazine | Herbicide, Microbiocide | 8.12 | 228.9 138.0 214.0 71.0 172.9 138.1 | 15 20 5 | 1 |
| Terbuthylazine-desethyl | Herbicide, Microbiocide | 7.36 | 186.2 104.0 186.2 83.1 145.1 110.1 | 15 20 10 | 1 |
| Tetraconazole | Fungicide | 9.99 | 336.0 217.9 336.0 203.8 170.9 136.0 | 20 30 10 | 2 |
| Thiabendazole | Fungicide | 10.73 | 201.9 175.0 201.0 130.0 173.9 65.0 | 15 30 30 | n.d. |
| Thiacloprid | Insecticide | 17.50 | 126.0 99.1 126.0 90.1 126.0 73.0 | 10 5 20 | n.d. |
| Tolclofos-methyl | Fungicide | 9.14 | 267.0 252.0 267.0 93.0 267.0 63.0 | 15 30 45 | 2 |
| Tralkoxydim | Herbicide | 14.75 | 282.1 226.0 268.2 143.0 226.0 143.0 | 10 40 25 | n.d. |
| Triadimenol | Fungicide | 10.73 | 129.9 102.0 129.9 65.0 112.0 58.0 | 15 25 10 | n.d. |
| Triallate | Herbicide | 8.57 | 270.0 228.1 268.0 226.1 268.0 184.1 | 10 10 20 | n.d. |
| Triclopyr-methyl ester | Herbicide | 7.51 | 209.9 145.9 209.9 109.9 145.9 110.0 | 20 35 15 | n.d. |
| Trifloxystrobin | Fungicide | 12.94 | 186.0 145.1 172.0 145.1 172.0 95.0 | 15 15 30 | n.d. |
| Triflumizole | Fungicide | 10.81 | 345.0 302.0 239.1 66.9 132.0 90.0 | 10 40 35 | n.d. |
| Trinexapac-ethyl | Herbicide | 9.50 | 224.0 151.0 224.0 95.0 207.0 68.9 | 5 25 25 | n.d. |
| Triticonazole | Fungicide | 14.51 | 237.0 182.0 237.0 167.1 234.8 182.1 | 10 25 10 | n.d. |
| Warfarin | Rodenticide | *2.27 2.07 | 307.0 161.0 307.0 250.0 | -26 -32 | *0.5 0.2 |
| Zoxamide | Fungicide | 13.47 | 259.9 189.0 257.9 187.1 189.0 161.1 | 10 10 15 | 4 |

S2 Anticoagulant rodenticides

Table S2_a: Configuration of LC-MS/MS

| LIQUID CHROMATOGRAPHY | Agilent 1290 II Infinity |  |  |
| --- | --- | --- | --- |
| Autosampler temperature | 10 °C |  |  |
| Injection volume | 5 µL |  |  |
| Analytical column | Agilent Zorbax Eclipse C18 (1.8 µm, 50 mm, 2.1 mm i.d.) |  |  |
| Column temperature | 40 °C |  |  |
| Mobile phase A | Water +1 mmol NH4F |  |  |
| Mobile phase B | Methanol /Acetonitrile (65/35) |  |  |
| Gradient program | Time (min) | A (%) | B (%) |
|  | 0.0 | 98 | 2 |
|  | 3.0 | 2 | 98 |
|  | 5.0 | 2 | 98 |
|  | 5.1 | 98 | 2 |
|  | 6.0 | 98 | 2 |
| Flow rate | 500 µL/min |  |  |
| MASS SPECTROMETER | QTRAP 6500+ (SCIEX) |  |  |
| Mode | negative ESI |  |  |
| Ion spray potential | -4500 V |  |  |
| Source temperature | 550 °C |  |  |
| Scan type | Multiple Reaction Monitoring (MRM) / Enhanced Product Ion (EPI) |  |  |
| Dwell time | 20 ms |  |  |
| Software | Analyst 1.7.1 |  |  |
| Quantification | relative peak area |  |  |

Table S2_b: LC-MS/MS - MRM-conditions of rodenticides, surrogates (Surr) and internal standards (IS); precursor (Q1), product ions (Q3) and mass range for EPI in m/z and declustering potential (DP), entrance potential (EP), collision energy (CE) and cell exit potential (CXP) in all in [V]; Quantifier transitions are marked in bold. EPI – enhanced product ion for complete MS/MS spectra.

| Q1 Mass | Q3 Mass | Analyte | DP | EP | CE | CXP | IS Assignment |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 377.0 | 201.0 | Chlorophacinone-d4 (IS) | -120 | -10 | -32 | -15 |  |
| 312.0 | 161.0 | Warfarin-d5 (IS) | -95 | -10 | -28 | -9 |  |
| 352.0 | 265.0 | Acenocoumarol (Surr) | -70 | -10 | -40 | -13 | Warfarin-d5 |
| 341.0 | 160.0 | Coumachlor (Surr) | -60 | -10 | -30 | -21 | Warfarin-d5 |
| 343.0 | 167.0 | Diphacinone-d4 (Surr) | -115 | -10 | -32 | -15 | Chlorophacinone-d4 |
| 279.0 | 250.0 | Phenprocoumon (Surr) | -55 | -10 | -32 | -17 | Warfarin-d5 |
| 521.0 | 78.8 | Brodifacoum | -20 | -10 | -128 | -11 | Warfarin-d5 |
| 521.0 | 135.0 |  | -180 | -10 | -46 | -17 | Warfarin-d5 |
| 527.0 | 250.0 | Bromadiolone | -30 | -10 | -50 | -19 | Warfarin-d5 |
| 525.0 | 181.0 |  | -85 | -10 | -46 | -9 | Warfarin-d5 |
| 373.0 | 201.0 | Chlorophacinone | -75 | -10 | -30 | -13 | Chlorophacinone-d4 |
| 291.0 | 141.0 | Coumatetralyl | -125 | -10 | -38 | -13 | Warfarin-d5 |
| 291.0 | 247.0 |  | -125 | -10 | -32 | -21 | Warfarin-d5 |
| 443.0 | 135.0 | Difenacoum | -55 | -10 | -46 | -15 | Warfarin-d5 |
| 443.0 | 93.0 |  | -55 | -10 | -80 | -15 | Warfarin-d5 |
| 539.0 | 80.8 | Difethialone | -20 | -10 | -92 | -13 | Warfarin-d5 |
| 539.0 | 151.0 |  | -20 | -10 | -46 | -13 | Warfarin-d5 |
| 541.0 | 382.0 | Flocoumafen | -65 | -10 | -36 | -29 | Warfarin-d5 |
| 541.0 | 161.0 |  | -65 | -10 | -48 | -17 | Warfarin-d5 |
| 307.0 | 161.0 | Warfarin | -45 | -10 | -26 | -19 | Warfarin-d5 |
| 307.0 | 250.0 |  | -90 | -10 | -32 | -19 | Warfarin-d5 |
| EPI- mass range |  |  | DP | EP | CE |  |  |
| [m/z] |  |  | [V] |  |  |  |  |
| 50 - 450 |  |  | -50 | -10 | -30(±15) |  |  |

Table S2_c1: FIRST ANALYSIS CAMPAIGN: Validation – Recovery with pig liver (n = 5), Limit of quantification (LOQ) and recovery levels (Rec) ± relative standard deviations (RSD). No findings in control samples (n = 2).

| Analyte | LOQ | 5 [µg kg-1] |  | 10 [µg kg-1] |  | 1000 [µg kg-1] |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | [µg kg-1] | Rec | RSD | Rec | RSD | Rec | RSD |
|  |  | [%] | [%] | [%] | [%] | [%] | [%] |
| Brodifacoum | 2.0 | 60 | 9 | 65 | 7 | 65 | 3 |
| Bromadiolone | 2.0 | 76 | 13 | 81 | 7 | 83 | 3 |
| Chlorophacinone | 10.0 | 37 | 18 | 63 | 8 | 81 | 3 |
| Coumatetralyl | 1.0 | 94 | 7 | 95 | 8 | 99 | 4 |
| Difenacoum | 1.0 | 85 | 6 | 85 | 5 | 86 | 3 |
| Difethialone | 1.0 | 45 | 6 | 52 | 12 | 53 | 1 |
| Flocoumafen | 1.0 | 80 | 7 | 82 | 6 | 80 | 1 |
| Warfarin | 0.5 | 101 | 7 | 102 | 5 | 107 | 4 |

Table S2_c2: FIRST ANALYSIS CAMPAIGN: Validation – Hunted liver samples batch (Recovery of Surrogates, 0.1 µg/sample) in quality control samples (pig liver, n = 17) and in hunted liver samples (n = 88), recovery levels (Rec) ± relative standard deviations (RSD).

| Surrogate | Quality control |  | Hunted liver samples |  |
| --- | --- | --- | --- | --- |
|  | Rec | RSD | Rec | RSD |
|  | [%] | [%] | [%] | [%] |
| Acenocoumarol | 99 | 10 | 98 | 13 |
| Coumachlor | 100 | 9 | 97 | 12 |
| Diphacinone-d4 | 81 | 8 | 91 | 15 |
| Phenprocoumon | 84 | 9 | 75 | 14 |

Table S2_d1: SECOND ANALYSIS CAMPAIGN: Validation – Recovery with calf liver (n = 10), LOQ and recovery levels (Rec) ± relative standard deviations (RSD). No findings in control samples (n = 2).

| Analyte | LOQ | Quality control 100 [µg kg-1] |  |  |
| --- | --- | --- | --- | --- |
|  |  | Rec | RSD |
|  | [µg kg-1] | [%] | [%] |
| Brodifacoum | 1.0 | 93 | 26 |
| Bromadiolone | 1.0 | 90 | 26 |
| Chlorophacinone | 2.0 | 109 | 8 |
| Coumatetralyl | 0.2 | 112 | 12 |
| Difenacoum | 0.2 | 90 | 35 |
| Difethialone | 1.0 | 80 | 34 |
| Flocoumafen | 0.4 | 87 | 41 |
| Warfarin | 0.2 | 128 | 7 |

Table S2_d2: SECOND ANALYSIS CAMPAIGN: Validation – Hunted liver samples batch (Recovery of Surrogates, 0.1 µg/sample) in quality control samples (calf liver, n = 10) and in hunted liver samples (n = 97), recovery levels (Rec) ± relative standard deviations (RSD).

| Surrogate | Quality control |  | Hunted liver samples |  |
| --- | --- | --- | --- | --- |
|  | Rec | RSD | Rec | RSD |
|  | [%] | [%] | [%] | [%] |
| Acenocoumarol | 127 | 12 | 114 | 13 |
| Coumachlor | 128 | 8 | 115 | 12 |
| Diphacinone-d4 | 128 | 29 | 111 | 28 |
| Phenprocoumon | 108 | 13 | 93 | 17 |

Table S3_a: Configuration of GC-MS/MS for the analysis of herbicide, herbicide safener, insecticide, fungicide, nematicide, microbiocide, plant growth regulator, PCB

| GAS CHROMATOGRAPHY | Agilent 7890 |  |  |  |
| --- | --- | --- | --- | --- |
| Autosampler temperature | room temperature |  |  |  |
| Injection volume | 1 µL |  |  |  |
| Analytical column | 2 Agilent HP-5ms ultra inert columns (15 m × 0.25 mm × 0.25 μm) |  |  |  |
| Split ratio | splitless |  |  |  |
| Inlet gradient program | Time (min) | Rate (°C min-1) | Value (°C) | Hold time (min) |
|  | 0 | - | 60 | 0.2 |
|  | 0.2 | 900 | 280 | 0 |
| Oven gradient program | Time (min) | Rate (°C min-1) | Value (°C) | Hold time (min) |
|  | 0 | - | 60 | 1 |
|  | 1 | 40 | 170 | 0 |
|  | 3.75 | 10 | 310 | 3 |
| Run time | 20.75 min |  |  |  |
| Post run time | 5 min |  |  |  |
| GC cycle time | 31 min |  |  |  |
| Carrier gas | Helium 5.0 (Air Liquide) |  |  |  |
| Carrier gas flow rate | Column 1: 0.9 mL min-1 Column 2: 1.1 mL min-1 |  |  |  |
| Back flush flow rate | Column 1: -4.0 mL min-1 Column 2: 4.4 mL min-1 |  |  |  |
| Vent flow | 100 mL min-1 |  |  |  |
| MASS SPEKTROMETER | 7010B (AGILENT) |  |  |  |
| Mode | EI, positive |  |  |  |
| Ion source voltage | 70 eV |  |  |  |
| Source temperature | 230 °C |  |  |  |
| Transfer line temperature | 280 °C |  |  |  |
| Quadrupole temperature | 150° Q1 and Q2 |  |  |  |
| Collision gas | Argon 4.5 (Air Liquide) 0.9 mL min-1 |  |  |  |
| Quench gas | Helium 5.0 (Air Liquide) 2.25 mL min-1 |  |  |  |
| Solvent delay | 4 min |  |  |  |
| Scan type | multiple reaction monitoring (MRM) |  |  |  |
| Software | Agilent Mass Hunter B.08.00 |  |  |  |
| Quantification | relative peak area |  |  |  |

S4 PFAS

Table S4_a: Configuration of LC-MS/MS for the analysis of PFAS

| LIQUID CHROMATOGRAPHY | Agilent 1200 |  |  |
| --- | --- | --- | --- |
| Autosampler temperature | room temperature |  |  |
| Injection volume | 10 µL |  |  |
| Analytical column | Restek Raptor Polar X (50 × 2.1 mm, 2.5 µm) |  |  |
| Column temperature | 40 °C |  |  |
| Mobile phase A | 0.5% (v/v) formic acid in water |  |  |
| Mobile phase B | 0.5% (v/v) formic acid in acetonitril |  |  |
| Gradient program | isocratic | 25% A | 75% B |
| Flow rate | 0.5 mL min-1 |  |  |
| MASS SPEKTROMETER | API3200 (SCIEX) |  |  |
| Mode | negative ESI |  |  |
| Ion spray potential | -4500 V |  |  |
| Source temperature | 500 °C |  |  |
| Scan type | multiple reaction monitoring (MRM) |  |  |
| Dwell time | Quantifier 100 ms Qualifier 20 ms |  |  |
| Software | Analyst 1.6.1 |  |  |
| Quantification | relative peak area |  |  |

Table S4_b: LC-MS/MS - MRM-conditions of the analytes and internal standard for the analysis of PFAS; precursor (Q1), product ions (Q3) in m/z; declustering potential (DP), entrance potential (EP), collision energy (CE) and cell exit potential (CXP) in V; Quantifier transitions are marked in bold.

| Q1 Mass | Q3 Mass | Analyte | DP | EP | CE | CXP | IS Assignment |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 265.0 | 97.0 | Sodium dodecyl sulfate (SDS, IS) | -55 | -7 | -40 | 0 |  |
| 265.0 | 96.3 |  | -55 | -7 | -78 | -58 |  |
| 265.0 | 80.0 |  | -55 | -7 | -76 | 0 |  |
| 412.8 | 368.9 | Perfluorooctanoic acid (PFOA) | -30 | -4.5 | -16 | -4 | SDS |
| 412.8 | 219.0 |  | -30 | -4.5 | -20 | -4 |  |
| 412.8 | 168.9 |  | -30 | -4.5 | -22 | 0 |  |
| 498.8 | 129.9 | Perfluorooctaesulfonic acid (PFOS) | -100 | -9 | -58 | 0 | SDS |
| 498.8 | 98.8 |  | -100 | -9 | -58 | -2 |  |
| 498.8 | 80.0 |  | -100 | -9 | -82 | -2 |  |
| 398.7 | 118.9 | Perflurohexanesulfonic acid (PFHxS) | -75 | -5.5 | -46 | 0 | SDS |
| 398.7 | 98.9 |  | -75 | -5.5 | -48 | 0 |  |
| 398.7 | 79.9 |  | -75 | -5.5 | -78 | 0 |  |