Supplementary Table 1: Scope of instrument repairs conducted at a single centre

Total number of repairs successfully completed across eleven year period at The Royal London Hospital, by instrument type, indicating those performed at onsite repair centre versus those repaired externally offsite. ‘Further information/ parts indicates repair centre unable to identify fault, or sent with missing or incorrect parts. We found that 14,210 out of 20,150 (70.5%, Wilson binomial confidence interval 69.9%- 71.2%) of instruments sent for repair were successfully repaired (13,241=93% of successful repairs repaired onsite, 969=7% repaired offsite). Due to lack of instrument tracing, it was not possible to determine the number of repairs a given instrument underwent across its lifespan. Of those instruments successfully repaired (either on or offsite), over half were general surgical scissors such as Mayo, Metzenbaum or McIndoe scissors (7,320: 52%), followed by osteotomes (866: 6%), needle holders (853: 6%), retractors (553:4%), and clamps (520: 4%)

| Instrument | Repaired onsite |  | Repaired offsite |  | Total repairable |  | Further parts/ information |  | Irreparable (corrosion, damage, unusable) |  | Irreparable (require further decontamination) |  | Total |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Count | % | Count | % | Count | % | Count | % | Count | % | Count | % | Count |
| General scissors | 7318 | 84.6 | 2 | 0.0 | 7320 | 84.6 | 17 | 0.2 | 1281 | 14.8 | 30 | 0.3 | 8648 |
| Osteotomes/ chisels/ mallets | 861 | 86.9 | 5 | 0.5 | 866 | 87.4 | 7 | 0.7 | 108 | 10.9 | 10 | 1.0 | 991 |
| Needle holders | 652 | 34.8 | 201 | 10.7 | 853 | 45.5 | 147 | 7.8 | 863 | 46.0 | 12 | 0.6 | 1875 |
| Retractors | 515 | 56.0 | 38 | 4.1 | 553 | 60.2 | 104 | 11.3 | 259 | 28.2 | 3 | 0.3 | 919 |
| Clamps | 481 | 51.1 | 39 | 4.1 | 520 | 55.2 | 103 | 10.9 | 313 | 33.2 | 6 | 0.6 | 942 |
| Vessel scissors | 506 | 87.5 | 1 | 0.2 | 507 | 87.7 | 3 | 0.5 | 63 | 10.9 | 5 | 0.9 | 578 |
| Artery forceps | 501 | 46.1 | 0 | 0.0 | 501 | 46.1 | 64 | 5.9 | 516 | 47.5 | 5 | 0.5 | 1086 |
| Endoscope telescopes | 15 | 3.5 | 349 | 80.8 | 364 | 84.3 | 65 | 15.0 | 3 | 0.7 | 0 | 0.0 | 432 |
| Micro scissors | 340 | 83.5 | 6 | 1.5 | 346 | 85.0 | 6 | 1.5 | 55 | 13.5 | 0 | 0.0 | 407 |
| Micro forceps | 314 | 60.6 | 8 | 1.5 | 322 | 62.2 | 27 | 5.2 | 166 | 32.0 | 3 | 0.6 | 518 |
| Tissue forceps | 212 | 49.6 | 1 | 0.2 | 213 | 49.9 | 42 | 9.8 | 168 | 39.3 | 4 | 0.9 | 427 |
| Levers/ rasps/ curettes | 203 | 44.2 | 5 | 1.1 | 208 | 45.3 | 30 | 6.5 | 219 | 47.7 | 2 | 0.4 | 459 |
| Dissecting forceps | 201 | 49.5 | 2 | 0.5 | 203 | 50.0 | 23 | 5.7 | 177 | 43.6 | 3 | 0.7 | 406 |
| Endoscope instruments | 95 | 32.9 | 94 | 32.5 | 189 | 65.4 | 40 | 13.8 | 54 | 18.7 | 6 | 2.1 | 289 |
| Bone cutting/ holding | 136 | 52.1 | 35 | 13.4 | 171 | 65.5 | 20 | 7.7 | 63 | 24.1 | 7 | 2.7 | 261 |
| Rongeurs | 154 | 66.1 | 12 | 5.2 | 166 | 71.2 | 16 | 6.9 | 45 | 19.3 | 6 | 2.6 | 233 |
| Picks/ probes/ hooks | 159 | 56.2 | 1 | 0.4 | 160 | 56.5 | 39 | 13.8 | 83 | 29.3 | 1 | 0.4 | 283 |
| Pliers/ cutters/ benders/ spanners | 129 | 50.6 | 31 | 12.2 | 160 | 62.7 | 28 | 11.0 | 67 | 26.3 | 0 | 0.0 | 255 |
| Punches | 77 | 68.1 | 6 | 5.3 | 83 | 73.5 | 13 | 11.5 | 17 | 15.0 | 0 | 0.0 | 113 |
| Air/ battery tools | 31 | 46.3 | 24 | 35.8 | 55 | 82.1 | 5 | 7.5 | 5 | 7.5 | 2 | 3.0 | 67 |
| Scalpels/ knives/ snares | 51 | 37.8 | 4 | 3.0 | 55 | 40.7 | 16 | 11.9 | 59 | 43.7 | 5 | 3.7 | 135 |
| Debakey forcep/ clamp | 53 | 50.5 | 0 | 0.0 | 53 | 50.5 | 36 | 34.3 | 15 | 14.3 | 1 | 1.0 | 105 |
| Dilators/ introducers/ catheters | 36 | 20.1 | 11 | 6.1 | 47 | 26.3 | 30 | 16.8 | 102 | 57.0 | 0 | 0.0 | 179 |
| Electro-medical | 14 | 19.4 | 25 | 34.7 | 39 | 54.2 | 5 | 6.9 | 28 | 38.9 | 0 | 0.0 | 72 |
| Measuring / jigs | 26 | 50.0 | 7 | 13.5 | 33 | 63.5 | 8 | 15.4 | 11 | 21.2 | 0 | 0.0 | 52 |
| Needles | 27 | 77.1 | 4 | 11.4 | 31 | 88.6 | 0 | 0.0 | 4 | 11.4 | 0 | 0.0 | 35 |
| Fibre lights | 7 | 11.1 | 23 | 36.5 | 30 | 47.6 | 20 | 31.7 | 13 | 20.6 | 0 | 0.0 | 63 |
| Skin graft | 16 | 47.1 | 12 | 35.3 | 28 | 82.4 | 4 | 11.8 | 2 | 5.9 | 0 | 0.0 | 34 |
| Biopsy forceps | 20 | 31.7 | 7 | 11.1 | 27 | 42.9 | 9 | 14.3 | 23 | 36.5 | 4 | 6.3 | 63 |
| Suction | 24 | 52.2 | 0 | 0.0 | 24 | 52.2 | 7 | 15.2 | 15 | 32.6 | 0 | 0.0 | 46 |
| Syringes | 14 | 23.0 | 3 | 4.9 | 17 | 27.9 | 17 | 27.9 | 27 | 44.3 | 0 | 0.0 | 61 |
| Sterilisation cases | 11 | 61.1 | 4 | 22.2 | 15 | 83.3 | 2 | 11.1 | 1 | 5.6 | 0 | 0.0 | 18 |
| Table accessories | 10 | 38.5 | 5 | 19.2 | 15 | 57.7 | 6 | 23.1 | 5 | 19.2 | 0 | 0.0 | 26 |
| Trocar/ cannula | 12 | 46.2 | 2 | 7.7 | 14 | 53.8 | 3 | 11.5 | 9 | 34.6 | 0 | 0.0 | 26 |
| Laparoscopic scissors | 12 | 66.7 | 1 | 5.6 | 13 | 72.2 | 0 | 0.0 | 5 | 27.8 | 0 | 0.0 | 18 |
| Complete sets and kits | 4 | 36.4 | 1 | 9.1 | 5 | 45.5 | 4 | 36.4 | 2 | 18.2 | 0 | 0.0 | 11 |
| Vein strippers | 2 | 50.0 | 0 | 0.0 | 2 | 50.0 | 1 | 25.0 | 1 | 25.0 | 0 | 0.0 | 4 |
| Air tools | 1 | 50.0 | 0 | 0.0 | 1 | 50.0 | 1 | 50.0 | 0 | 0.0 | 0 | 0.0 | 2 |
| Measuring | 1 | 50.0 | 0 | 0.0 | 1 | 50.0 | 1 | 50.0 | 0 | 0.0 | 0 | 0.0 | 2 |
| Optical instrument | 0 | 0.0 | 0 | 0.0 | 0 | 0.0 | 0 | 0.0 | 9 | 100.0 | 0 | 0.0 | 9 |
| Total | 13,241 | 65.7 | 969 | 4.8 | 14,210 | 70.5 | 969 | 4.8 | 4,856 | 24.1 | 115 | 0.6 | 20,150 |

Supplementary table 2: Life cycle inventory processes chosen.

Legend: Processes selected for surgical scissors. Ecoinvent (version 3.6)- allocation, cut-off by classification- unit library was selected for all processes. ‘Transformation’ versions of processes within SimaPro were chosen for stainless steel (which do not include transportation), and adapted to account for the country of manufacture and transportation method. The ‘market’ versions of processes were used for all other products and processes, which include global averages for transportation from raw material extraction to manufacture. GLO= Global, RER= Rest of Europe, ROW= rest of world.

| Material | Process Name |  |  | Used for process/ item; comments |
| --- | --- | --- | --- | --- |
| Cardboard | Corrugated board box {RoW}¦ market for corrugated board box ¦ Cut-off, U |  |  | Reusable scissor packaging |
| High density polyethylene | Polyethylene, high density, granulate {GLO}¦ market for ¦ Cut-off, U |  |  | Sterile barrier system for scissor decontamination |
| Low density polyethylene | Packaging film, low density polyethylene {GLO}¦ market for ¦ Cut-off, U |  |  | Reusable scissor packaging, scissor repair (onsite and offsite) |
| Paper | Kraft paper, bleached {GLO}¦ market for ¦ Cut-off, U |  |  | Sterile barrier system for scissor decontamination |
| Paraffin | Paraffin {GLO}¦ market for ¦ Cut-off, U |  |  | Scissor repair (onsite and offsite); Oil spray for surgical instruments assumed to be paraffin |
| Polypropylene | Polypropylene, granulate {GLO}¦ market for ¦ Cut-off, U |  |  | Sterile barrier system for scissor decontamination |
| Polystyrene | Polystyrene, general purpose {GLO}¦ market for ¦ Cut-off, U |  |  | Reusable scissor packaging |
| Potassium hydroxide | Potassium hydroxide {GLO}¦ market for ¦ Cut-off, U |  |  | Virudet Eco sterilisation detergent; detergent packaging indicates contains 5% potassium silicate- not available in Ecoinvent, potassium hydroxide used instead |
| Natural gas (German) | Natural gas, high pressure {DE}¦ market for ¦ Cut-off, U |  |  | Manufacture of steel (German); high pressure is default option |
| Natural gas (UK) | Natural gas, high pressure {GB}¦ market for ¦ Cut-off, U |  |  | Decontamination reusable scissors; high pressure is default option |
| Natural gas (for steam production) | Process steam from natural gas, heat plant, consumption mix, at plant, MJ GB S |  |  | Production of steam from natural gas; used European Life Cycle Database, as steam processing not available in Ecoinvent |
| Stainless steel (German) | Steel, chromium steel 18/8 {RER}¦ steel production, converter, chromium steel 18/8 ¦ Cut-off, U |  |  | Reusable scissor manufacture, process copied and adapted for German manufacture (electricity, water, natural gas altered); Chromium steel 18/8 contains 18% chromium and 8% Nickel, so this process chosen as closest match- surgical instruments contain around 11.5-19% chromium, and 0.5-13.5% Nickel(1) |
| Stainless steel | Steel, chromium steel 18/8 {GLO}¦ market for ¦ Cut-off, U |  |  | Sterile barrier system for scissor decontamination |
| Sodium hydroxide | Sodium hydroxide, without water, in 50% solution state {GLO}¦ market for ¦ Cut-off, U |  |  | Virudet Eco sterilisation detergent; detergent packaging indicates contains 5% sodium hydroxide |
| Tissue paper | Tissue paper {GLO}¦ market for ¦ Cut-off, U |  |  | Scissor repair (onsite and offsite) |
| Energy | Process Name |  |  | Used for process/ item; comments |
| Electricity (German) | Electricity, medium voltage {DE}¦ market for ¦ Cut-off, U |  |  | Adapting manufacture stainless steel process (German); medium voltage chosen as is this is used for production industry |
| Electricity (UK) | Electricity, medium voltage {GB}¦ market for ¦ Cut-off, U |  |  | Decontamination reusable scissors, scissor repair (onsite and offsite); as above |
| Transportation | Process Name |  |  | Used for process/ item; comments |
| Courier (Europe) | Transport, freight, light commercial vehicle {RER}¦ market group for transport, freight, light commercial vehicle ¦ Cut-off, U |  |  | Transport via courier within Europe, across scenarios |
| Flight (short haul) | Transport, freight, aircraft, short haul {GLO}¦ market for transport, freight, aircraft, short haul ¦ Cut-off, U |  |  | Transport via flight from Germany to UK |
| Heavy goods vehicle (Europe) | Transport, freight, lorry, unspecified {RER}¦ market for transport, freight, lorry, unspecified ¦ Cut-off, U |  |  | Transport via heavy goods vehicle within Europe, across scenarios |
| Train | Transport, freight train {RER}¦ market group for transport, freight train ¦ Cut-off, U |  |  | Transport via train reusable scissors within Germany |
| Water | Process Name |  |  | Used for process/ item; comments |
| Tap water (German, UK) | Tap water {RER}¦ market group for ¦ Cut-off, U |  |  | Adapting manufacture stainless steel process (German), decontamination reusable scissors (UK) |
| Wastewater (German, UK) | Wastewater, average {Europe without Switzerland}¦ market for wastewater, average ¦ Cut-off, U |  |  | Adapting manufacture stainless steel process (German), decontamination reusable scissors (UK) |
| Processing | Process 1 Name | Process 2 Name (Subprocesses of process 1) | Process 3 Name (Subprocesses of process 2) | Used for process/ item; comments |
| Manufacture stainless steel (German) | Metal working, average for chromium steel product manufacturing {RER}¦ processing ¦ Cut-off, U | Energy and auxilliary inputs, metal working factory {RER}¦ market for energy and auxilliary inputs, metal working factory ¦ Cut-off, U | Energy and auxilliary inputs, metal working factory {RER}¦ with heating from hard coal ¦ Cut-off, U | Manufacture reusable scissors; adapted for German manufacture through changing electricity source to German supply |
|  |  |  | Energy and auxilliary inputs, metal working factory {RER}¦ with heating from heavy fuel oil ¦ Cut-off, U |  |
|  |  |  | Energy and auxilliary inputs, metal working factory {RER}¦ with heating from light fuel oil ¦ Cut-off, U |  |
|  |  |  | Energy and auxilliary inputs, metal working factory {RER}¦ with heating from natural gas ¦ Cut-off, U |  |
|  |  | Energy and auxilliary inputs, metal working machine {RER}¦ market for energy and auxilliary inputs, metal working machine ¦ Cut-off, U | Energy and auxilliary inputs, metal working machine {RER}¦ with process heat from hard coal ¦ Cut-off, U |  |
|  |  |  | Energy and auxilliary inputs, metal working machine {RER}¦ with process heat from heavy fuel oil ¦ Cut-off, U |  |
|  |  |  | Energy and auxilliary inputs, metal working machine {RER}¦ with process heat from light fuel oil ¦ Cut-off, U |  |
|  |  |  | Energy and auxilliary inputs, metal working machine {RER}¦ with process heat from natural gas ¦ Cut-off, U |  |
|  |  | Metal working factory {RER}¦ construction ¦ Cut-off, U |  |  |
|  |  | Metal working machine, unspecified {RER}¦ market for metal working machine, unspecified ¦ Cut-off, U |  |  |
|  |  | Steel, chromium steel 18/8, hot rolled {GLO}¦ market for ¦ Cut-off, U |  | Steel ‘wasted’ within the manufacturing process |

1.	British Stainless Steel Association. Selection of stainless steels for surgical instruments. Available at: https://bssa.org.uk/bssa_articles/9-selection-of-stainless-steels-for-surgical-instruments/ (accessed 13 April 2021).

Supplementary table 3: Parameters for decontamination of reusable surgical scissors

Legend: Parameters for decontamination of reusable surgical scissors. kWh= kilowatt hours, m3= meters cubed, kg= kilograms. Data based on research published by the author group elsewhere.(2) The mean weight of instruments across three typical sets used for common operations (tonsillectomy set, minor op set used for carpal tunnel decompression, and basic major orthopaedic set) was 66.7 g. The mean number of instruments per set across sets was 29 (determined using a retrospective audit of instrument decontaminations over one year at our local hospital conducted 1/7/18- 30/6/19). A typical surgical set was therefore assumed to contain 2 kg instruments, allowing the energy and material inputs to be determined.

| Process | Sub-process | Input | Amount (per kg of surgical instruments decontaminated) |
| --- | --- | --- | --- |
| Sterile barrier system | Basket | Stainless steel | 4.54 g |
|  | Inner wrap | Polypropylene | 27.98 g |
|  | Identification tag | High Density Polyethylene | 0.09 g |
|  | Outer wrap | Paper | 38.39 g |
|  | Label and tape |  |  |
|  | Kit list |  |  |
| Decontamination (washing and disinfection) | Detergent | Sodium hydroxide in 50% solution state | 0.93 g |
|  |  | Potassium hydroxide | 0.46 g |
|  | Electricity | Electricity (UK) | 0.63 kWh |
|  | Fuel | Natural gas (UK) | 0.15 m3 =5.76MJ =1.60kWh |
|  | Water | Tap water (UK) | 38.1 kg |
|  |  | Wastewater (UK) |  |

2. Rizan C, Lillywhite R, Reed M, Bhutta M (2021 C) Minimising carbon footprint and financial costs of decontamination (steam sterilisation) and packaging reusable surgical instruments. British Journal of Surgery. Online ahead of print https://academic.oup.com/bjs/advance-article/doi/10.1093/bjs/znab406/6445122 Accessed 13 Dec 2021. https://doi.org/10.1093/bjs/znab406

Supplementary Table 4: Alternative waste scenarios

Legend: Processes selected for waste scenarios. Ecoinvent (version 3.6)- allocation, cut-off by classification- unit library was selected for all processes. ‘Waste types’ as defined in Ecoinvent (version 3.6)- allocation, cut-off by classification- unit. GLO= global, PE= polyethylene, PET= polyethylene terephthalate, PP= polypropylene, PS= polystyrene.

| Waste Scenario | Process Name | Waste type applied to |
| --- | --- | --- |
| Hazardous incineration | Hazardous waste, for incineration {Europe without Switzerland}¦ treatment of hazardous waste, hazardous waste incineration ¦ Cut-off, U | All waste types + waste streams remaining after separation (100%) |
| Recycling | Aluminium (waste treatment) {GLO}¦ recycling of aluminium ¦ Cut-off, U | Non-ferro (100%) |
|  | Core board (waste treatment) {GLO}¦ recycling of core board ¦ Cut-off, U | Cardboard (100%) |
|  | Mixed plastics (waste treatment) {GLO}¦ recycling of mixed plastics ¦ Cut-off, U | Waste streams remaining after separation (100%) |
|  | Paper (waste treatment) {GLO}¦ recycling of paper ¦ Cut-off, U | Paper (100%) |
|  | PE (waste treatment) {GLO}¦ recycling of PE ¦ Cut-off, U | PE (100%) |
|  | PET (waste treatment) {GLO}¦ recycling of PET ¦ Cut-off, U | PET (100%) |
|  | PP (waste treatment) {GLO}¦ recycling of PP ¦ Cut-off, U | PP (100%) |
|  | PS (waste treatment) {GLO}¦ recycling of PS ¦ Cut-off, U | PS (100%) |
|  | Steel and iron (waste treatment) {GLO}¦ recycling of steel and iron ¦ Cut-off, U | Steel (100%) |
| Segregation (landfill) | Core board (waste treatment) {GLO}¦ recycling of core board ¦ Cut-off, U | Cardboard (100%) |
|  | Hazardous waste, for incineration {Europe without Switzerland}¦ treatment of hazardous waste, hazardous waste incineration ¦ Cut-off, U | Steel (100%) |
|  | Inert waste {Europe without Switzerland}¦ treatment of inert waste, sanitary landfill ¦ Cut-off, U | PE (100%), PET (100%), PP (100%), PS (100%), Non-ferro (100%), Waste streams remaining after separation (100%) |
|  | Paper (waste treatment) {GLO}¦ recycling of paper ¦ Cut-off, U | Paper (100%) |
| Segregation (municipal incineration) | Core board (waste treatment) {GLO}¦ recycling of core board ¦ Cut-off, U | Cardboard (100%) |
|  | Hazardous waste, for incineration {Europe without Switzerland}¦ treatment of hazardous waste, hazardous waste incineration ¦ Cut-off, U | Steel (100%) |
|  | Municipal solid waste {GB}¦ treatment of, incineration ¦ Cut-off, U | PE (100%), PET (100%), PP (100%), PS (100%), Non-ferro (100%), Waste streams remaining after separation (100%) |
|  | Paper (waste treatment) {GLO}¦ recycling of paper ¦ Cut-off, U | Paper (100%) |

Supplementary Table 5: Environmental impact of surgical scissors (midpoint categories): normalised results. Normalised results reported as environmental impact relative to the global average person’s contribution to the impact category over one hour, comparing base-line scenarios. All results reported per scissor use, with highest impact per category in red and lowest in green.

| Impact category | Reusable scissors, no repair | Reusable scissors, offsite repair | Reusable scissors, onsite repair |
| --- | --- | --- | --- |
| Global warming | 0.08 | 0.06 | 0.06 |
| Stratospheric ozone depletion | 0.00 | 0.00 | 0.00 |
| Ionizing radiation | 0.19 | 0.19 | 0.18 |
| Ozone formation, Human health | 0.05 | 0.04 | 0.04 |
| Fine particulate matter formation | 0.03 | 0.02 | 0.02 |
| Ozone formation, Terrestrial ecosystems | 0.06 | 0.04 | 0.04 |
| Terrestrial acidification | 0.03 | 0.02 | 0.02 |
| Freshwater eutrophication | 0.23 | 0.14 | 0.14 |
| Marine eutrophication | 0.03 | 0.03 | 0.03 |
| Terrestrial ecotoxicity | 3.34 | 1.05 | 1.04 |
| Freshwater ecotoxicity | 14.25 | 7.96 | 7.79 |
| Marine ecotoxicity | 23.28 | 12.81 | 12.54 |
| Human carcinogenic toxicity | 15.17 | 6.70 | 6.63 |
| Human non-carcinogenic toxicity | 2.74 | 1.71 | 1.68 |
| Land use | 0.01 | 0.01 | 0.01 |
| Mineral resource scarcity | 0.00 | 0.00 | 0.00 |
| Fossil resource scarcity | 0.21 | 0.19 | 0.19 |
| Water consumption | 0.02 | 0.02 | 0.02 |

Supplementary Table 6: Environmental impact of surgical scissors (midpoint categories): normalised results. Normalised results reported as environmental impact relative to the global average person’s contribution to the impact category over one hour, comparing base-line scenarios. All results reported per scissor use, with highest impact per category in red and lowest in green.

| Damage category | Reusable scissors, no repair | Reusable scissors, offsite repair | Reusable scissors, onsite repair |
| --- | --- | --- | --- |
| Human health | 0.0539 | 0.0366 | 0.0360 |
| Ecosystems | 0.0038 | 0.0031 | 0.0030 |
| Resources | 0.0024 | 0.0022 | 0.0022 |

Supplementary Table 7. Environmental impact of one use of surgical scissors-scenario uncertainty modelling

Legend: Environmental impact of scenario uncertainty models. 1,4-DCB =dichlorobenzene, CFC11= Trichlorofluoromethane, CO2e= carbon dioxide equivalents, Cu= copper, eq= equivalents, kBq Co-60 eq = kilobecquerel Cobalt-60, m2a = square meter years, N= nitrogen, NOx= nitrous oxides, P=phosphate, PM2.5 = particulate matter <2.5 micrometers, SO2= sulphur dioxide. Where uncertainty scenario not applicable to the base scenario, the base scenario itself should be used as the point of comparison. Base scenarios are compared with the following models, in which only one assumption is changed. All results reported per scissor use, with highest impact per category in red and lowest in green.

- 1: reducing number of reuses of reusables to 10 (from 40)

- 2: increasing number of reuses of reusables to 400 (from 40)

- 3: reducing number of repairs to once (from nine)

- 4: Increased distance of travel for offsite recycling to 800km (from 80km)

- 5: Electricity supply with lower proportion of renewables (Australian) assumed for decontamination of reusable instruments

- 6: all waste assumed to be recycled

- 7: waste segregation, with sharps waste (scissors) sent for high temperature incineration, cardboard and paper sent for recycling, and all other waste sent to landfill

- 8: waste segregation, with sharps waste (scissors) sent for high temperature incineration, cardboard and paper sent for recycling, and all other waste sent to municipal incineration

Supplementary table 7A: Reusable scissors (no repair) base scenario and alternative assumption models

| Impact category | Unit | Base scenario | 1 | 2 | 5 | 6 | 7 | 8 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Global warming | g CO2e | 70.3 | 121.7 | 54.9 | 95.8 | 65.5 | 69.7 | 69.9 |
| Stratospheric ozone depletion | g CFC11 eq | 0.000026 | 0.000044 | 0.000021 | 0.000050 | 0.000024 | 0.000026 | 0.000026 |
| Ionizing radiation | Bq Co-60 eq | 10 | 13 | 10 | 2 | 10 | 10 | 10 |
| Ozone formation, Human health | g NOx eq | 0.11 | 0.22 | 0.08 | 0.16 | 0.11 | 0.11 | 0.11 |
| Fine particulate matter formation | g PM2.5 eq | 0.08 | 0.21 | 0.05 | 0.11 | 0.08 | 0.08 | 0.08 |
| Ozone formation, Terrestrial ecosystems | g NOx eq | 0.12 | 0.23 | 0.08 | 0.17 | 0.11 | 0.12 | 0.12 |
| Terrestrial acidification | g SO2 eq | 0.15 | 0.31 | 0.10 | 0.24 | 0.14 | 0.15 | 0.15 |
| Freshwater eutrophication | g P eq | 0.02 | 0.04 | 0.01 | 0.08 | 0.02 | 0.02 | 0.02 |
| Marine eutrophication | g N eq | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 |
| Terrestrial ecotoxicity | g 1,4-DCB | 395 | 1,310 | 120 | 401 | 390 | 394 | 394 |
| Freshwater ecotoxicity | g 1,4-DCB | 2.0 | 5.1 | 1.1 | 3.6 | 1.9 | 2.0 | 2.0 |
| Marine ecotoxicity | g 1,4-DCB | 2.7 | 7.1 | 1.4 | 5.0 | 2.6 | 2.7 | 2.7 |
| Human carcinogenic toxicity | g 1,4-DCB | 4.8 | 13.9 | 2.1 | 8.1 | 4.3 | 4.7 | 4.7 |
| Human non-carcinogenic toxicity | g 1,4-DCB | 47 | 109 | 28 | 112 | 44 | 46 | 46 |
| Land use | m2a crop eq | 0.0051 | 0.0066 | 0.0046 | 0.0035 | 0.0051 | 0.0051 | 0.0051 |
| Mineral resource scarcity | g Cu eq | 1.3 | 4.4 | 0.3 | 1.3 | 1.3 | 1.3 | 1.3 |
| Fossil resource scarcity | g oil eq | 23 | 34 | 20 | 28 | 23 | 23 | 23 |
| Water consumption | m3 | 0.0006 | 0.0008 | 0.0005 | 0.0006 | 0.0006 | 0.0006 | 0.0006 |

Supplementary Table 7B: Reusable scissors repaired offsite, base scenario and alternative assumption models

| Impact category | Unit | Base scenario | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Global warming | g CO2e | 57.0 | 68.3 | 53.6 | 62.9 | 57.1 | 83.1 | 56.5 | 56.9 | 56.9 |
| Stratospheric ozone depletion | g CFC11 eq | 0.000021 | 0.000025 | 0.000020 | 0.000023 | 0.000021 | 0.000046 | 0.000021 | 0.000021 | 0.000021 |
| Ionizing radiation | Bq Co-60 eq | 10 | 11 | 10 | 10 | 10 | 1 | 10 | 10 | 10 |
| Ozone formation, Human health | g NOx eq | 0.09 | 0.11 | 0.08 | 0.10 | 0.09 | 0.14 | 0.09 | 0.09 | 0.09 |
| Fine particulate matter formation | g PM2.5 eq | 0.05 | 0.07 | 0.04 | 0.07 | 0.05 | 0.08 | 0.05 | 0.05 | 0.05 |
| Ozone formation, Terrestrial ecosystems | g NOx eq | 0.09 | 0.11 | 0.08 | 0.10 | 0.09 | 0.14 | 0.09 | 0.09 | 0.09 |
| Terrestrial acidification | g SO2 eq | 0.10 | 0.13 | 0.09 | 0.12 | 0.10 | 0.19 | 0.10 | 0.10 | 0.10 |
| Freshwater eutrophication | g P eq | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.08 | 0.01 | 0.01 | 0.01 |
| Marine eutrophication | g N eq | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 |
| Terrestrial ecotoxicity | g 1,4-DCB | 124 | 229 | 93 | 245 | 127 | 131 | 124 | 124 | 124 |
| Freshwater ecotoxicity | g 1,4-DCB | 1.1 | 1.6 | 1.0 | 1.5 | 1.1 | 2.8 | 1.1 | 1.1 | 1.1 |
| Marine ecotoxicity | g 1,4-DCB | 1.5 | 2.2 | 1.3 | 2.1 | 1.5 | 3.8 | 1.5 | 1.5 | 1.5 |
| Human carcinogenic toxicity | g 1,4-DCB | 2.1 | 3.2 | 1.8 | 3.3 | 2.1 | 5.5 | 2.1 | 2.1 | 2.1 |
| Human non-carcinogenic toxicity | g 1,4-DCB | 29 | 39 | 26 | 37 | 29 | 96 | 29 | 29 | 29 |
| Land use | m2a crop eq | 0.0049 | 0.0058 | 0.0046 | 0.0050 | 0.0049 | 0.0033 | 0.0049 | 0.0049 | 0.0049 |
| Mineral resource scarcity | g Cu eq | 0.3 | 0.7 | 0.2 | 0.8 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 |
| Fossil resource scarcity | g oil eq | 21 | 25 | 20 | 22 | 21 | 26 | 21 | 21 | 21 |
| Water consumption | m3 | 0.0005 | 0.0006 | 0.0005 | 0.0005 | 0.0005 | 0.0005 | 0.0005 | 0.0005 | 0.0005 |

Supplementary Table 7C: Reusable scissors repaired onsite: base scenario and alternative assumption models

| Impact category | Unit | Base scenario | 1 | 2 | 3 | 5 | 6 | 7 | 8 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Global warming | g CO2e | 56.3 | 65.7 | 53.5 | 62.5 | 82.4 | 55.8 | 56.2 | 56.3 |
| Stratospheric ozone depletion | g CFC11 eq | 0.000021 | 0.000024 | 0.000020 | 0.000023 | 0.000045 | 0.000021 | 0.000021 | 0.000021 |
| Ionizing radiation | Bq Co-60 eq | 10 | 11 | 10 | 10 | 1 | 10 | 10 | 10 |
| Ozone formation, Human health | g NOx eq | 0.08 | 0.10 | 0.08 | 0.10 | 0.14 | 0.08 | 0.08 | 0.08 |
| Fine particulate matter formation | g PM2.5 eq | 0.05 | 0.07 | 0.04 | 0.06 | 0.08 | 0.05 | 0.05 | 0.05 |
| Ozone formation, Terrestrial ecosystems | g NOx eq | 0.09 | 0.10 | 0.08 | 0.10 | 0.14 | 0.09 | 0.09 | 0.09 |
| Terrestrial acidification | g SO2 eq | 0.10 | 0.12 | 0.09 | 0.12 | 0.19 | 0.10 | 0.10 | 0.10 |
| Freshwater eutrophication | g P eq | 0.01 | 0.01 | 0.01 | 0.01 | 0.08 | 0.01 | 0.01 | 0.01 |
| Marine eutrophication | g N eq | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 |
| Terrestrial ecotoxicity | g 1,4-DCB | 122 | 221 | 93 | 243 | 129 | 122 | 122 | 122 |
| Freshwater ecotoxicity | g 1,4-DCB | 1.1 | 1.5 | 1.0 | 1.5 | 2.8 | 1.1 | 1.1 | 1.1 |
| Marine ecotoxicity | g 1,4-DCB | 1.5 | 2.0 | 1.3 | 2.0 | 3.8 | 1.5 | 1.5 | 1.5 |
| Human carcinogenic toxicity | g 1,4-DCB | 2.1 | 3.1 | 1.8 | 3.3 | 5.5 | 2.0 | 2.1 | 2.1 |
| Human non-carcinogenic toxicity | g 1,4-DCB | 29 | 37 | 26 | 37 | 95 | 28 | 29 | 29 |
| Land use | m2a crop eq | 0.0048 | 0.0053 | 0.0046 | 0.0049 | 0.0032 | 0.0048 | 0.0048 | 0.0048 |
| Mineral resource scarcity | g Cu eq | 0.3 | 0.7 | 0.2 | 0.8 | 0.3 | 0.3 | 0.3 | 0.3 |
| Fossil resource scarcity | g oil eq | 21 | 23 | 20 | 22 | 26 | 21 | 21 | 21 |
| Water consumption | m3 | 0.0005 | 0.0006 | 0.0005 | 0.5425 | 0.0005 | 0.0005 | 0.0005 | 0.0005 |