### Data preprocessing of Odor profile analysis

Prior to multivariate analysis, raw sensor signals underwent the following preprocessing steps:

- Baseline correction:
Each sensor response was normalized against its initial baseline value using:

R =  G-G0G0 × 100

where G = conductance at time t, and G₀ = initial conductance.

- Averaging interval selection:
To avoid transient noise, the steady-state region (last 20 s) of the response curve was extracted and averaged for each sensor.

- Autoscaling (z-score normalization):
Each sensor variable was then standardized using:

Z =  ϰ-µσ × 100

ensuring equal weighting of sensors prior to PCA.

- Outlier screening:
Mahalanobis distance was used to screen for outliers; no sample exceeded the 95% confidence limit.

PCA Validation and Classification Performance

To validate the robustness of the E-nose–based separation among reheating treatments, a supervised classification model was constructed using leave-one-out cross-validation (LOOCV) applied to a linear discriminant analysis (LDA) classifier. Prior to model training, the same preprocessed sensor matrix used for PCA (baseline-corrected, autoscaled, and averaged steady-state signals) was used as input.

The LDA–LOOCV model achieved the following results:

- Overall classification accuracy: 94.4%

- Misclassification rate: 5.6%

- Confusion matrix: Correct grouping was observed for all RF and RR samples; minor overlapping occurred between RB and RS, consistent with their proximity in the PCA plot.

To further assess grouping reliability, we computed:

- Permutation test (n = 1000 permutations): p < 0.01

- Cross-validated Mahalanobis distance: All dry-heat samples (RF, RR) showed significantly larger distances from moist-heat samples (RB, RS) (p < 0.05).

- Silhouette coefficient: 0.71, indicating strong cluster separation.

### Derivatization protocol of free amino acids

FAAs were derivatized online using o-phthalaldehyde (OPA) for primary amino acids and 9-fluorenylmethyl chloroformate (FMOC) for secondary amino acids, following the Agilent standard post-column derivatization procedure:

- OPA reagent: OPA + 2-mercaptoethanol in borate buffer

- FMOC reagent: FMOC in acetonitrile

- Reaction coil temperature: 40°C

- Reaction time: ~30 s prior to UV detection

### Chromatographic conditions

Analyses were conducted on an Agilent 1100 HPLC system equipped with a ZORBAX Eclipse AAA C18 column (4.6 × 150 mm, 5 μm). Mobile phases were:

- A: 40 mM sodium acetate buffer (pH 7.2)

- B: acetonitrile:methanol:water (45:45:10, v/v/v)

Gradient elution (Agilent AAA protocol):

- 0–1.9 min: 0% B

- 1.9–18.1 min: 0–57% B

- 18.1–18.6 min: 57–100% B

- 18.6–22.3 min: 100% B

- 22.3–26.0 min: re-equilibration at 0% B

Other instrument parameters:

- Flow rate: 1.0 mL/min

- Column temperature: 40°C

- Injection volume: 10 μL

- UV detection: 338 nm (OPA derivatives) and 262 nm (FMOC derivatives)

### Calibration and quantification

Quantification was performed using external calibration curves prepared from a certified amino acid standard mixture (Sigma-Aldrich A9906):

- Five concentration levels (0.5, 1, 5, 10, 20 μmol/L)

- Calibration fitted using least-squares linear regression

- Mean R² values: 0.991–0.999

### Identification criteria

Each amino acid was positively identified based on:

- Retention time match (±0.2 min) with standards

- UV absorption spectra for OPA/FMOC derivatives

- Derivatization response consistent with primary/secondary amine chemistry

- Peak purity checks using Agilent ChemStation software

### Quality control procedures

To ensure dataset reliability:

- System suitability was evaluated using standard mix injections every 10 samples.

- Method precision (repeatability): RSD < 5% for peak areas.

- Spike recovery test:

  - Fortification at 5 and 15 μmol/L

  - Recoveries ranged 92.4–105.1%, confirming accuracy

- Blank TCA extractions were analyzed to ensure no contamination.

- Duplicate injections were performed for each sample.

Results were expressed as mg FAA per 100 g of meat (mean ± SD).

Methodological Validation for Oxidation Assays

Validation of TBARS, Carbonyl, and Total Sulfhydryl Assays

To ensure analytical reliability, all oxidation assays (TBARS, carbonyls, and total sulfhydryls) were validated in accordance with standard meat-science QA/QC procedures.

1. Linearity and Calibration

- TBARS: Calibration was performed using MDA standard solutions (0–5 μmol/L). The curve showed excellent linearity (R² = 0.996).

- Carbonyls: Calibration was verified using DNPH-derivatized BSA standards across 0–10 nmol/mg. Linearity was high (R² = 0.992).

- Sulfhydryls: Validation was conducted using L-cysteine standard curves (0–50 μmol/L), yielding R² = 0.995.

2. Limit of Detection (LOD) and Limit of Quantification (LOQ)

LOD and LOQ were determined using signal-to-noise ratios of 3 and 10, respectively.

| Assay | LOD | LOQ |
| --- | --- | --- |
| TBARS | 0.02 mg MDA/kg | 0.06 mg MDA/kg |
| Carbonyls | 0.08 nmol/mg | 0.25 nmol/mg |
| Total SH | 0.10 nmol/mg | 0.30 nmol/mg |

3. Recovery and Accuracy

Recovery was assessed by spiking raw beef homogenates with known concentrations of analytes:

- TBARS recovery: 92.6–103.4%

- Carbonyl recovery: 88.1–101.7%

- Total SH recovery: 90.4–105.2%

All recovery values fell within the acceptable analytical range (80–110%).

4. Precision and Repeatability

Inter- and intra-assay precision were evaluated using six replicates:

- TBARS:

  - Intra-assay CV = 3.2%

  - Inter-assay CV = 4.7%

- Carbonyls:

  - Intra-assay CV = 4.1%

  - Inter-assay CV = 5.3%

- Total SH:

  - Intra-assay CV = 3.9%

  - Inter-assay CV = 5.0%

All CV values were < 10%, indicating strong precision.