	case_id
	finding_id
	Age Range
	Evaluation of AI Finding
	Diameter of AI Finding (in mm)
	Location
	Description

	A01
	A01_1
	41 – 45
	False Positive
	2.6
	Right M1 segment
	Infundibulum

	A02
	A02_1
	51 – 55
	True Positive
	2.8
	Left MCA bifurcation
	Saccular aneurysm

	A03
	A03_1
	51 – 55
	True Positive
	3.7
	Right MCA bifurcation
	Saccular aneurysm

	A03
	A03_2
	51 – 55
	False Positive
	1.7
	Left M1 segment
	Infundibulum

	A04
	A04_2
	56 – 60
	False Positive
	1.3
	Left M2 segment
	Vascular loop

	A05
	A05_1
	51 – 55
	True Positive
	4.6
	Right ICA
	Saccular aneurysm

	A06
	A06_1
	71 – 75
	False Positive
	2.8
	Right ICA
	Infundibulum

	A07
	A07_1
	66 – 70
	True Positive
	3.9
	Right ICA
	Saccular aneurysm

	A08
	A08_1
	81 – 85
	True Positive
	2.5
	Acom
	Saccular aneurysm

	A09
	A09_1
	76 - 80
	True Positive
	4.7
	Acom
	Saccular aneurysm

	A09
	A09_2
	76 – 80
	True Positive
	3.0
	Right MCA bifurcation
	Saccular aneurysm

	A10
	A10_1
	71 – 75
	False Positive
	2.9
	Acom
	Acom fenestration

	B01
	B01_1
	76 – 80
	False Positive
	1.4
	Right ICA
	Vascular loop

	B02
	B02_1
	16 – 20
	False Positive
	2.2
	Left M2 segment
	Vascular loop

	B03
	B03_1
	71 – 75
	True Positive
	5.6
	Left MCA bifurcation
	Saccular aneurysm

	B03
	B03_2
	71 – 75
	True Positive
	3.7
	Left ICA
	Saccular aneurysm

	B04
	B04_1
	56 – 60
	True Positive
	3.1
	Left MCA bifurcation
	Fusiform aneurysm

	B05
	B05_1
	31 – 35
	False Positive
	1.5
	Left M1 segment
	Perforator

	B06
	B06_1
	26 – 30 
	False Positive
	2.0
	Left M1 segment
	Perforator

	B07
	B07_1
	31 – 35
	True Positive
	4.0
	Left ICA
	Saccular aneurysm

	B08
	B08_1
	26 – 30
	False Positive
	1.2
	Acom
	Vascular loop

	B09
	B09_1
	56 – 60
	True Positive
	4.0
	Left MCA bifurcation
	Saccular aneurysm

	B10
	B10_1
	31 – 35
	False Positive
	2.2
	Left MCA bifurcation
	Vascular loop
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	Aneurysm Present
	Aneurysm Absent
	Predictive Value
	Total

	Test Positive
	True Positive
1.82% (182/10,000)
	False Positive
16.17% (1,617/10,000)
	Positive Predictive Value
11.3% (182/1,799)
	1,799

	Test Negative
	False Negative
0.18% (18/10,000)
	True Negative
81.83% (8,183/10,000)
	Negative Predictive Value
99.8% (8,183/8,201)
	8,201

	Accuracy Metrics
	Sensitivity
91.2% (182/200)
	Specificity
83.5% (8,183/9,800)
	
	

	Total
	200
	9,800
	
	10,000


Supplement 2: Contingency table for AI-based cerebral aneurysm detection. A real-world scenario with 10,000 cases was simulated. A prevalence of 2% was assumed. A sensitivity of 91.2% and specificity of 83.5% were assumed based on a recent meta-analysis (Din et al. 2023). These conditions result in a very low positive predictive value (PPV) of 11.3%, implying that only around one in nine AI-detected findings represents a true aneurysm. Numbers were rounded to the nearest whole number.
