I. Analytical procedure and parameters

The linear range of calibration function reached from the detection limit up 0.0; 1.0; 2.0; 5.0; 10.0 Pb µg/L and 0.0; 0.5; 1.0; 2.0 µg/L for Cd. The  values of the correlation coefficient (R) are good indicators of the linearity for AAS instruments giving precise and accurate results [7, 8]. The correlation coefficients obtained (0.998 for Pb and 0.998 for Cd) indicated that the analyses were both precise and accurate.

The recoveries were 98.2 % for Pb and 97.5 % for Cd. The recoveries were calculated as the quotient of the determined level and the known amount of the determined element expressed as a percentage. The calculated limits of detection (LODs) were 0.45 µg/L for Pb and 0.15 µg/L for Cd. The calculated limits of quantification (LOQs) were 0.95 µg/L for Pb and 0.34 µg/L for Cd.

II. Conditions of apparatus

1. The AAS operating conditions.

| Operating parameters | heavy metal |  |
| --- | --- | --- |
|  | Pb | Cd |
| Wavelength [nm] | 283.3 | 228.8 |
| Slit width [nm] | 0.7 | 0.7 |
| Lamp current [mA] | 8 | 5 |
| Optimum working range [µg/kg] | 1.0-10.0 | 0.02-0.20 |

2. Time-temperature program in the graphite furnace atomic absorption spectrometer in Pb and Cd determination.

Pb

| Step | Temperature [°C] | Ramp [s] | Hold [s] | Gas Flow [mL/min] |
| --- | --- | --- | --- | --- |
| 1 | 120 | 1 | 30 | 250 |
| 2 | 950 | 10 | 20 | 250 |
| 3 | 1450 | 0 | 5 | 0 |
| 5 | 2400 | 1 | 2 | 250 |

Cd

| Step | Temperature [°C] | Ramp [s] | Hold [s] | Gas Flow [mL·min-1] |
| --- | --- | --- | --- | --- |
| 1 | 120 | 10 | 25 | 250 |
| 2 | 300 | 5 | 15 | 250 |
| 3 | 1600 | 0 | 3 | 0 |
| 5 | 2400 | 1 | 2 | 250 |