Supplement 2:

Table S1: ICP-MS measurements of multiple elements in commercial reference materials for whole blood matrix (analytical values or acceptable ranges of which are provided). The given SD values are calculated from the total number of measurements (n) during a period of 10 months.

|  |  | Seronorm™ Trace Elements Whole Blood L-1 (lot: 1706877) |  | Seronorm™ Trace Elements Whole Blood L-2 (lot: 1706878) |  |
| --- | --- | --- | --- | --- | --- |
|  | LOD | Analytical/Target value | Measured value | Analytical/Target value | Measured value |
|  |  | (Acceptable range) | (N=51) | (Acceptable range) | (N=41) |
|  | (µg/L) | (µg/L) | (µg/L) | (µg/L) | (µg/L) |
| Mn | 0.20 | 16.3 | 15.6±2.0 | 22.5 | 23.6±1.7 |
|  |  | (13-19.6) |  | (20.7-24.3) |  |
| Cu | 0.30 | 650 | 588±44 | 920 | 855±7 |
|  |  | 520-780 |  | (830-1000) |  |
| Zn | 25 | 4500 | 4533±279 | 6200 | 6477±427 |
|  |  | (3600-5400) |  | (5000-7500) |  |
| As | 0.03 | 2.5 | 2.16±0.06 | 11.4 | 11.2±0.4 |
|  |  | (2.0-3.0) |  | (9.1-13.7) |  |
| Se | 0.40 | 67 | 66.1±2.4 | 91 | 97.7±5.9 |
|  |  | (47-87) |  | (63-118) |  |
| Mo | 0.04 | 0.28 | 0.30±0.05 | 5.3 | 5.3±0.3 |
|  |  | (0.22-0.33) |  | (4.3-6.4) |  |
| Cd | 0.01 | 0.29 | 0.272±0.006 | 5.0 | 5.0±0.1 |
|  |  | (0.23-0.35) |  | (4.0-6.0) |  |
| Hg | 0.04 | 1.63 | 1.41±0.08 | 15 | 14.3±1.4 |
|  |  | (1.30-1.96) |  | (12.0-18.0) |  |
| Pb | 0.04 | 10.3 | 9.17±0.33 | 295.0 | 280±15 |
|  |  | (8.2-12.4) |  | (264-325) |  |

Table S2: ICP-MS measurements of multiple elements in commercial reference materials for serum matrix (analytical values or acceptable ranges of which are provided). The given SD values are calculated from the total number of measurements (n) during a period of 10 months.

|  |  | Seronorm™ Trace Elements Serum L-1 (lot: 1801802) |  | Seronorm™ Trace Elements Serum L-2 (lot: 1801803) |  |
| --- | --- | --- | --- | --- | --- |
|  | LOD | Analytical/Target value | Measured value | Analytical/Target value | Measured value |
|  |  | (Acceptable range) | (N=52) | (Acceptable range) | (N=44) |
|  | (µg/L) | (µg/L) | (µg/L) | (µg/L) | (µg/L) |
| Mn | 0.2 | 9.4 | 10.0±0.8 | 14.2 | 15.4±0.5 |
|  |  | (7.5-11.2) |  | (11.3-17.0) |  |
| Cu | 0.4 | 1128 | 1061±89 | 2175 | 2124±38 |
|  |  | (901-1355) |  | (1737-2613) |  |
| Zn | 50 | 1460 | 1543±164 | 2090 | 2102±121 |
|  |  | (1160-1750) |  | (1670-2510) |  |
| As | 0.03 | 0.32 | 0.32±0.04 | 0.38 | 0.40±0.02 |
| Se | 0.2 | 95 | 89.4±8.5 | 139 | 142.0±5.5 |
|  |  | (76-114) |  | (111-167) |  |
| Mo | 0.04 | 0.41 | 0.36±0.01 | 0.91 | 0.95±0.07 |
| Cd | 0.01 | 0.135 | 0.15±0.01 | 0.122 | 0.127±0.003 |
| Hg | 0.04 | 1.20 | 1.16±0.17 | 2.27 | 2.21±0.19 |
|  |  | (0.84-1.57) |  | (1.59-2.95) |  |
| Pb | 0.03 | 0.75 | 0.82±0.07 | 1.29 | 1.26±0.06 |

Table S3: ICP-MS measurements of multiple elements in commercial reference materials for urine matrix (analytical values or acceptable ranges of which are provided). The given SD values are calculated from the total number of measurements (n) during a period of 10 months.

|  |  | ClinChek® Urine Control for Trace Elements, Level I (lot: 8847) |  | Seronorm™ Trace Elements Urine L-2 (lot: 1706878) |  |
| --- | --- | --- | --- | --- | --- |
|  | LOD | Analytical/Target value | Measured value | Analytical/Target value | Measured value |
|  |  | (Acceptable range) | (N=39) | (Acceptable range) | (N=36) |
|  | (µg/L) | (µg/L) | (µg/L) | (µg/L) | (µg/L) |
| Mn | 0.3 | 4.09 | 4.1±0.2 | 9.3 | 10.6±0.4 |
|  |  | (3.27-4.90) |  | (7.4-11.2) |  |
| Cu | 0.3 | 58.2 | 63.3±1.5 | 68 | 73.6±2.3 |
|  |  | (46.6-69.9) |  | (62-74) |  |
| Zn | 30 | 195 | 199±1 | 1179 | 1339±82 |
|  |  | (156-234) |  | (984-1374) |  |
| As | 0.1 | 17.0 | 17.0±0.9 | 198 | 191.4±9.6 |
|  |  | (13.6-20.3) |  | (173-223) |  |
| Se | 0.2 | 26.4 | 26.2±0.7 | 66.2 | 73.4±3.6 |
|  |  | (19.8-33.0) |  | (51.9-80.4) |  |
| Mo | 0.04 | 20.2 | 21.5±1.2 | 73.8 | 82.3±5.9 |
|  |  | (16.2-24.3) |  | (68.1-79.5) |  |
| Cd | 0.01 | 2.56 | 2.74±0.1 | 4.7 | 5.41±0.21 |
|  |  | (2.05-3.08) |  | (3.7-5.6) |  |
| Hg | 0.03 | 1.92 | 2.18±0.18 | 41.5 | 40.3±3.0 |
|  |  | (1.15-2.68) |  | (33.2-49.9) |  |
| Pb | 0.06 | 26.4 | 24.5±0.8 | 78.4 | 78.2±3.3 |
|  |  | (21.1-31.6) |  | (70.4-86.4) |  |