MAD=i=1n|pi-yi|n		     	                                          	( S1)

MD=i=1n(pi-yi)n							( S2)

SD=i=1npi-yi2n							( S3)

The actual changes in blood pressure that were observed in the gold standard sphygmomanometer readings are presented as box plots and as a table with means, median, 25th percentile, 75th percentile, and standard deviations.

Figure S1 Normal Subjects – Box plot of BP corresponding to induced BP change stimulus changes.

Figure S2 Prehypertension Subjects – Box plot of BP corresponding to induced BP change stimulus changes.

Figure S3 Stage 1 Hypertension Subjects – Box plot of BP corresponding to induced BP change stimulus changes.

Figure S4 Stage 2 Hypertension Subjects – Box plot of BP corresponding to induced BP change stimulus changes.

Figure S5 Assessment of device performance in tracking BP recovery after an induced change. Illustration depicting which measurements were used for the performance analysis.

Figure S6 Overall assessment of device performance inclusive of ability to track induced change and recovery of BP. Illustration depicting which measurements were used for the performance analysis.

Table S1 Shows the data presented in the box plots.

| BP Type | Clinical Stratification |  | Condition |  | Mean (mmHg) |  | Standard Deviation (mmHg) |  | Median (mmHg) |  | 25th percentile (mmHg) |  | 75th percentile (mmHg) |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Systolic |  | Normal |  | Static |  | 106.38 |  | 9.21 |  | 108.00 |  | 100.00 |  | 113.00 |
|  |  |  |  | Warm Stimulus |  | 105.39 |  | 11.28 |  | 108.00 |  | 100.00 |  | 113.75 |
|  |  |  |  | Cold Stimulus |  | 110.07 |  | 11.84 |  | 110.00 |  | 102.50 |  | 118.00 |
|  |  |  |  | Rest |  | 106.94 |  | 11.37 |  | 108.00 |  | 99.00 |  | 115.25 |
|  |  | Prehypertension |  | Static |  | 127.89 |  | 7.21 |  | 128.00 |  | 122.75 |  | 133.25 |
|  |  |  |  | Warm Stimulus |  | 126.44 |  | 10.11 |  | 125.00 |  | 120.00 |  | 130.38 |
|  |  |  |  | Cold Stimulus |  | 134.86 |  | 12.60 |  | 132.00 |  | 126.00 |  | 140.00 |
|  |  |  |  | Rest |  | 128.21 |  | 10.60 |  | 129.50 |  | 122.00 |  | 136.00 |
|  |  | Stage 1 Hypertension |  | Static |  | 142.61 |  | 10.06 |  | 144.00 |  | 140.00 |  | 150.00 |
|  |  |  |  | Warm Stimulus |  | 138.34 |  | 12.28 |  | 140.00 |  | 133.75 |  | 146.00 |
|  |  |  |  | Cold Stimulus |  | 142.49 |  | 13.09 |  | 143.00 |  | 138.00 |  | 150.00 |
|  |  |  |  | Rest |  | 138.23 |  | 9.87 |  | 138.00 |  | 133.75 |  | 144.13 |
|  |  | Stage 2 Hypertension |  | Static |  | 164.32 |  | 11.52 |  | 163.00 |  | 160.00 |  | 170.00 |
|  |  |  |  | Warm Stimulus |  | 161.97 |  | 15.35 |  | 160.00 |  | 152.00 |  | 171.00 |
|  |  |  |  | Cold Stimulus |  | 167.44 |  | 14.93 |  | 164.00 |  | 158.00 |  | 180.00 |
|  |  |  |  | Rest |  | 161.07 |  | 15.72 |  | 162.50 |  | 150.00 |  | 170.00 |
| Diastolic |  | Normal |  | Static |  | 67.63 |  | 8.59 |  | 70.00 |  | 63.50 |  | 73.00 |
|  |  |  |  | Warm Stimulus |  | 66.64 |  | 9.79 |  | 68.00 |  | 60.00 |  | 73.50 |
|  |  |  |  | Cold Stimulus |  | 69.43 |  | 9.10 |  | 70.00 |  | 63.25 |  | 78.00 |
|  |  |  |  | Rest |  | 68.16 |  | 8.81 |  | 68.00 |  | 62.00 |  | 74.00 |
|  |  | Prehypertension |  | Static |  | 77.39 |  | 8.81 |  | 80.00 |  | 72.00 |  | 84.00 |
|  |  |  |  | Warm Stimulus |  | 77.36 |  | 8.69 |  | 80.00 |  | 73.50 |  | 82.00 |
|  |  |  |  | Cold Stimulus |  | 81.01 |  | 10.45 |  | 80.00 |  | 74.00 |  | 88.00 |
|  |  |  |  | Rest |  | 79.14 |  | 10.45 |  | 80.00 |  | 74.00 |  | 84.00 |
|  |  | Stage 1 Hypertension |  | Static |  | 88.27 |  | 6.60 |  | 90.00 |  | 85.00 |  | 90.00 |
|  |  |  |  | Warm Stimulus |  | 86.72 |  | 8.77 |  | 89.50 |  | 80.38 |  | 90.00 |
|  |  |  |  | Cold Stimulus |  | 89.02 |  | 11.11 |  | 90.00 |  | 80.00 |  | 94.00 |
|  |  |  |  | Rest |  | 87.10 |  | 9.79 |  | 87.00 |  | 80.00 |  | 92.00 |
|  |  | Stage 2 Hypertension |  | Static |  | 99.04 |  | 13.57 |  | 100.00 |  | 90.00 |  | 105.00 |
|  |  |  |  | Warm Stimulus |  | 96.77 |  | 17.66 |  | 91.00 |  | 82.00 |  | 108.00 |
|  |  |  |  | Cold Stimulus |  | 96.44 |  | 14.29 |  | 98.00 |  | 82.88 |  | 108.00 |
|  |  |  |  | Rest |  | 95.60 |  | 13.69 |  | 91.00 |  | 83.25 |  | 109.50 |

Table S2 Test set blood pressure distribution satisfying ISO 81060-2 Clause 5.1.5

| ISO 81060-2 Clause 5.1.5 Blood Pressure Distribution Requirements | Percentages in test data |
| --- | --- |
| At least 5 % of the reference blood pressure readings shall have a systolic blood pressure ≤100 mmHg (13,33 kPa). | 9.03% |
| At least 5 % of the reference blood pressure readings shall have a systolic blood pressure ≥160 mmHg (21,33 kPa). | 10.13% |
| At least 20 % of the reference blood pressure readings shall have a systolic blood pressure ≥140 mmHg (18,66 kPa). | 32.15% |
| At least 5 % of the reference blood pressure readings shall have a diastolic blood pressure ≤60 mmHg (8,0 kPa). | 7.59% |
| At least 5 % of the reference blood pressure readings shall have a diastolic blood pressure ≥100 mmHg (13,33 kPa). | 7.26% |
| At least 20 % of the reference blood pressure readings shall have a diastolic blood pressure ≥85 mmHg (11,33 kPa). | 29.73% |

Table S3 Number of subjects exhibiting each type of change and combinations thereof.

| Change Description | Normal (Number of Subjects) | Prehypertension (Number of Subjects) | Stage 1 Hypertension (Number of Subjects) | Stage 2 Hypertension (Number of Subjects) | Total Subjects in each change type from all clinical stratifications |
| --- | --- | --- | --- | --- | --- |
| Subjects Exhibiting Exactly One Change (across all protocol time points) |  |  |  |  |  |
| SBP Increase ≥ 15 mmHg only | 2 | 4 | 2 | 3 | 11 |
| DBP Increase ≥ 10 mmHg only | 4 | 4 | 8 | 3 | 19 |
| SBP Decrease ≤ -15 mmHg only | 2 | 1 | 2 | 6 | 11 |
| DBP Decrease ≤ -10 mmHg only | 10 | 1 | 3 | 3 | 17 |
| Subjects Exhibiting Exactly Two Changes (across all protocol time points) |  |  |  |  |  |
| SBP Increase ≥ 15 mmHg and DBP Increase ≥ 10 mmHg only | 4 | 1 | 3 | 3 | 11 |
| SBP Increase ≥ 15 mmHg and SBP Decrease ≤ -15 mmHg only | 2 | 7 | 4 | 5 | 18 |
| SBP Increase ≥ 15 mmHg and DBP Decrease ≤ -10 mmHg only | 2 | 0 | 0 | 0 | 2 |
| DBP Increase ≥ 10 mmHg and SBP Decrease ≤ -15 mmHg only | 0 | 0 | 1 | 1 | 2 |
| DBP Increase ≥ 10 mmHg and DBP Decrease ≤ -10 mmHg only | 3 | 0 | 1 | 4 | 8 |
| SBP Decrease ≤ -15 mmHg and DBP Decrease ≤ -10 mmHg only | 1 | 0 | 0 | 4 | 5 |
| Subjects Exhibiting Exactly Three Changes (across all protocol time points) |  |  |  |  |  |
| SBP Increase ≥ 15 mmHg and DBP Increase ≥ 10 mmHg and SBP Decrease ≤ -15 mmHg only | 1 | 2 | 5 | 1 | 9 |
| DBP Increase ≥ 10 mmHg and SBP Decrease ≤ -15 mmHg and DBP Decrease ≤ -10 mmHg only | 3 | 0 | 2 | 1 | 6 |
| SBP Increase ≥ 15 mmHg and DBP Increase ≥ 10 mmHg and DBP Decrease ≤ -10 mmHg only | 1 | 2 | 0 | 2 | 5 |
| SBP Increase ≥ 15 mmHg and SBP Decrease ≤ -15 mmHg and DBP Decrease ≤ -10 mmHg only | 2 | 3 | 1 | 2 | 8 |
| Subjects Exhibiting Exactly Four Changes (across all protocol time points) |  |  |  |  |  |
| SBP Increase ≥ 15 mmHg and DBP Increase ≥ 10 mmHg and SBP Decrease ≤ -15 mmHg and DBP Decrease ≤ -10 mmHg only | 3 | 5 | 5 | 0 | 13 |
| Totals |  |  |  |  |  |
| Total Subjects in each clinical stratification across all change types | 40 | 30 | 37 | 38 | Total Unique Subjects = 145 |

The following tables S4 and S5 present errors for systolic blood pressure and diastolic blood pressure, respectively, for each clinical category and consequent model selected for estimating blood pressure for each of the conditions used to induce blood pressure variation.

Table S4 Errors for systolic blood pressure for each clinical category and consequent model selected for estimating blood pressure for each of the conditions used to induce blood pressure variation.

| Category | Measurement Type | MAD(mmHg) | MD(mmHg) | SD(mmHg) |
| --- | --- | --- | --- | --- |
| Normal | Static | 3.77 | -1.40 | 4.42 |
|  | Variation 1 (Lower BP) | 4.51 | 0.05 | 5.10 |
|  | Variation 2 (Increase BP) | 4.32 | -0.03 | 5.16 |
|  | Rest | 5.33 | -1.06 | 6.36 |
| Prehypertension | Static | 3.59 | 1.26 | 4.27 |
|  | Variation 1 (Lower BP) | 4.36 | -1.02 | 5.03 |
|  | Variation 2 (Increase BP) | 5.67 | 2.75 | 6.59 |
|  | Rest | 5.68 | 0.06 | 6.44 |
| Stage 1 Hypertension | Static | 3.14 | 0.55 | 3.75 |
|  | Variation 1 (Lower BP) | 3.73 | -2.06 | 4.33 |
|  | Variation 2 (Increase BP) | 4.55 | 1.21 | 5.15 |
|  | Rest | 4.05 | -0.67 | 5.05 |
| Stage 2 Hypertension | Static | 5.66 | 0.83 | 6.61 |
|  | Variation 1 (Lower BP) | 6.38 | -1.16 | 7.59 |
|  | Variation 2 (Increase BP) | 4.61 | 3.71 | 5.39 |
|  | Rest | 5.58 | 3.37 | 5.98 |

Table S5 errors for diastolic blood pressure for each clinical category and consequent model selected for estimating blood pressure for each of the conditions used to induce blood pressure variation.

| Category | Measurement Type | MAD(mmHg) | MD(mmHg) | SD(mmHg) |
| --- | --- | --- | --- | --- |
| Normal | Static | 4.53 | -0.78 | 5.61 |
|  | Variation 1 (Lower BP) | 5.14 | -0.30 | 6.21 |
|  | Variation 2 (Increase BP) | 5.35 | 0.63 | 6.59 |
|  | Rest | 5.20 | -0.41 | 6.45 |
| Prehypertension | Static | 3.22 | -0.14 | 4.28 |
|  | Variation 1 (Lower BP) | 4.79 | -1.66 | 6.09 |
|  | Variation 2 (Increase BP) | 4.95 | -0.60 | 6.30 |
|  | Rest | 4.34 | -0.29 | 5.43 |
| Stage 1 Hypertension | Static | 2.94 | 0.63 | 3.96 |
|  | Variation 1 (Lower BP) | 4.09 | -0.47 | 5.30 |
|  | Variation 2 (Increase BP) | 5.22 | 2.02 | 5.91 |
|  | Rest | 4.65 | -0.21 | 6.04 |
| Stage 2 Hypertension | Static | 5.70 | -3.43 | 6.04 |
|  | Variation 1 (Lower BP) | 5.17 | -3.63 | 5.39 |
|  | Variation 2 (Increase BP) | 5.01 | -3.74 | 6.00 |
|  | Rest | 4.57 | -1.84 | 5.52 |

Table S6 The patient-specific parameters used by the systolic pressure model and the performance for the range of these parameters.

| Confounder | Range (All observed measures split into three intervals, low, medium, and high) | Number of Subjects | Number of measures | MAD (mmHg) | MD(mmHg) | SD(mmHg) |
| --- | --- | --- | --- | --- | --- | --- |
| Age | 20.00 - 40.00 | 36 | 282 | 5.50 | -1.14 | 6.66 |
|  | 40.00 - 60.00 | 51 | 426 | 5.47 | 0.33 | 6.78 |
|  | 60.00 - 80.00 | 31 | 200 | 5.24 | 0.56 | 6.49 |
| Height | 55.00 - 62.50 | 25 | 184 | 6.20 | -0.32 | 7.42 |
|  | 62.50 - 70.00 | 59 | 472 | 5.15 | 0.25 | 6.37 |
|  | 70.00 - 77.50 | 34 | 252 | 5.39 | -0.51 | 6.78 |
| Weight | 90.00 - 190.00 | 76 | 604 | 5.47 | -0.22 | 6.72 |
|  | 190.00 - 290.00 | 38 | 281 | 5.34 | 0.39 | 6.68 |
|  | 290.00 - 390.00 | 4 | 23 | 5.35 | -1.79 | 6.66 |
| Category | Normal | 42 | 370 | 5.41 | -1.19 | 6.51 |
|  | Prehypertensive | 30 | 219 | 5.47 | 1.30 | 6.58 |
|  | Stage 1 Hypertension | 25 | 198 | 4.72 | -0.74 | 5.95 |
|  | Stage 2 hypertension | 21 | 121 | 6.55 | 1.97 | 7.81 |
| Gender | Male | 67 | 531 | 5.20 | -0.29 | 6.48 |
|  | Female | 51 | 377 | 5.75 | 0.24 | 7.03 |
| Ethnicity | Black | 32 | 271 | 6.08 | -0.07 | 7.42 |
|  | Asian | 31 | 238 | 5.12 | -0.85 | 6.20 |
|  | Hispanic | 22 | 144 | 5.22 | 0.98 | 6.54 |
|  | White | 33 | 255 | 5.15 | 0.05 | 6.43 |
| BMI | 10.00 - 25.00 | 35 | 244 | 5.74 | -0.83 | 6.96 |
|  | 25.00 - 35.00 | 69 | 576 | 5.32 | 0.35 | 6.59 |
|  | 35.00 - 55.00 | 14 | 88 | 5.29 | -0.71 | 6.69 |

Table S7 The patient-specific parameters used by the diastolic pressure model and the performance for the range of these parameters.

| Confounder | Range (All observed measures split into three intervals, low, medium, and high) | Number of Subjects | Number of measures | MAD(mmHg) | MD(mmHg) | SD(mmHg) |
| --- | --- | --- | --- | --- | --- | --- |
| Age | 20.00 - 40.00 | 36 | 282 | 5.26 | -1.35 | 6.37 |
|  | 40.00 - 60.00 | 51 | 426 | 4.12 | 0.66 | 5.13 |
|  | 60.00 - 80.00 | 31 | 200 | 5.03 | -2.18 | 6.13 |
| Height | 55.00 - 62.00 | 11 | 86 | 4.95 | -0.06 | 5.87 |
|  | 62.00 - 69.00 | 64 | 493 | 4.79 | -1.11 | 6.02 |
|  | 69.00 - 76.00 | 43 | 329 | 4.44 | 0.06 | 5.62 |
| Weight | 90.00 - 190.00 | 76 | 604 | 4.83 | -0.71 | 6.04 |
|  | 190.00 - 290.00 | 38 | 281 | 4.40 | -0.37 | 5.60 |
|  | 290.00 - 390.00 | 4 | 23 | 3.85 | 0.23 | 5.39 |
| Category | Normal | 42 | 370 | 5.03 | -0.24 | 6.19 |
|  | Prehypertensive | 30 | 219 | 4.27 | -0.67 | 5.52 |
|  | Stage 1 Hypertension | 25 | 198 | 4.16 | 0.45 | 5.36 |
|  | Stage 2 hypertension | 21 | 121 | 5.16 | -3.17 | 5.72 |
| Gender | Male | 67 | 531 | 4.76 | -0.52 | 6.00 |
|  | Female | 51 | 377 | 4.56 | -0.68 | 5.74 |
| Ethnicity | Black | 32 | 271 | 4.22 | -0.51 | 5.29 |
|  | Asian | 31 | 238 | 4.82 | -0.77 | 6.03 |
|  | Hispanic | 22 | 144 | 4.91 | 0.57 | 6.11 |
|  | White | 33 | 255 | 4.91 | -1.15 | 6.17 |
| BMI | 10.00 - 25.00 | 35 | 244 | 5.19 | -2.31 | 6.07 |
|  | 25.00 - 35.00 | 69 | 576 | 4.55 | 0.20 | 5.75 |
|  | 35.00 - 55.00 | 14 | 88 | 4.10 | -0.99 | 5.30 |

### Induced Blood Pressure Changes – Device Accuracy

### Clinical Stratification Transitions

Table S8 Transition counts for subjects with Normal Stratification based on static reading of BP (Total 14 subjects)

| Number of Subjects | Transition type (from before change point to change point) |
| --- | --- |
| 5 | From Normal to Prehypertension |
| 1 | From Prehypertension to Stage 1 Hypertension |
| 8 | No Transition |

Table S9 Transition counts for subjects with Prehypertension Stratification based on static reading of BP (Total 16 subjects)

| Number of Subjects | Transition type (from before change point to change point) |
| --- | --- |
| 2 | From Normal to Prehypertension |
| 1 | From Normal to Stage 1 Hypertension |
| 7 | From Prehypertension to Stage 1 Hypertension |
| 1 | From Prehypertension to Stage 2 Hypertension |
| 2 | From Stage 1 Hypertension to Stage 2 Hypertension |
| 3 | No Transition |

Table S10 Transition counts for subjects with Stage 1 Hypertension Stratification based on static reading of BP (Total 14 subjects)

| Number of Subjects | Transition type (from before change point to change point) |
| --- | --- |
| 1 | From Normal to Stage 1 Hypertension |
| 7 | From Prehypertension to Stage 1 Hypertension |
| 1 | From Prehypertension to Stage 2 Hypertension |
| 3 | From Stage 1 Hypertension to Stage 2 Hypertension |
| 2 | No Transition |

Table S11 Transition counts for subjects with Stage 2 Stratification based on static reading of BP (Total 10 subjects)

| Number of Subjects | Transition type (from before change point to change point) |
| --- | --- |
| 1 | From Stage 1 Hypertension to Stage 2 Hypertension |
| 9 | No Transition |

Table S12 Error Performance with and without averaging of three consecutive readings for assessment of BP accuracy around a timepoint of induced BP change.

| Category | BP type | Number of Subjects | Not Averaged | Averaged |
| --- | --- | --- | --- | --- |
|  |  |  | MAD(mmHg) | MD(mmHg) | SD(mmHg) | MAD(mmHg) | MD(mmHg) | SD(mmHg) |
| Overall | Systolic | 54 | 4.68 | -1.12 | 5.27 | 3.84 | -1.12 | 4.42 |
|  | Diastolic |  | 3.59 | -1.01 | 4.07 | 2.76 | -1.01 | 3.21 |
| Normal | Systolic | 14 | 4.54 | 2.7 | 4.68 | 3.66 | 2.77 | 3.74 |
|  | Diastolic |  | 3.40 | 1.4 | 3.63 | 2.50 | 1.44 | 2.73 |
| Prehypertension | Systolic | 16 | 5.02 | 2.08 | 5.14 | 4.12 | 2.08 | 4.22 |
|  | Diastolic |  | 3.53 | 1.17 | 4.08 | 2.69 | 1.17 | 3.38 |
| Stage 1 Hypertension | Systolic | 14 | 4.58 | 1.67 | 5.02 | 3.70 | 1.58 | 4.05 |
|  | Diastolic |  | 3.76 | 1.16 | 4.31 | 2.95 | 1.1 | 3.44 |
| Stage 2 Hypertension | Systolic | 10 | 4.43 | -3.71 | 3.8 | 3.86 | -3.71 | 2.87 |
|  | Diastolic |  | 3.76 | -0.05 | 4.27 | 2.98 | -0.05 | 3.25 |

Figure S7 Bland Altman Plots with bias, limits of agreement, and respective 95% CI for SBP and DBP in the sub analysis 1 Part a (without averaging of three BP values)

Table S13 Performance on protocol time points with significant change only.

|  | Only protocol time points with significant changes included |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | Description of Change | Clinical Stratification | Number of Subjects | Number of Measurements | MD(mmHg) | SD(mmHg) |
| Systolic | SBP Increase ≥ 15 mmHg | Normal | 17 | 51 | -3.96 | 3.69 |
|  |  | Prehypertension | 24 | 77 | -4.74 | 2.87 |
|  |  | Stage 1 Hypertension | 20 | 62 | -4.91 | 3.53 |
|  |  | Stage 2 Hypertension | 16 | 51 | -4.61 | 4.44 |
|  |  | Overall | 77 | 241 | -4.59 | 3.58 |
|  | SBP Decrease ≤ -15 mmHg | Normal | 14 | 47 | 4.72 | 3.91 |
|  |  | Prehypertension | 18 | 57 | 4.66 | 2.86 |
|  |  | Stage 1 Hypertension | 20 | 80 | 4.48 | 3.22 |
|  |  | Stage 2 Hypertension | 20 | 63 | 2.95 | 4.86 |
|  |  | Overall | 72 | 247 | 4.18 | 3.81 |
| Diastolic | DBP Increase ≥ 10 mmHg | Normal | 19 | 66 | -2.43 | 3.55 |
|  |  | Prehypertension | 14 | 50 | -2.57 | 3.93 |
|  |  | Stage 1 Hypertension | 25 | 79 | -2.55 | 3.57 |
|  |  | Stage 2 Hypertension | 15 | 54 | -2.56 | 4.18 |
|  |  | Overall | 73 | 249 | -2.52 | 3.75 |
|  | DBP Decrease ≤ -10 mmHg | Normal | 25 | 83 | 3.51 | 3.02 |
|  |  | Prehypertension | 11 | 39 | 3.55 | 2.35 |
|  |  | Stage 1 Hypertension | 12 | 42 | 3.61 | 2.83 |
|  |  | Stage 2 Hypertension | 16 | 55 | 2.77 | 4.21 |
|  |  | Overall | 64 | 219 | 3.35 | 3.23 |

Figure S8 Bland Altman Plots with bias, limits of agreement, and respective 95% CI for SBP and DBP in the sub analysis 1 Part b (without averaging of three BP values)

Table S14 Error Performance with and without averaging of three consecutive readings for assessment of BP accuracy overall including points of significant induced BP changes.

| Category | BP type | Number of Subjects | Not Averaged |  |  | Averaged |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | MAD(mmHg) | MD(mmHg) | SD(mmHg) | MAD(mmHg) | MD(mmHg) | SD(mmHg) |
| Overall | Systolic | 54 | 4.55 | -1.73 | 4.93 | 3.94 | -1.84 | 4.13 |
|  | Diastolic |  | 3.62 | -1.5 | 3.93 | 3.08 | -1.79 | 3.2 |
| Normal | Systolic | 14 | 4.68 | 3.19 | 4.3 | 4.22 | 3.39 | 3.34 |
|  | Diastolic |  | 3.84 | 2.12 | 3.76 | 3.08 | 2.27 | 2.77 |
| Prehypertension | Systolic | 16 | 4.89 | 2.55 | 4.82 | 4.26 | 2.81 | 3.94 |
|  | Diastolic |  | 3.44 | 1.36 | 3.94 | 3.09 | 1.93 | 3.34 |
| Stage 1 Hypertension | Systolic | 14 | 4.19 | 1.96 | 4.46 | 3.43 | 1.95 | 3.36 |
|  | Diastolic |  | 3.49 | 1.35 | 3.86 | 3.03 | 1.69 | 3.24 |
| Stage 2 Hypertension | Systolic | 10 | 4.39 | -2.64 | 4.55 | 3.71 | -2.57 | 3.75 |
|  | Diastolic |  | 3.82 | 1.03 | 4.22 | 3.11 | 0.92 | 3.44 |

Figure S9 Bland Altman Plots with bias, limits of agreement, and respective 95% CI for SBP and DBP in the sub analysis 2 Part a (without averaging of three BP values)

Table S15 Performance on protocol time points with all protocol time points included.

|  |  | All Protocol time points (Overall) included |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Measurement Type | Clinical Stratification | Number of Subjects | Number of Measurements | MD(mmHg) | SD(mmHg) |
| Systolic | Normal | 41 | 479 | 1.42 | 4.81 |
|  | Prehypertension | 30 | 353 | -0.76 | 4.92 |
|  | Stage 1 Hypertension | 37 | 434 | 0.42 | 4.53 |
|  | Stage 2 Hypertension | 39 | 441 | -0.99 | 5.08 |
|  | Overall | 147 | 1707 | 0.09 | 4.93 |
| Diastolic | Normal | 41 | 479 | 0.97 | 4.06 |
|  | Prehypertension | 30 | 353 | -0.09 | 3.82 |
|  | Stage 1 Hypertension | 37 | 434 | -0.18 | 3.75 |
|  | Stage 2 Hypertension | 39 | 441 | 0.54 | 4.22 |
|  | Overall | 147 | 1707 | 0.35 | 4 |

Figure S10 Bland Altman Plots with bias, limits of agreement, and respective 95% CI for SBP and DBP in the sub analysis 2 Part b (without averaging of three BP values) with all protocol time points included.

Table S16 Performance on protocol time points with protocol time points with nominal changes only included.

|  |  | Only protocol time points with nominal changes |  |  |  |
| --- | --- | --- | --- | --- | --- |
|  |  | (SBP Change ≤ ±15 mmHg; DBP Change ≤ ±10 mmHg) |  |  |  |
| Measurement Type | Clinical Stratification | Number of Subjects | Number of Measurements | MD(mmHg) | SD(mmHg) |
| Systolic | Normal | 41 | 289 | 1.55 | 4.72 |
|  | Prehypertension | 30 | 180 | -0.23 | 4.35 |
|  | Stage 1 Hypertension | 37 | 232 | 0.21 | 4.22 |
|  | Stage 2 Hypertension | 39 | 266 | -1.14 | 5 |
|  | Overall | 147 | 967 | 0.16 | 4.73 |
| Diastolic | Normal | 41 | 289 | 0.93 | 4.05 |
|  | Prehypertension | 30 | 180 | 0 | 3.58 |
|  | Stage 1 Hypertension | 37 | 232 | 0.14 | 3.6 |
|  | Stage 2 Hypertension | 39 | 266 | 0.64 | 4.09 |
|  | Overall | 147 | 967 | 0.49 | 3.88 |

Figure S11 Bland Altman Plots with bias, limits of agreement, and respective 95% CI for SBP and DBP in the sub analysis 2 Part b (without averaging of three BP values) with only nominal or no change protocol time points included.

Table S17 Error performance measures for the validity of the calibration

| Category | BP type | Number of subjects | MAD(mmHg) | MD(mmHg) | SD(mmHg) |
| --- | --- | --- | --- | --- | --- |
| Overall | Systolic | 91 | 4.37 | -1.81 | 5.02 |
|  | Diastolic |  | 3.26 | -0.3 | 4.03 |
| Normal | Systolic | 24 | 2.94 | 0.55 | 3.77 |
|  | Diastolic |  | 2.86 | 0.64 | 3.5 |
| Prehypertension | Systolic | 23 | 3.66 | -1.09 | 4.21 |
|  | Diastolic |  | 2.98 | 1.25 | 3.54 |
| Stage 1 Hypertension | Systolic | 23 | 4.65 | -3.14 | 4.37 |
|  | Diastolic |  | 3.06 | -2.35 | 2.99 |
| Stage 2 Hypertension | Systolic | 21 | 6.47 | -3.82 | 6.37 |
|  | Diastolic |  | 4.24 | -0.81 | 4.99 |