Details of  HMs and PAH concentration determination and quality control of analyzes

The concentrations of Cd, Pb and Zn were measured, following digestion of the sample in a mixture (1:3 v/v) of concentrated acids (HClO4 and HNO3) in a closed microwave oven (Multiwave 3000, Anton Paar), and passing it through a 100-mesh sieve. The digestion was carried out at 1400 W and 240oC, with the time to reaching maximum power being 5 min, and the time of maintaining maximum power being 15 min. The concentrations of the elements were measured using an atomic-emission spectrometer (PerkinElmer ICP-OES Optima 7300 DV) and a multi-element standard solution (ICP-IV, Merck). The quality of the data was verified using the certified reference material (CRM) 023–050 – Trace Metals – Sandy Loam 7 (RT Corporation). The precision and accuracy of the method for determining HMs in the soils  were controlled as follows:

Precision = standard deviationmean element content x 100%                                            [1]

Accuracy = measured content-reference contentreference content x 100%                           [2]

Table 1S. Parameters of the validation method of HMs determination

| Element | Content |  | Precision | Accuracy | LOD2 |
| --- | --- | --- | --- | --- | --- |
|  | Measured | CRM1 |  |  |  |
|  | mg kg-1 |  | % |  | µg dm-1 |
| Zn | 22.1 ± 0.2 | 20.6 ± 2.2 | 1.31 | 7.00 | 0.0071 |
| Pb | 7.15 ± 0.07 | 7.10 ± 1.10 | 3.22 | 0.70 | 0.0482 |
| Cd | 0.16 ± 0.02 | 0.14 ± 0.06 | 6.78 | 14.3 | 0.0038 |

1certified reference material, 2limit of detection

2.2.3. PAH determination

Sample preparation for PAH analysis

The PAH fraction was extracted from the soil samples according to Application Note 133 of Marcherey-Nagel GmbH Co. KG’s Sample Preparation sheet (1998). An aliquot of 30 ml of 2-propanol was added to 5 g of air-dried soil, ground to pass through a 100-mesh sieve. The sample was homogenised for 20 min in an ultrasonic bath, then shaken for 20 min before being filtered into a 250-ml flask already filled with redistilled water. Each sample was passed through conditioned Chromabond C18ec columns, then 1 ml of the mixture to 1 ml of redistilled water/2-propanol (85:15, v/v) was passed twice through an SPE-column bed to wash out impurities from the bed. The SPE column bed was dried using a stream of N to remove the excess water. The PAHs were eluted using small amounts of dichloromethane until 2 x 500 μl of eluate was collected. The separated PAH fraction was analysed by gas chromatography. The soil samples were analysed in three replications. Where variation in the results exceeded 10%, the analysis was repeated. The CRM assays were performed in five replicates.

PAH detection: Gas chromatography/mass spectrometry

Determination of the PAHs was performed on a Varian GC/MS/MS 4000 equipped with a capillary column, using FactorFour VF-17MS, with a length of 20 m, internal diameter of 0.15 mm and film thickness 0.05 μm. The analysis was performed at the initial temperature of the column (40ºC) for 2 min, then the temperature was increased to 250ºC at a rate of 30ºC/min, and then again from 250 to 320ºC at a rate of 10ºC/min, with the final temperature of 320ºC being maintained for 0.5 min. The total analysis time was 16.5 min. The dispenser had a temperature of 320ºC, the ion source 150ºC, the ion trap 150ºC and the transfer line 325ºC. The carrier gas (He) flowed through the column at a rate of 1 ml/min. The mass spectrometer detector was used in the SCAN mode, monitoring mass in the range 70–300 m/z. The PAH content was measured using a calibration curve prepared from the external standard solution Restek, 610 PAH Calibration Mix A (containing 16 PAHs in methylene chloride at 500–1000 μg/ml). The detection limit (LOD) for the measured PAHs was 2±1.5 µg kg-1 d.m. The quantification limit (LOQ) was 6±4 µg kg-1 d.m. The accuracy of the method was controlled using the CRM for PAHs in Loamy clay 1 – CRM 141 – 50 g (ISO/EC 17025 Cert# AT-1467). The recovery rate of the 16 measured PAHs using this method ranged from 77% (for Ant) to 97.5% (for Naf), and this was used to correct the raw data.