Supporting Information

Table S1. Target respiratory viruses and controls in different wells and detection channels

| Well group |  | Control | Sample wells (2–10) |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Well number |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 |
| Detection channel | FAM | Negative control | PIV 1 | PIV 4 | AdV | INF A H1N1 pdm09 | INF A | RSV | HCoV 229E, NL63 | SARS-CoV-2 (RdRP) | SARS-CoV-2 (N) |
|  | ROX | Internal control (Check chip) | PIV 2 | PIV 3 | HBoV | INF A H3N2 | INF B | HMPV | HCoV OC43 | HRV/EV | Internal control (Check sample) |

FAM: Carboxyfluorescein; ROX: 6-Carboxy-X-Rhodamine; INF A: Influenza A virus; INF B: Influenza B virus; SARS-CoV-2: Severe acute respiratory syndrome coronavirus 2; HCoV: Human coronavirus; HMPV: Human metapneumovirus; HBoV: Human bocavirus; AdV: Adenovirus; HRV/EV: Human rhinovirus and enterovirus; RSV: Respiratory syncytial virus; PIV: Parainfluenza virus

Table S2. Diagnostic criteria for a respiratory virus panel based on a cut-off Ct value

| Diagnostic criteria | Internal control | SARS-CoV-2 (RdRp, N) | HCoV (229E, NL63) | HCoV (OC43) | INF A | INF B | PIV | HMPV | HBoV | RSV | AdV | HRV/ EV |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Positive (+) Ct value | ≤35 | ≤33 | ≤32 | ≤32 | ≤32 | ≤32 | ≤32 | ≤32 | ≤32 | ≤32 | ≤32 | ≤32 |
| Negative (-) Ct value | Not detected | >33 or Not detected | >32 or Not detected | >32 or Not detected | >32or Not detected | >32or Not detected | >33 or Not detected | >32 or Not detected | >32 or Not detected | >32 or Not detected | >32 or Not detected | >32 or Not detected |

INF A: Influenza A virus; INF B: Influenza B virus; SARS-CoV-2: Severe acute respiratory syndrome coronavirus 2; HCoV: Human coronavirus; HMPV: Human metapneumovirus; HBoV: Human bocavirus; AdV: Adenovirus; HRV/EV: Human rhinovirus and enterovirus; RSV: Respiratory syncytial virus; PIV: Parainfluenza virus

Table S3. Comparison of amplification performance between solution-phase and pre-loaded reagent formats using identical PCR chemistry and primer/probe sets.

| Target | Solution-phase Ct | Pre-loaded Ct | ΔCt (Solution − Pre-loaded) |
| --- | --- | --- | --- |
| SARS-CoV-2 (RdRP) | 35.55 | 35.11 | 0.44 |
| SARS-CoV-2 (N) | 35.14 | 32.99 | 2.15 |
| Influenza A | 32.14 | 31.57 | 0.57 |
| Influenza A H1N1pdm09 | 32.96 | 32.17 | 0.79 |
| Influenza A H3N2 | 33.03 | 32.81 | 0.22 |
| Influenza B | 41.89 | 41.30 | 0.59 |
| Respiratory syncytial virus | 34.92 | 35.06 | -0.14 |
| Human coronavirus 229E, NL63 | 36.83 | 34.26 | 2.57 |
| Human coronavirus OC43 | 33.07 | 29.97 | 3.10 |
| Human metapneumovirus | 36.89 | 34.05 | 2.84 |
| Adenovirus | 37.83 | 35.84 | 1.99 |
| Parainfluenza Virus 1 | 32.54 | 31.75 | 0.79 |
| Parainfluenza Virus 2 | 36.25 | 35.76 | 0.49 |
| Parainfluenza Virus 3 | 35.34 | 33.42 | 1.92 |
| Parainfluenza Virus 4 | 35.02 | 35.21 | -0.19 |
| Human bocavirus | 37.12 | 34.91 | 2.21 |
| Human rhinovirus/Enterovirus | 36.94 | 34.15 | 2.79 |

Table S4. Interference test for the Genoplex chip-based PCR platform

| Target | Interference materials [Ct (Hit rate)] |  |  |
| --- | --- | --- | --- |
|  | Zanamivir [5 mg/mL] | hgDNA [5 ug/mL] | Mucin [0.4 mg/mL] |
| SARS-CoV-2(RdRp) | 25.35 (100%) | 26.41 (100%) | 25.17 (100%) |
| SARS-CoV-2 (N) | 27.30 (100%) | 27.44 (100%) | 28.14 (100%) |
| Human coronavirus 229E, NL63 | 29.08 (100%) | 30.16 (100%) | 30.48 (100%) |
| Human coronavirus OC43 | 27.31 (100%) | 27.45 (100%) | 27.16 (100%) |
| Influenza A | 27.32 (100%) | 27.69 (100%) | 27.02 (100%) |
| Influenza A H1N1pdm09 | 27.93 (100%) | 27.36 (100%) | 28.34 (100%) |
| Influenza A H3N2 | 28.01 (100%) | 28.15 (100%) | 30.07 (100%) |
| Influenza B | 32.11 (100%) | 32.35 (100%) | 32.38 (100%) |
| Respiratory syncytial virus | 27.25 (100%) | 27.34 (100%) | 27.14 (100%) |
| Human metapneumovirus | 25.87 (100%) | 25.89 (100%) | 26.78 (100%) |
| Parainfluenza virus type 1 | 28.86 (100%) | 32.34 (100%) | 28.15 (100%) |
| Parainfluenza virus type 2 | 28.09 (100%) | 28.21 (100%) | 28.17 (100%) |
| Parainfluenza virus type 4 | 27.12 (100%) | 27.7 (100%) | 28.01 (100%) |
| Parainfluenza virus type 3 | 27.83 (100%) | 28.03 (100%) | 28.63 (100%) |
| Adenovirus | 28.54 (100%) | 28.03 (100%) | 29.48 (100%) |
| Human bocavirus | 28.42 (100%) | 28.13 (100%) | 27.62 (100%) |
| Human rhinovirus and enterovirus | 27.08 (100%) | 27.37 (100%) | 29.21 (100%) |

hgDNA: Human genomic DNA

Table S5. Mean Ct values and corresponding coefficients of variation (CV, %) across high, medium, and low concentration levels for each viral target.

| Target | High |  | Medium |  | Low |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | Mean Ct | CV (%) | Mean Ct | CV (%) | Mean Ct | CV (%) |
| SARS-CoV-2 (RdRP) | 27.50 | 1.62 | 28.73 | 0.75 | 30.44 | 0.70 |
| SARS-CoV-2 (N) | 27.16 | 0.31 | 28.25 | 0.42 | 29.58 | 1.37 |
| Influenza A | 29.40 | 0.62 | 30.34 | 0.76 | 31.50 | 0.63 |
| Influenza A H1N1pdm09 | 27.36 | 0.99 | 28.32 | 0.25 | 29.67 | 1.98 |
| Influenza A H3N2 | 27.55 | 1.32 | 28.78 | 0.32 | 29.48 | 1.37 |
| Influenza B | 29.28 | 0.25 | 30.68 | 1.18 | 32.39 | 0.50 |
| Respiratory syncytial virus | 27.81 | 1.40 | 29.36 | 3.49 | 31.30 | 2.94 |
| Human coronavirus 229E, NL63 | 30.55 | 0.51 | 31.61 | 0.89 | 32.84 | 0.84 |
| Human coronavirus OC43 | 26.83 | 1.48 | 28.90 | 1.51 | 29.75 | 1.43 |
| Human metapneumovirus | 24.28 | 1.01 | 26.38 | 1.15 | 27.07 | 0.24 |
| Adenovirus | 31.06 | 0.96 | 32.34 | 1.27 | 33.48 | 0.70 |
| Parainfluenza virus 1 | 28.71 | 0.85 | 29.72 | 1.26 | 31.00 | 0.55 |
| Parainfluenza virus 2 | 24.77 | 0.91 | 25.49 | 1.81 | 27.16 | 2.89 |
| Parainfluenza virus 3 | 30.39 | 0.72 | 31.57 | 1.43 | 32.46 | 1.06 |
| Parainfluenza virus 4 | 29.19 | 0.27 | 30.46 | 1.01 | 31.49 | 0.62 |
| Human bocavirus | 27.51 | 1.38 | 28.25 | 0.71 | 29.41 | 1.21 |
| Human rhinovirus/enterovirus | 29.39 | 0.95 | 30.28 | 0.60 | 31.18 | 0.31 |

Table S6. Details of the discrepancy between results obtained when using the Genoplex chip-based PCR platform and comparator kits

| Specimen No. | Genoplex chip-based PCR platform (Ct value of target gene) | Routine clinical laboratory test (Ct value of target gene) |
| --- | --- | --- |
| Specimens with discrepant results for false-positive (n = 5) |  |  |
| 337 | HCoV (16.91) | Negative |
| 47 | HRV/EV (32.27) | Negative |
| 69 | PIV (19.91) | Negative |
| 677 | PIV (29.61) | Negative |
| 666 | SARS-CoV-2 (29.16) | Negative |
| Specimens with discrepant results for false-negative (n = 22) |  |  |
| 142 | Negative | AdV (17.24) |
| 254 | Negative | AdV (16.11) |
| 255 | Negative | AdV (20.72) |
| 361 | Negative | AdV (20.91) |
| 366 | Negative | AdV (18.03) |
| 378 | Negative | AdV (21.21) |
| 404 | Negative | AdV (19.03) |
| 575 | Negative | AdV (18.89) |
| 615 | Negative | AdV (20.78) |
| 616 | Negative | AdV (22.39) |
| 644 | Negative | AdV (7.52) |
| 205 | Negative | HBoV (19.17) |
| 471 | Negative | HCoV (20.46) |
| 565 | Negative | HCoV (20.80) |
| 66 | Negative | HRV/EV (21.72/18.50) |
| 241 | Negative | HRV/EV (13.38/19.71) |
| 281 | Negative | HRV/EV (18.83/16.55) |
| 452 | Negative | HRV/EV (21.86/17.79) |
| 606 | Negative | HRV/EV (20.54/13.98) |
| 619 | Negative | HRV/EV (15.54/19.79) |
| 163 | Negative | PIV (23.58) |
| 274 | Negative | SARS-CoV-2 (35.95) |
| Specimens with discrepant results for multiple respiratory virus infections (n = 13) |  |  |
| 539 | INF A (14.33) and HCoV (19.52) | INF A (6.69) |
| 70 | INF A (17.79) and AdV (5.61) | INF A (11.67) |
| 11 | INF A (23.33) and HBoV (24.23) | INF A (15.83) |
| 72 | INF A (19.62) and HBoV (21.05) | INF A (10.82) |
| 54 | HBoV (17.99) and HCoV (15.52) | HBoV (10.04) |
| 525 | HBoV (12.37) and HCoV (18.16) | HBoV (6.16) |
| 150 | HMPV (20.20) and SARS-CoV-2 (22.21) | HMPV (12.10) |
| 183 | HRV/EV (15.07) and H3N2 (18.56) | HRV/EV (8.57/20.89) |
| 425 | PIV (18.12) and HCoV (20.92) and HRV/EV (24.17) | PIV (16.27) |
| 224 | RSV (15.13) and INF B (29.42) | RSV (11.77) |
| 259 | RSV (16.05) and INF A (20.85) | RSV (17.70) |
| 313 | RSV (14.28) and INF A (31.19) | RSV (14.45) |
| 83 | SARS-CoV-2 (14.76) and HRV/EV (30.70) | SARS-CoV-2 (21.77) |
| Specimens with discrepant results for different viral infections (n = 1) |  |  |
| 610 | INF A (21.39) | HRV/EV (21.86/17.11) |
| Total number of specimens with discrepant results (n = 41) |  |  |

Note: Both the Genoplex assay and the AdvanSure assay employ two-phase amplification protocols; however, the number of initial pre-amplification cycles differs between assays (5 cycles for the Genoplex assay and 10 cycles for the AdvanSure assay). Accordingly, reported Ct values are assay-dependent and should be interpreted with caution.

For SARS-CoV-2–positive specimens, when both RdRp and N gene targets were amplified, the Ct value of the RdRp gene was reported as the representative value. In specimens No. 83 and No. 666, only the N gene was amplified. In specimen No. 183, the influenza A virus target was not detected, whereas only the influenza A(H3N2) subtyping signal was detected.

SARS-CoV-2: Severe acute respiratory syndrome coronavirus 2; INF A: Influenza A virus; INF B: Influenza B virus; RSV: Respiratory syncytial virus; HCoV: Human coronavirus; HMPV: Human metapneumovirus; AdV: Adenovirus; PIV 1–4: Parainfluenza virus types 1–4; HBoV: Human bocavirus; HRV/EV: Human rhinovirus and enterovirus.

Table S7. Comparison of viral target composition

| Ability to detect pathogen viruses | Genoplex | FilmArray | Verigene | NxTAG-RPP | Allplex |
| --- | --- | --- | --- | --- | --- |
| Severe acute respiratory syndrome coronavirus 2 | ✓ | ✓ |  | ✓ |  |
| Adenovirus | ✓ | ✓ | ✓ | ✓ | ✓ |
| Coronavirus HKU1 |  | ✓ |  | ✓ |  |
| Coronavirus NL63 | ✓ | ✓ |  | ✓ | ✓ |
| Coronavirus 229E | ✓ | ✓ |  | ✓ | ✓ |
| Coronavirus OC43 | ✓ | ✓ |  | ✓ | ✓ |
| Human bocavirus | ✓ |  |  |  | ✓ |
| Human metapneumovirus | ✓ | ✓ | ✓ | ✓ | ✓ |
| Influenza A virus | ✓ | ✓ | ✓ | ✓ | ✓ |
| Subtype H1 |  | ✓ | ✓ | ✓ | ✓ |
| Subtype H3 | ✓ | ✓ | ✓ | ✓ | ✓ |
| Subtype 2009 H1N1 | ✓ | ✓ |  | ✓ | ✓ |
| Influenza B virus | ✓ | ✓ | ✓ | ✓ | ✓ |
| Parainfluenza virus 1 | ✓ | ✓ | ✓ | ✓ | ✓ |
| Parainfluenza virus 2 | ✓ | ✓ | ✓ | ✓ | ✓ |
| Parainfluenza virus 3 | ✓ | ✓ | ✓ | ✓ | ✓ |
| Parainfluenza virus 4 | ✓ | ✓ | ✓ | ✓ | ✓ |
| Respiratory syncytial virus | ✓ | ✓ |  |  |  |
| Respiratory syncytial virus A |  |  | ✓ | ✓ | ✓ |
| Respiratory syncytial virus B |  |  | ✓ | ✓ | ✓ |
| Rhinovirus/enterovirus | ✓ | ✓ | ✓ | ✓ | ✓ |

RPP, respiratory pathogen panel. RPP, respiratory pathogen panel. BioFire FilmArray RP2.1 (bioMérieux, Marcy-l’Étoile, France); Luminex NxTAG RPPv2 (Luminex Corporation, Austin, TX, USA); Luminex Verigene RP Flex (Luminex Corporation, Austin, TX, USA); Seegene Allplex Respiratory Panel (Seegene, Seoul, Korea).