Supplementary Material

Supplementary Table 1. List of origin of samples, sites, and devices used for testing.

| Sample provision | Samples, N total (N PCR-positive) |
| --- | --- |
| Stuttgart | 1331 (95) |
| Hamburg | 1058 (58) |
| Ludwigshafen | 647 (151) |
| Commercial vendor | 103 (88) |
| Testing site | Samples, N total |
| Hamburg (cobas e 411) | 1808 |
| Augsburg (cobas e 801) | 1331 |

Ct, cycle threshold; N, number of samples

Supplementary Table 2. Analytic performance of Elecsys SARS-CoV-2 Antigen stratified by symptoms, duration of symptoms, and according to viral load (copies/mL).

|  | Description | N total | Elecsys SARS-CoV-2 Antigen negative | Elecsys SARS-CoV-2 Antigen positive | Relative sensitivity, % (95%, CI two-sided) |
| --- | --- | --- | --- | --- | --- |
| Symptomatic | All | 313 | 99 | 214 | 68.4 (62.9–73.5) |
|  | Symptomatic <5 days DPSO Symptomatic >5 DPSO Symptomatic with unknown DPSO | 149 109 55 | 25 59 15 | 124 50 40 | 83.2 (76.2–88.8) 45.9 (36.3–55.7) 72.7 (59.0–83.9) |
|  | T2 >104 copies/mL (~Ct 29.9) | 197 | 8 | 189 | 95.9 (92.2–98.2) |
|  | Symptomatic <5 days DPSO Symptomatic >5 DPSO Symptomatic with unknown DPSO | 118 42 37 | 2 5 1 | 116 37 36 | 98.3 (94.0–99.8) 88.1 (74.4–96.0) 97.3 (85.8–99.9) |
|  | T2 >105 copies/mL (~Ct 26.6) | 153 | 0 | 153 | 100.0 (97.6–100.0) |
|  | Symptomatic <5 days DPSO Symptomatic >5 DPSO Symptomatic with unknown DPSO | 99 28 26 | 0 0 0 | 99 28 26 | 100.0 (96.3–100.0) 100.0 (87.7–100.0) 100.0 (86.8–100.0) |
|  | T2 >106 copies/mL (~Ct 23.0) | 109 | 0 | 109 | 100.0 (96.7–100.0) |
|  | Symptomatic <5 days DPSO Symptomatic >5 DPSO Symptomatic with unknown DPSO | 75 20 14 | 0 0 0 | 75 20 14 | 100.0 (95.2–100.0) 100.0 (83.2–100.0) 100.0 (76.8–100.0) |
| Asymptomatic (Known / suspected exposure + screening) | All | 79 | 57 | 22 | 27.8 (18.3–39.1) |
|  | T2 >104 copies/mL (~Ct 29.9) | 26 | 6 | 20 | 76.9 (56.4–91.0) |
|  | T2 >105 copies/mL (~Ct 26.6) | 18 | 0 | 18 | 100.0 (81.5–100.0) |
|  | T2 >106 copies/mL (~Ct 23.0) | 13 | 0 | 13 | 100.0 (75.3–100.0) |

CI, confidence interval; Ct, cycle threshold; DPSO, days post symptom onset; N, number of samples; T2, Target-2