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Supplementary Table 1. List of origin of samples, sites, and devices used for testing.
	Sample provision
	Samples, N total (N PCR-positive)

	Stuttgart
	1331 (95)

	Hamburg
	1058 (58)

	Ludwigshafen
	647 (151)

	Commercial vendor
	103 (88)

	Testing site 
	Samples, N total

	Hamburg (cobas e 411)
	1808

	Augsburg (cobas e 801)
	1331


Ct, cycle threshold; N, number of samples


Supplementary Table 2. Analytic performance of Elecsys SARS-CoV-2 Antigen stratified by symptoms, duration of symptoms, and according to viral load (copies/mL).
	
	Description
	N total
	Elecsys SARS-CoV-2 Antigen negative
	Elecsys SARS-CoV-2 
Antigen positive
	Relative sensitivity, %
(95%, CI two-sided)

	Symptomatic
	All
	313
	99
	214
	68.4 (62.9–73.5)

	
	Symptomatic <5 days DPSO 
Symptomatic >5 DPSO
Symptomatic with unknown DPSO
	149
109
55
	25
59
15
	124
50
40
	83.2 (76.2–88.8)
45.9 (36.3–55.7)
72.7 (59.0–83.9)

	
	T2 >104 copies/mL (~Ct 29.9)
	197
	8
	189
	95.9 (92.2–98.2)

	
	Symptomatic <5 days DPSO 
Symptomatic >5 DPSO
Symptomatic with unknown DPSO
	118
42
37
	2
5
1
	116
37
36
	98.3 (94.0–99.8)
88.1 (74.4–96.0)
97.3 (85.8–99.9)

	
	T2 >105 copies/mL (~Ct 26.6)
	153
	0
	153
	100.0 (97.6–100.0)

	
	Symptomatic <5 days DPSO 
Symptomatic >5 DPSO
Symptomatic with unknown DPSO
	99
28
26
	0
0
0
	99
28
26
	100.0 (96.3–100.0)
100.0 (87.7–100.0)
100.0 (86.8–100.0)

	
	T2 >106 copies/mL (~Ct 23.0)
	109
	0
	109
	100.0 (96.7–100.0)

	
	Symptomatic <5 days DPSO 
Symptomatic >5 DPSO
Symptomatic with unknown DPSO
	75
20
14
	0
0
0
	75
20
14
	100.0 (95.2–100.0)
100.0 (83.2–100.0)
100.0 (76.8–100.0)

	Asymptomatic
(Known / suspected exposure + screening)
	All
	79
	57
	22
	27.8 (18.3–39.1)

	
	T2 >104 copies/mL (~Ct 29.9)
	26
	6
	20
	76.9 (56.4–91.0)

	
	T2 >105 copies/mL (~Ct 26.6)
	18
	0
	18
	100.0 (81.5–100.0)

	
	T2 >106 copies/mL (~Ct 23.0)
	13
	0
	13
	100.0 (75.3–100.0)


CI, confidence interval; Ct, cycle threshold; DPSO, days post symptom onset; N, number of samples; T2, Target-2
