Supporting Information

S1 Table. Background reading and zero-adjusted data for Task 1 in parts per million (ppm) and task duration in minutes. The maximum reading for each sample during the entire task execution is indicated in bold.

| CSTD Device | BG0 | Reading 1 | Reading 2 | Reading 3 | Reading 4 | Reading 5 | Reading 6 | Task Duration |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Background Reading | Attach Vial Spike | Fill Transfer Syringe | Fill Dose Syringe | IV Bag Transfer 1 | Fill 2nd Dose Syringe | IV Bag Transfer 2 |  |
| ChemoLock |  |  |  |  |  |  |  |  |
| Sample 1 | -0.5 | 0.1 | 0.1 | 0.2 | 0.1 | 0.1 | 0.2 | 2.87 |
| Sample 2 | -0.2 | 0.1 | 0.2 | 0.1 | 0.1 | 0.2 | 0.2 | 2.67 |
| Sample 3 | -0.1 | 0.0 | 0.0 | 0.1 | 0.1 | 0.1 | 0.1 | 2.63 |
| Sample 4 | -0.1 | 0.0 | 0.2 | 0.1 | 0.2 | 0.2 | 0.2 | 2.75 |
| Sample 5 | -0.1 | 0.1 | 0.1 | 0.2 | 0.2 | 0.3 | 0.3 | 2.57 |
| Sample 6 | -0.1 | 0.1 | 0.1 | 0.2 | 0.2 | 0.3 | 0.3 | 2.60 |
| PhaSeal |  |  |  |  |  |  |  |  |
| Sample 1 | 0.1 | 0.0 | 0.1 | 0.2 | 0.2 | 0.4 | 0.3 | 4.13 |
| Sample 2 | 0.2 | 0.1 | 0.3 | 0.3 | 0.5 | 0.6 | 0.6 | 4.43 |
| Sample 3 | 0.0 | 0.1 | 0.3 | 0.5 | 0.6 | 0.8 | 0.8 | 4.67 |
| Sample 4 | 0.2 | 0.0 | 0.2 | 0.3 | 0.4 | 0.5 | 0.5 | 4.47 |
| Sample 5 | 0.2 | 0.1 | 0.3 | 0.5 | 0.6 | 0.8 | 0.8 | 4.12 |
| Sample 6 | 0.0 | 0.1 | 0.4 | 0.5 | 0.5 | 0.7 | 0.7 | 4.83 |
| Equashield |  |  |  |  |  |  |  |  |
| Sample 1 | 0.1 | 0.1 | 0.1 | 0.2 | 0.2 | 0.2 | 0.2 | 3.25 |
| Sample 2 | 0.0 | 0.0 | 0.1 | 0.1 | 0.0 | 0.1 | 0.1 | 2.85 |
| Sample 3 | 0.1 | 0.0 | 0.1 | 0.2 | 0.3 | 0.3 | 0.3 | 3.02 |
| Sample 4 | 0.0 | 0.0 | 0.1 | 0.2 | 0.2 | 0.3 | 0.3 | 3.17 |
| Sample 5 | 0.0 | 0.0 | 0.1 | 0.1 | 0.2 | 0.2 | 0.2 | 2.85 |
| Sample 6 | -0.1 | 0.0 | 0.1 | 0.1 | 0.2 | 0.2 | 0.2 | 3.27 |
| Positive Control |  |  |  |  |  |  |  |  |
| Sample 1 | 0.0 | 0.0 | 0.5 | 1.0 | 10.7 | 16.5 | 24.8 | 6.22 |
| Sample 2 | 0.2 | 0.1 | 1.0 | 1.3 | 3.5 | 3.0 | 10.0 | 5.05 |

S2 Table. Background reading and zero-adjusted data for Task 2 in parts per million (ppm) and task duration in minutes. The maximum reading for each sample during the entire task execution is indicated in bold.

| CSTD Device | BG0 | Reading 1 | Reading 2 | Reading 3 | Reading 4 | Reading 5 | Reading 6 | Task Duration |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Background Reading | Attach Vial Spikes | Fill Transfer Syringe | Fill Bolus Syringe 3 | Fill Bolus Syringe 4 | Y-site Push 1 | Y-site Push 2 |  |
| ChemoLock |  |  |  |  |  |  |  |  |
| Sample 1 | -0.1 | 0.0 | 0.1 | 0.2 | 0.2 | 0.2 | 0.1 | 2.85 |
| Sample 2 | -0.1 | 0.2 | 0.3 | 0.5 | 0.5 | 0.6 | 0.6 | 2.73 |
| Sample 3 | 0.1 | 0.1 | 0.3 | 0.4 | 0.4 | 0.4 | 0.5 | 2.32 |
| Sample 4 | 0.1 | 0.1 | 0.2 | 0.3 | 0.3 | 0.4 | 0.5 | 2.48 |
| Sample 5 | 0.1 | 0.1 | 0.3 | 0.5 | 0.6 | 0.7 | 0.7 | 2.75 |
| Sample 6 | 0.1 | 0.1 | 0.3 | 0.5 | 0.5 | 0.6 | 0.6 | 2.72 |
| PhaSeal |  |  |  |  |  |  |  |  |
| Sample 1 | 0.1 | 0.1 | 0.2 | 0.3 | 0.3 | 0.3 | 0.3 | 4.10 |
| Sample 2 | 0.2 | 0.1 | 0.1 | 0.5 | 0.6 | 0.7 | 0.7 | 4.05 |
| Sample 3 | 0.2 | 0.1 | 0.1 | 0.2 | 0.2 | 0.2 | 0.3 | 4.98 |
| Sample 4 | 0.1 | 0.1 | 0.3 | 0.3 | 0.4 | 0.4 | 0.4 | 4.62 |
| Sample 5 | 0.0 | 0.1 | 0.3 | 0.4 | 0.5 | 0.5 | 0.4 | 4.68 |
| Sample 6 | 0.0 | 0.1 | 0.3 | 0.5 | 0.5 | 0.5 | 0.4 | 4.72 |
| Equashield |  |  |  |  |  |  |  |  |
| Sample 1 | 0.0 | 0.1 | 0.2 | 0.2 | 0.2 | 0.3 | 0.3 | 3.42 |
| Sample 2 | -0.2 | 0.1 | 0.2 | 0.2 | 0.3 | 0.3 | 0.5 | 3.25 |
| Sample 3 | 0.2 | 0.1 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 4.40 |
| Sample 4 | -0.1 | 0.1 | 0.2 | 0.3 | 0.3 | 0.3 | 0.2 | 3.22 |
| Sample 5 | -0.1 | 0.0 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 3.60 |
| Sample 6 | 0.0 | 0.1 | 0.2 | 0.3 | 0.3 | 0.4 | 0.4 | 4.00 |
| Positive Control |  |  |  |  |  |  |  |  |
| Sample 1 | 0.2 | 0.2 | 0.8 | 4.1 | 12.7 | 19.8 | 28.3 | 6.38 |
| Sample 2 | 0.1 | 0.5 | 1.4 | 4.9 | 7.4 | 14.1 | 18.6 | 5.77 |