Personalized Cancer Monitoring Assay for the Detection of ctDNA in Patients with Solid Tumors

Jianhua Zhao1, Jacquelyn Reuther1, Kaylee Scozzaro1, Megan Hawley1, Emily Metzger1,  Matthew Emery1, Ingrid Chen1, Michelle Barbosa1, Laura Johnson1*, Alijah O’Connor1, Mike Washburn1*, Luke Hartje1*, Erik Reckase1*, Verity Johnson1*, Yuhua Zhang1, Emily Westheimer1, William O’Callaghan1, Nirav Malani1, Adrian Chesh1, Michael Moreau1, Robert Daber1

1 Invitae Corp., 1400 16th Street, San Francisco, CA 94103

*Affiliated with Invitae Corp. at the time of the study, currently employees at Integrated DNA Technologies, 1710 Commercial Park, Coralville, IA 52241.

Corresponding author: Jianhua Zhao PhD FACMG, jianhua.zhao@invitae.com

Acknowledgement: We thank Aaron Garnett, Paula Roberts, Nicole Williams, Dhruiti Legare, Angela Oliviero, and Laura Hyland for their contributions to the development of the technology.

Supplementary Table 1. False-negative results in accuracy study for whole exome sequencing

| Lab Sample ID | Gene | Chromosome | Position (hg19/GRCh37) | Reference | Alternate | Expected Allele Frequency | c. Nomenclature | p. Nomenclature | Note |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| GEV1800221 | BRCA2 | 13 | 32910721 | T | C | 21.7 | NM_000059.3:c.2229T>C | p.His743His | False negative, below 20% allele frequency |
| GEV1800205 | FLT3 | 13 | 28610183 | A | G | 13.1 | NM_004119.2:c.1310-3T>C | p.? | False negative, below 20% allele frequency |
| GEV1800205 | BRCA2 | 13 | 32911888 | A | G | 14.8 | NM_000059.3:c.3396A>G | p.Lys1132Lys | False negative, below 20% allele frequency |
| GEV1800231 | TERT | 5 | 1295228 | G | A | 42.4 | NM_198253.2:c.-124C>T | p.= | Poor coverage |
| GEV1800231 | CEBPA | 19 | 33793430 | G | A | 34.2 | NM_004364.4:c.-110C>T | p.? | False negative, below 20% allele frequency |
| GEV1800236 | PDGFRB | 5 | 149499672 | T | C | 14.7 | NM_002609.3:c.2601A>G | p.Leu867Leu | Poor coverage |
| GEV1800236 | DDX41 | 5 | 176940384 | G | A | 11.6 | NM_016222.3:c.1200C>T | p.Arg400Arg | False negative, below 20% allele frequency |
| GEV1800236 | TP53 | 17 | 7578210 | T | C | 15.4 | NM_000546.5:c.639A>G | p.Arg213Arg | False negative, below 20% allele frequency |
| GEV1800235 | PDGFRB | 5 | 149499672 | T | C | 30.7 | NM_002609.3:c.2601A>G | p.Leu867Leu | False negative, below 20% allele frequency |
| GEV1800235 | MET | 7 | 116411848 | C | CCT | 46.8 | NM_001127500.2:c.2942-54_2942-53dup | p.? | False negative, below 20% allele frequency |
| GEV1800235 | BLM | 15 | 91310209 | A | G | 22.4 | NM_001287246.1:c.2263A>G | p.Lys755Glu | Poor coverage |
| GEV1800230 | TERT | 5 | 1295250 | G | A | 38.8 | NM_198253.2:c.-146C>T | p.= | Poor coverage |
| GEV1800218 | FLT3 | 13 | 28610183 | A | G | 12.1 | NM_004119.2:c.1310-3T>C | p.? | False negative, below 20% allele frequency |
| GEV1800218 | BRCA2 | 13 | 32907129 | T | C | 11.4 | NM_000059.3:c.1514T>C | p.Ile505Thr | False negative, below 20% allele frequency |
| GEV1800692 | NPM1 | 5 | 170837543 | C | CTCTG | 6.2 | NM_002520.6:c.860_863dup | p.Trp288Cysfs*12 | Poor coverage |

Supplementary Table 2. Limit of detection of whole exome sequencing: variant allele frequency

| Variant Allele Frequency | Number of Expected Variants | Number of Variants Detected | Sensitivity (%) |
| --- | --- | --- | --- |
| 5-7% | 89 | 83 | 93.3 |
| 10% | 91 | 91 | 100 |
| 10-15% | 101 | 101 | 100 |

Supplementary Table 3. Limit of blank of whole exome sequencing for panel design

| A. |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Panel Design Metrics | Passed Design Metric Acceptance Criterion: n/N (%) | Mean | Standard Deviation | Minimum | Maximum |
| Normal SNP ID Called Positions | 120/120 (100.0%) | 30.74 | 4.43 | 20 | 44 |
| Tumor/Germline Variant Concordance | 120/120 (100.0%) | 99.46 | 0.52 | 94.02 | 99.82 |
| Targeted PSP Variant Count | 0/120 (0.0%) | 0.02 | 0.16 | 0 | 1 |

| B. |  |  |  |  |
| --- | --- | --- | --- | --- |
| Expected Panel Design Failure: n/N (%) | Observed Panel Design Failure: n/N (%) | Agreement between Expected & Observed Failure: n/N (%) | 95% CI |  |
| 120/120 (100.0%) | 120/120 (100.0%) | 120/120 (100.0%) | 97-100% |  |

Supplementary Table 4. Summary of expected and observed ctDNA calls by cfDNA Input and allele frequency with baseline threshold in part 1 of accuracy study of ctDNA detection

| Library ID | cfDNA Input | Allele Frequency | Panel ID | Expected ctDNA Result | Observed ctDNA Result | Total Unique On-Target Reads with Depth of >= 5 Reads Covering Specific Variant Positions |
| --- | --- | --- | --- | --- | --- | --- |
| Patient-1-NUC-7288-GEC2007961EP-D1-D1 | 10 | 0.100 | abef6d7a6 | PRESENT | PRESENT | 87252 |
| Patient-1-NUC-7289-GEC2007962EP-D1-D1 | 10 | 0.100 | abef6d7a6 | PRESENT | PRESENT | 101729 |
| Patient-1-NUC-7290-GEC2007963EP-D1-D1 | 10 | 0.100 | abef6d7a6 | PRESENT | PRESENT | 93680 |
| Patient-1-NUC-7291-GEC2007964EP-D1-D1 | 10 | 0.050 | abef6d7a6 | PRESENT | PRESENT | 59881 |
| Patient-1-NUC-7292-GEC2007965EP-D1-D1 | 10 | 0.050 | abef6d7a6 | PRESENT | PRESENT | 87413 |
| Patient-1-NUC-7293-GEC2007966EP-D1-D1 | 10 | 0.050 | abef6d7a6 | PRESENT | PRESENT | 84071 |
| Patient-1-NUC-7294-GEC2007967EP-D1-D1 | 25 | 0.050 | abef6d7a6 | PRESENT | PRESENT | 210666 |
| Patient-1-NUC-7295-GEC2007968EP-D1-D1 | 25 | 0.050 | abef6d7a6 | PRESENT | PRESENT | 190585 |
| Patient-1-NUC-7296-GEC2007993EP-D1-D1 | 25 | 0.050 | abef6d7a6 | PRESENT | PRESENT | 172468 |
| Patient-1-NUC-7297-GEC2007994EP-D1-D1 | 10 | 0.010 | abef6d7a6 | PRESENT | PRESENT | 55224 |
| Patient-1-NUC-7298-GEC2007995EP-D1-D1 | 10 | 0.010 | abef6d7a6 | PRESENT | PRESENT | 72912 |
| Patient-1-NUC-7299-GEC2007996EP-D1-D1 | 10 | 0.010 | abef6d7a6 | PRESENT | PRESENT | 74724 |
| Patient-1-NUC-7300-GEC2007997EP-D1-D1 | 25 | 0.010 | abef6d7a6 | PRESENT | PRESENT | 173265 |
| Patient-1-NUC-7301-GEC2007998EP-D1-D1 | 25 | 0.010 | abef6d7a6 | PRESENT | PRESENT | 184160 |
| Patient-1-NUC-7302-GEC2007999EP-D1-D1 | 25 | 0.010 | abef6d7a6 | PRESENT | PRESENT | 169394 |
| Patient-1-NUC-7303-GEC2008000EP-D1-D1 | 60 | 0.008 | abef6d7a6 | PRESENT | PRESENT | 324742 |
| Patient-1-NUC-7304-GEC2007969EP-D1-D1 | 60 | 0.008 | abef6d7a6 | PRESENT | PRESENT | 438272 |
| Patient-1-NUC-7305-GEC2007970EP-D1-D1 | 10 | 0.008 | abef6d7a6 | PRESENT | PRESENT | 86686 |
| Patient-1-NUC-7306-GEC2007971EP-D1-D1 | 10 | 0.008 | abef6d7a6 | PRESENT | PRESENT | 85022 |
| Patient-1-NUC-7307-GEC2007972EP-D1-D1 | 10 | 0.008 | abef6d7a6 | PRESENT | PRESENT | 86916 |
| Patient-1-NUC-7308-GEC2007973EP-D1-D1 | 25 | 0.008 | abef6d7a6 | PRESENT | PRESENT | 200334 |
| Patient-1-NUC-7309-GEC2007974EP-D1-D1 | 25 | 0.008 | abef6d7a6 | PRESENT | PRESENT | 212083 |
| Patient-1-NUC-7310-GEC2007975EP-D1-D1 | 25 | 0.008 | abef6d7a6 | PRESENT | PRESENT | 194068 |
| Patient-1-NUC-7311-GEC2007976EP-D1-D1 | 60 | 0.005 | abef6d7a6 | PRESENT | PRESENT | 371109 |
| Patient-1-NUC-7312-GEC2008001EP-D1-D1 | 60 | 0.005 | abef6d7a6 | PRESENT | PRESENT | 374675 |
| Patient-2-NUC-7313-GEC2008002EP-D1-D1 | 10 | 0.100 | 17e2d6534 | PRESENT | PRESENT | 93422 |
| Patient-2-NUC-7314-GEC2008003EP-D1-D1 | 10 | 0.100 | 17e2d6534 | PRESENT | PRESENT | 81416 |
| Patient-2-NUC-7315-GEC2008004EP-D1-D1 | 10 | 0.100 | 17e2d6534 | PRESENT | PRESENT | 72760 |
| Patient-2-NUC-7316-GEC2008005EP-D1-D1 | 10 | 0.050 | 17e2d6534 | PRESENT | PRESENT | 68875 |
| Patient-2-NUC-7317-GEC2008006EP-D1-D1 | 10 | 0.050 | 17e2d6534 | PRESENT | PRESENT | 66175 |
| Patient-2-NUC-7318-GEC2008007EP-D1-D1 | 10 | 0.050 | 17e2d6534 | PRESENT | PRESENT | 65213 |
| Patient-2-NUC-7319-GEC2008008EP-D1-D1 | 25 | 0.050 | 17e2d6534 | PRESENT | PRESENT | 138819 |
| Patient-2-NUC-7320-GEC2007977EP-D1-D1 | 25 | 0.050 | 17e2d6534 | PRESENT | PRESENT | 200178 |
| Patient-2-NUC-7321-GEC2007978EP-D1-D1 | 25 | 0.050 | 17e2d6534 | PRESENT | PRESENT | 205121 |
| Patient-2-NUC-7322-GEC2007979EP-D1-D1 | 10 | 0.010 | 17e2d6534 | PRESENT | PRESENT | 88127 |
| Patient-2-NUC-7323-GEC2007980EP-D1-D1 | 10 | 0.010 | 17e2d6534 | PRESENT | PRESENT | 81935 |
| Patient-2-NUC-7324-GEC2007981EP-D1-D1 | 10 | 0.010 | 17e2d6534 | PRESENT | PRESENT | 75867 |
| Patient-2-NUC-7325-GEC2007982EP-D1-D1 | 25 | 0.010 | 17e2d6534 | PRESENT | PRESENT | 202621 |
| Patient-2-NUC-7326-GEC2007983EP-D1-D1 | 25 | 0.010 | 17e2d6534 | PRESENT | PRESENT | 193434 |
| Patient-2-NUC-7327-GEC2007984EP-D1-D1 | 25 | 0.010 | 17e2d6534 | PRESENT | PRESENT | 184097 |
| Patient-2-NUC-7328-GEC2008009EP-D1-D1 | 60 | 0.008 | 17e2d6534 | PRESENT | PRESENT | 403878 |
| Patient-2-NUC-7329-GEC2008010EP-D1-D1 | 60 | 0.008 | 17e2d6534 | PRESENT | PRESENT | 389136 |
| Patient-2-NUC-7330-GEC2008011EP-D1-D1 | 10 | 0.008 | 17e2d6534 | PRESENT | PRESENT | 84396 |
| Patient-2-NUC-7331-GEC2008012EP-D1-D1 | 10 | 0.008 | 17e2d6534 | PRESENT | PRESENT | 86460 |
| Patient-2-NUC-7332-GEC2008013EP-D1-D1 | 10 | 0.008 | 17e2d6534 | PRESENT | PRESENT | 86205 |
| Patient-2-NUC-7333-GEC2008014EP-D1-D1 | 25 | 0.008 | 17e2d6534 | PRESENT | PRESENT | 191719 |
| Patient-2-NUC-7334-GEC2008015EP-D1-D1 | 25 | 0.008 | 17e2d6534 | PRESENT | PRESENT | 201607 |
| Patient-2-NUC-7335-GEC2008016EP-D1-D1 | 25 | 0.008 | 17e2d6534 | PRESENT | PRESENT | 176337 |
| Patient-2-NUC-7336-GEC2007985EP-D1-D1 | 60 | 0.005 | 17e2d6534 | PRESENT | PRESENT | 436213 |
| Patient-2-NUC-7337-GEC2007986EP-D1-D1 | 60 | 0.005 | 17e2d6534 | PRESENT | PRESENT | 471430 |
| Patient-3-NUC-7338-GEC2007987EP-D1-D1 | 10 | 0.100 | edaa8a4db | PRESENT | PRESENT | 91399 |
| Patient-3-NUC-7339-GEC2007988EP-D1-D1 | 10 | 0.100 | edaa8a4db | PRESENT | PRESENT | 91337 |
| Patient-3-NUC-7340-GEC2007989EP-D1-D1 | 10 | 0.100 | edaa8a4db | PRESENT | PRESENT | 106635 |
| Patient-3-NUC-7341-GEC2007990EP-D1-D1 | 10 | 0.050 | edaa8a4db | PRESENT | PRESENT | 107091 |
| Patient-3-NUC-7342-GEC2007991EP-D1-D1 | 10 | 0.050 | edaa8a4db | PRESENT | PRESENT | 107307 |
| Patient-3-NUC-7343-GEC2007992EP-D1-D1 | 10 | 0.050 | edaa8a4db | PRESENT | PRESENT | 106019 |
| Patient-3-NUC-7344-GEC2008017EP-D1-D1 | 25 | 0.050 | edaa8a4db | PRESENT | PRESENT | 210138 |
| Patient-3-NUC-7345-GEC2008018EP-D1-D1 | 25 | 0.050 | edaa8a4db | PRESENT | PRESENT | 221336 |
| Patient-3-NUC-7346-GEC2008019EP-D1-D1 | 25 | 0.050 | edaa8a4db | PRESENT | PRESENT | 224621 |
| Patient-3-NUC-7347-GEC2008020EP-D1-D1 | 10 | 0.010 | edaa8a4db | PRESENT | PRESENT | 96387 |
| Patient-3-NUC-7348-GEC2008021EP-D1-D1 | 10 | 0.010 | edaa8a4db | PRESENT | PRESENT | 95239 |
| Patient-3-NUC-7349-GEC2008022EP-D1-D1 | 10 | 0.010 | edaa8a4db | PRESENT | PRESENT | 82968 |
| Patient-3-NUC-7350-GEC2008023EP-D1-D1 | 25 | 0.010 | edaa8a4db | PRESENT | PRESENT | 160052 |
| Patient-3-NUC-7351-GEC2008024EP-D1-D1 | 25 | 0.010 | edaa8a4db | PRESENT | PRESENT | 179017 |
| Patient-3-NUC-7352-GEC2008025EP-D1-D1 | 25 | 0.010 | edaa8a4db | PRESENT | PRESENT | 231688 |
| Patient-3-NUC-7353-GEC2008026EP-D1-D1 | 60 | 0.008 | edaa8a4db | PRESENT | PRESENT | 458499 |
| Patient-3-NUC-7354-GEC2008027EP-D1-D1 | 60 | 0.008 | edaa8a4db | PRESENT | PRESENT | 453454 |
| Patient-3-NUC-7355-GEC2008028EP-D1-D1 | 10 | 0.008 | edaa8a4db | PRESENT | ABSENT | 99009 |
| Patient-3-NUC-7356-GEC2008029EP-D1-D1 | 10 | 0.008 | edaa8a4db | PRESENT | PRESENT | 97635 |
| Patient-3-NUC-7357-GEC2008030EP-D1-D1 | 10 | 0.008 | edaa8a4db | PRESENT | PRESENT | 91053 |
| Patient-3-NUC-7358-GEC2008031EP-D1-D1 | 25 | 0.008 | edaa8a4db | PRESENT | PRESENT | 294298 |
| Patient-3-NUC-7359-GEC2008032EP-D1-D1 | 25 | 0.008 | edaa8a4db | PRESENT | PRESENT | 217738 |
| Patient-3-NUC-7360-GEC2008053EP-D1-D1 | 25 | 0.008 | edaa8a4db | PRESENT | PRESENT | 246636 |
| Patient-3-NUC-7361-GEC2008054EP-D1-D1 | 60 | 0.005 | edaa8a4db | PRESENT | PRESENT | 452255 |
| Patient-3-NUC-7362-GEC2008055EP-D1-D1 | 60 | 0.005 | edaa8a4db | PRESENT | PRESENT | 461734 |
| Patient-4-NUC-7363-GEC2008056EP-D1-D1 | 10 | 0.100 | 9fa706f70 | PRESENT | PRESENT | 91080 |
| Patient-4-NUC-7364-GEC2008057EP-D1-D1 | 10 | 0.100 | 9fa706f70 | PRESENT | PRESENT | 94781 |
| Patient-4-NUC-7365-GEC2008058EP-D1-D1 | 10 | 0.100 | 9fa706f70 | PRESENT | PRESENT | 85575 |
| Patient-4-NUC-7366-GEC2008059EP-D1-D1 | 10 | 0.050 | 9fa706f70 | PRESENT | PRESENT | 83864 |
| Patient-4-NUC-7367-GEC2008060EP-D1-D1 | 10 | 0.050 | 9fa706f70 | PRESENT | PRESENT | 88258 |
| Patient-4-NUC-7368-GEC2008033EP-D1-D1 | 10 | 0.050 | 9fa706f70 | PRESENT | PRESENT | 78069 |
| Patient-4-NUC-7369-GEC2008034EP-D1-D1 | 25 | 0.050 | 9fa706f70 | PRESENT | PRESENT | 177698 |
| Patient-4-NUC-7370-GEC2008035EP-D1-D1 | 25 | 0.050 | 9fa706f70 | PRESENT | PRESENT | 193466 |
| Patient-4-NUC-7371-GEC2008036EP-D1-D1 | 25 | 0.050 | 9fa706f70 | PRESENT | PRESENT | 190016 |
| Patient-4-NUC-7372-GEC2008037EP-D1-D1 | 10 | 0.010 | 9fa706f70 | PRESENT | PRESENT | 74907 |
| Patient-4-NUC-7373-GEC2008038EP-D1-D1 | 10 | 0.010 | 9fa706f70 | PRESENT | PRESENT | 90951 |
| Patient-4-NUC-7374-GEC2008039EP-D1-D1 | 10 | 0.010 | 9fa706f70 | PRESENT | ABSENT | 81614 |
| Patient-4-NUC-7375-GEC2008040EP-D1-D1 | 25 | 0.010 | 9fa706f70 | PRESENT | PRESENT | 205811 |
| Patient-4-NUC-7376-GEC2008061EP-D1-D1 | 25 | 0.010 | 9fa706f70 | PRESENT | PRESENT | 193114 |
| Patient-4-NUC-7377-GEC2008062EP-D1-D1 | 25 | 0.010 | 9fa706f70 | PRESENT | PRESENT | 200233 |
| Patient-4-NUC-7378-GEC2008063EP-D1-D1 | 60 | 0.008 | 9fa706f70 | PRESENT | PRESENT | 382785 |
| Patient-4-NUC-7379-GEC2008064EP-D1-D1 | 60 | 0.008 | 9fa706f70 | PRESENT | PRESENT | 367380 |
| Patient-4-NUC-7380-GEC2008065EP-D1-D1 | 10 | 0.008 | 9fa706f70 | PRESENT | PRESENT | 81853 |
| Patient-4-NUC-7382-GEC2008067EP-D1-D1 | 10 | 0.008 | 9fa706f70 | PRESENT | ABSENT | 87334 |
| Patient-4-NUC-7383-GEC2008068EP-D1-D1 | 25 | 0.008 | 9fa706f70 | PRESENT | PRESENT | 181539 |
| Patient-4-NUC-7384-GEC2008049EP-D1-D1 | 25 | 0.008 | 9fa706f70 | PRESENT | PRESENT | 220975 |
| Patient-4-NUC-7385-GEC2008050EP-D1-D1 | 25 | 0.008 | 9fa706f70 | PRESENT | PRESENT | 205707 |
| Patient-4-NUC-7386-GEC2008051EP-D1-D1 | 60 | 0.005 | 9fa706f70 | PRESENT | PRESENT | 428183 |
| Patient-4-NUC-7387-GEC2008052EP-D1-D1 | 60 | 0.005 | 9fa706f70 | PRESENT | PRESENT | 418448 |
| Patient-NA-NUC-7388-GEC2008041EP-D1-D1 | 25 | 0.000 | abef6d7a6 | ABSENT | ABSENT | 162372 |
| Patient-NA-NUC-7389-GEC2008042EP-D1-D1 | 25 | 0.000 | abef6d7a6 | ABSENT | ABSENT | 102015 |
| Patient-NA-NUC-7390-GEC2008043EP-D1-D1 | 25 | 0.000 | 17e2d6534 | ABSENT | ABSENT | 186669 |
| Patient-NA-NUC-7391-GEC2008044EP-D1-D1 | 25 | 0.000 | 17e2d6534 | ABSENT | ABSENT | 106510 |
| Patient-NA-NUC-7392-GEC2008045EP-D1-D1 | 25 | 0.000 | edaa8a4db | ABSENT | ABSENT | 137143 |
| Patient-NA-NUC-7393-GEC2008046EP-D1-D1 | 25 | 0.000 | edaa8a4db | ABSENT | ABSENT | 231698 |
| Patient-NA-NUC-7394-GEC2008047EP-D1-D1 | 25 | 0.000 | 9fa706f70 | ABSENT | ABSENT | 160939 |
| Patient-NA-NUC-7395-GEC2008048EP-D1-D1 | 25 | 0.000 | 9fa706f70 | ABSENT | ABSENT | 202708 |

Supplementary Table 5. Summary of expected and observed ctDNA calls with baseline threshold in part 2 of accuracy study of ctDNA detection

| Lab Sample ID | Panel ID | ctDNA P value | Total Unique On-Target Reads with Depth of >= 5 Reads Covering Specific Variant Positions | Observed Result | Expected Result | Concordance |
| --- | --- | --- | --- | --- | --- | --- |
| GEV2006793 | 110041341 | 0.10 | 81865 | ABSENT | ABSENT | YES |
| GER2005721 | 110041341 | 0.01 | 12174 | PRESENT | PRESENT | YES |
| GEV2006794 | 110041385 | 0.88 | 884334 | ABSENT | ABSENT | YES |
| GER2005726 | 110041385 | <0.01 | 861256 | PRESENT | PRESENT | YES |
| GEV2006795 | 110041458 | 0.83 | 447065 | ABSENT | ABSENT | YES |
| GER2005268 | 110041458 | 0.00 | 219568 | PRESENT | PRESENT | YES |
| GEV2006796 | 110041626 | 0.81 | 170160 | ABSENT | ABSENT | YES |
| GER2005661 | 110041626 | <0.01 | 278509 | PRESENT | PRESENT | YES |
| GEV2006797 | 110041703 | 0.66 | 794158 | ABSENT | ABSENT | YES |
| GER2005413 | 110041703 | <0.01 | 1189204 | PRESENT | PRESENT | YES |
| GEV2006798 | 110041823 | 0.91 | 248029 | ABSENT | ABSENT | YES |
| GER2005689 | 110041823 | <0.01 | 604069 | PRESENT | PRESENT | YES |
| GEV2006799 | 110041824 | 0.77 | 440209 | ABSENT | ABSENT | YES |
| GER2005691 | 110041824 | 0.00 | 435353 | PRESENT | PRESENT | YES |
| GEV2006800 | 110041898 | 0.88 | 55898 | ABSENT | ABSENT | YES |
| GER2005690 | 110041898 | <0.01 | 166766 | PRESENT | PRESENT | YES |
| GEV2006785 | 110041921 | 0.70 | 87556 | ABSENT | ABSENT | YES |
| GER2005728 | 110041921 | 0.00 | 262377 | PRESENT | PRESENT | YES |
| GEV2006789 | 200016024 | 0.25 | 282712 | ABSENT | ABSENT | YES |
| GER2005663 | 200016024 | <0.01 | 207334 | PRESENT | PRESENT | YES |
| GEV2006786 | 200016102 | 0.26 | 82221 | ABSENT | ABSENT | YES |
| GER2005648 | 200016102 | 0.00 | 174603 | PRESENT | PRESENT | YES |
| GEV2006787 | 200016104 | 0.38 | 375384 | ABSENT | ABSENT | YES |
| GER2005662 | 200016104 | 0.00 | 70967 | PRESENT | PRESENT | YES |
| GEV2006788 | 200016105 | 0.01 | 210579 | ABSENT | ABSENT | YES |
| GER2005671 | 200016105 | 0.00 | 151380 | PRESENT | PRESENT | YES |
| GEV2006790 | 200016140 | 0.88 | 626317 | ABSENT | ABSENT | YES |
| GER2005737 | 200016140 | 0.00 | 155096 | PRESENT | PRESENT | YES |
| GEV2006791 | 200017118 | 0.99 | 614602 | ABSENT | ABSENT | YES |
| GER2005693 | 200017118 | <0.01 | 213784 | PRESENT | PRESENT | YES |
| GEV2006792 | 200017120 | 0.09 | 503975 | ABSENT | ABSENT | YES |
| GER2005736 | 200017120 | <0.01 | 431580 | PRESENT | PRESENT | YES |

Supplementary Table 6. Reproducibility and precision of ctDNA detection

| Operator | Planned Time | Sequencer | Panel ID | Observed Time | Total (N) | Number Agree | Percent Agree |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | 1 | NDX550351 | GPSA0083 | 1 | 2 | 2 | 100 |
| 1 | 1 | NDX550351 | GPSA0083 | 11 | 1 | 1 | 100 |
| 1 | 1 | NDX550351 | GPSA0084 | 11 | 1 | 1 | 100 |
| 1 | 1 | NDX550357 | GPSA0083 | 1 | 4 | 4 | 100 |
| 1 | 2 | NDX550351 | GPSA0083 | 2 | 6 | 6 | 100 |
| 1 | 2 | NDX550357 | GPSA0084 | 2 | 1 | 1 | 100 |
| 1 | 3 | NDX550351 | GPSA0084 | 3 | 6 | 6 | 100 |
| 1 | 3 | NDX550351 | GPSA0086 | 11 | 1 | 1 | 100 |
| 1 | 3 | NDX550351 | GPSA0087 | 11 | 1 | 1 | 100 |
| 1 | 3 | NDX550357 | GPSA0084 | 3 | 4 | 4 | 100 |
| 1 | 4 | NDX550351 | GPSA0085 | 11 | 1 | 1 | 100 |
| 1 | 4 | NDX550351 | GPSA0086 | 4 | 2 | 2 | 100 |
| 1 | 4 | NDX550351 | GPSA0088 | 11 | 1 | 1 | 100 |
| 1 | 4 | NDX550357 | GPSA0086 | 4 | 4 | 4 | 100 |
| 1 | 5 | NDX550351 | GPSA0086 | 5 | 6 | 6 | 100 |
| 1 | 5 | NDX550357 | GPSA0087 | 5 | 1 | 1 | 100 |
| 1 | 6 | NDX550351 | GPSA0087 | 6 | 6 | 6 | 100 |
| 1 | 6 | NDX550357 | GPSA0087 | 6 | 4 | 4 | 100 |
| 1 | 7 | NDX550351 | GPSA0085 | 7 | 2 | 2 | 100 |
| 1 | 7 | NDX550357 | GPSA0085 | 7 | 4 | 4 | 100 |
| 1 | 8 | NDX550351 | GPSA0085 | 10 | 6 | 6 | 100 |
| 1 | 8 | NDX550357 | GPSA0088 | 8 | 1 | 1 | 100 |
| 1 | 9 | NDX550351 | GPSA0088 | 9 | 6 | 6 | 100 |
| 1 | 9 | NDX550357 | GPSA0088 | 9 | 4 | 4 | 100 |
| 2 | 1 | NDX550349 | GPSA0083 | 1 | 2 | 2 | 100 |
| 2 | 1 | NDX550364 | GPSA0083 | 1 | 6 | 4 | 66.67 |
| 2 | 2 | NDX550349 | GPSA0083 | 10 | 2 | 2 | 100 |
| 2 | 2 | NDX550364 | GPSA0084 | 2 | 1 | 1 | 100 |
| 2 | 3 | NDX550349 | GPSA0084 | 3 | 4 | 4 | 100 |
| 2 | 3 | NDX550349 | GPSA0084 | 11 | 2 | 2 | 100 |
| 2 | 3 | NDX550364 | GPSA0084 | 3 | 4 | 4 | 100 |
| 2 | 4 | NDX550349 | GPSA0086 | 4 | 2 | 2 | 100 |
| 2 | 4 | NDX550364 | GPSA0086 | 4 | 6 | 6 | 100 |
| 2 | 5 | NDX550349 | GPSA0086 | 5 | 2 | 2 | 100 |
| 2 | 5 | NDX550364 | GPSA0087 | 5 | 1 | 1 | 100 |
| 2 | 6 | NDX550349 | GPSA0083 | 11 | 1 | 1 | 100 |
| 2 | 6 | NDX550349 | GPSA0084 | 11 | 1 | 1 | 100 |
| 2 | 6 | NDX550349 | GPSA0087 | 6 | 6 | 6 | 100 |
| 2 | 6 | NDX550364 | GPSA0087 | 6 | 4 | 4 | 100 |
| 2 | 7 | NDX550349 | GPSA0085 | 7 | 2 | 2 | 100 |
| 2 | 7 | NDX550349 | GPSA0086 | 11 | 1 | 1 | 100 |
| 2 | 7 | NDX550349 | GPSA0087 | 11 | 1 | 1 | 100 |
| 2 | 7 | NDX550364 | GPSA0085 | 7 | 6 | 6 | 100 |
| 2 | 8 | NDX550349 | GPSA0085 | 8 | 2 | 2 | 100 |
| 2 | 8 | NDX550364 | GPSA0088 | 8 | 1 | 1 | 100 |
| 2 | 9 | NDX550349 | GPSA0085 | 11 | 1 | 1 | 100 |
| 2 | 9 | NDX550349 | GPSA0088 | 9 | 6 | 6 | 100 |
| 2 | 9 | NDX550349 | GPSA0088 | 11 | 1 | 1 | 100 |
| 2 | 9 | NDX550364 | GPSA0088 | 9 | 4 | 4 | 100 |
| 1 | 1 | NDX550351 | GPSA0083 | 1 | 1 | 1 | 100 |
| 1 | 1 | NDX550357 | GPSA0083 | 1 | 2 | 2 | 100 |
| 1 | 2 | NDX550351 | GPSA0083 | 2 | 3 | 3 | 100 |
| 1 | 7 | NDX550351 | GPSA0085 | 7 | 1 | 1 | 100 |
| 1 | 7 | NDX550357 | GPSA0085 | 7 | 2 | 2 | 100 |
| 1 | 8 | NDX550351 | GPSA0085 | 10 | 3 | 3 | 100 |
| 2 | 1 | NDX550349 | GPSA0083 | 1 | 1 | 1 | 100 |
| 2 | 1 | NDX550364 | GPSA0083 | 1 | 3 | 2 | 66.67 |
| 2 | 2 | NDX550349 | GPSA0083 | 10 | 1 | 1 | 100 |
| 2 | 7 | NDX550349 | GPSA0085 | 7 | 1 | 1 | 100 |
| 2 | 7 | NDX550364 | GPSA0085 | 7 | 3 | 3 | 100 |
| 2 | 8 | NDX550349 | GPSA0085 | 8 | 1 | 1 | 100 |

Supplementary Table 7. End-to-end study: ctDNA results for each library of four patients

| Library ID | Patient ID | Panel ID | WES Input Mass | Observed ctDNA Result | Expected ctDNA Result |
| --- | --- | --- | --- | --- | --- |
| NUC-6226-10b87ff68-GEC2007500EP-D1-D1 | AP00002 | 10b87ff68 | 500 | PRESENT | PRESENT |
| NUC-6227-638ed3d27-GEC2007501EP-D1-D1 | AP00002 | 638ed3d27 | 500 | PRESENT | PRESENT |
| NUC-6228-0926e70f5-GEC2007502EP-D1-D1 | AP00002 | 0926e70f5 | 500 | PRESENT | PRESENT |
| NUC-6229-86484e7dd-GEC2007503EP-D1-D1 | AP00002 | 86484e7dd | 200 | PRESENT | PRESENT |
| NUC-6230-c301adef5-GEC2007504EP-D1-D1 | AP00002 | c301adef5 | 200 | PRESENT | PRESENT |
| NUC-6232-2d33d1d49-GEC2007506EP-D1-D1 | AP00004 | 2d33d1d49 | 200 | ABSENT | ABSENT |
| NUC-6233-1ddd3d01b-GEC2007507EP-D1-D1 | AP00004 | 1ddd3d01b | 500 | ABSENT | ABSENT |
| NUC-6234-7025fa6b7-GEC2007508EP-D1-D1 | AP00004 | 7025fa6b7 | 200 | ABSENT | ABSENT |
| NUC-6235-c7f11ce6d-GEC2007509EP-D1-D1 | AP00004 | c7f11ce6d | 200 | ABSENT | ABSENT |
| NUC-6236-bef668ce2-GEC2007510EP-D1-D1 | AP00004 | bef668ce2 | 500 | ABSENT | ABSENT |
| NUC-6237-7dac1f052-GEC2007511EP-D1-D1 | AP00004 | 7dac1f052 | 200 | ABSENT | ABSENT |
| NUC-6244-2a2682bb6-GEC2007512EP-D1-D1 | AP00006 | 2a2682bb6 | 200 | ABSENT | ABSENT |
| NUC-6245-33cef63ca-GEC2007513EP-D1-D1 | AP00006 | 33cef63ca | 500 | ABSENT | ABSENT |
| NUC-6246-49a79a60d-GEC2007514EP-D1-D1 | AP00006 | 49a79a60d | 500 | ABSENT | ABSENT |
| NUC-6247-445ddcf65-GEC2007515EP-D1-D1 | AP00006 | 445ddcf65 | 500 | ABSENT | ABSENT |
| NUC-6238-50242344a-GEC2007518EP-D1-D1 | AP00007 | 50242344a | 500 | PRESENT | PRESENT |
| NUC-6239-1e2a70da1-GEC2007519EP-D1-D1 | AP00007 | 1e2a70da1 | 200 | PRESENT | PRESENT |
| NUC-6240-2e90bf9a9-GEC2007520EP-D1-D1 | AP00007 | 2e90bf9a9 | 500 | PRESENT | PRESENT |
| NUC-6241-4be650dc4-GEC2007521EP-D1-D1 | AP00007 | 4be650dc4 | 500 | PRESENT | PRESENT |
| NUC-6248-4ed666804-GEC2007522EP-D1-D1 | AP00007 | 4ed666804 | 200 | PRESENT | PRESENT |

Supplementary Figure 1. Samples and specimens used in the studies