## Supplemental Information

Table S1: Operating points of the circulated APG CAST generator. The particle number size distribution is characterised by SMPS scans. Also, burner flow settings for each operating point are given.

| APG Op. Point | Classified Diameter | SMPS Mode | SMPS GSD | Propane | Mixing N2 | Oxidation Air | Burner Air | VPR Dil. Flow | C:O ratio |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| - | nm | nm | - | ml/min | ml/min | l/min | l/min | l/min |  |
| A | 10, 15 | 10 | 1.3 | 18 | 30 | 0.4 | 5 | 3 | 0.33 |
| B | 23, 33 | 12.5 | 1.41 | 18 | 20 | 0.4 | 5 | 3 | 0.33 |
| C | 41, 59 | 23 | 1.61 | 25 | 0 | 0.45 | 5 | 3 | 0.41 |
| D | 70, 103 | 46 | 1.94 | 20 | 15 | 0.66 | 5 | 3 | 0.22 |
| E | 100, 150 | 82 | 2.07 | 18 | 0 | 0.5 | 5 | 3 | 0.26 |

Table S2: Additional Operating Parameters of Laboratory Equipment

| Laboratory | Lab No | DMA Flow Sheath/Sample | DMA Pressure | Neutralizer | Reference Correction Factor | Ref. Flow [l/min] |
| --- | --- | --- | --- | --- | --- | --- |
| TSI | Mfr. | n/a | Under | n/a | 1,0006 | 3 |
| JRC | #1a/b | 10/1.5 | Under | Grimm 0,1mCi AM-241 | 0,9 | 1.5 |
| AVL | #2 | 15/2.0 | under | TSI 10mCi KR-85 | 1 | 1 |
| PTB | #3 | 15/1.5 | Over | TSI 10mCi KR-85 | 1 | 3 |
| BMW | #4 | 15/1.5 | Over | TSI 10mCi KR-85 | 1 | 1 |
| Ricardo E&E | #5 | 10/1.0 | Under | TSI 10mCi KR-85 | 1.053 - 1.073 (size dependent) | 1 |
| VW | #6 | 10/1.0 | Under | TSI 2mCi KR-85 | 1 | 1 |

Table S3: Operating points of PTB Mini-CAST triple flame miniCAST 5203C. The particle number size distribution is characterised by SMPS scans. Also, burner flow settings for each operating point are given.

| Op. Point | Propane | Oxidation Air | Mixing N2 | SMPS Mode | SMPS GSD |
| --- | --- | --- | --- | --- | --- |
| - | ml/min | l/min | l/min | nm |  |
| 1 | 135 | 0.90 | 3.00 | 11.85 | 1.32 |
| 2 | 140 | 0.87 | 3.25 | 22.61 | 1.35 |
| 3 | 160 | 0.73 | 3.55 | 40.48 | 1.37 |
| 4 | 150 | 0.65 | 3.30 | 55.01 | 1.43 |
| 5 | 155 | 0.50 | 3.00 | 60.88 | 1.41 |
| 6 | 155 | 0.40 | 3.00 | 71.92 | 1.46 |
| 7 | 145 | 0.43 | 3.00 | 79.02 | 1.47 |
| 8 | 140 | 0.40 | 3.00 | 83.70 | 1.48 |
| 9 | 155 | 0.40 | 3.50 | 90.13 | 1.50 |
| 10 | 155 | 0.30 | 3.50 | 100.41 | 1.50 |

Figure S1: crCPC sample flow measurements (diamonds) in chronological order. The room temperature (crosses) is plotted on the secondary y-axis for reference.