SUPPLEMENTARY INFORMATION

Tellurium Speciation via Frontal Chromatography ICP-MS: An Innovative and Effective Approach for the Ultratrace Detection in Environmental Matrices

Davide Spanua,*, Chiara Omodeia, Gilberto Bindab,a, Tommaso Grandea, Sandro Recchiaa

a Department of Science and High Technology, University of Insubria, Via Valleggio 11, 22100 Como, Italy

b Norwegian Institute for Water Research (NIVA), Økernveien 94, 0579 Oslo, Norway

* Corresponding Author

Mail: davide.spanu@uninsubria.it

Phone: +39 0312386428

Table S1. ICP-MS working parameters used under optimized conditions.

| Parameter | Value |
| --- | --- |
| RF Power | 1550 W |
| Nebulizer gas flow | 1.06 L/min |
| Auxiliary gas flow | 0.80 L/min |
| Cooling gas flow | 13.99 L/min |
| Dwell time | 160 ms |
| Monitored channels (m/z) | 126Te,73Ge and 129Xe(*) |
| Sample flow rate | 1.39 mL/min |
| Methanol flow rate | 0.20 mL/min |

(*) Other monitored isotopes during optimization stage: 125Te, 128Te, 130Te

Figure S1. Illustration of analytical estimation for resolution and analysis time. Frontal chromatogram depicting a Te(IV)-Te(VI) mixture solution (black line), derivative function utilized for simulating an elution chromatogram (green line). Each peak is modeled with a Gaussian function. Resolution and analysis time are quantified based on the equations provided in the Figure, adhering to the standard definitions for these parameters in elution chromatography.

Figure S2. Frontal chromatogram of a 0.25 M HNO3 solution containing 1 μg kg−1 of both Te(VI) and Te(IV) species (red line, raw 125Te signal). The signal of the internal standard 73Ge is depicted in blue. Experimental conditions are detailed in Figure 3.

Table S2. Chemical composition and physiochemical parameters of mineral waters analyzed in this work.

| Parameter | Wat-1 | Wat-2 | Wat-3 | Wat-4 | Wat-5 | Wat-6 |
| --- | --- | --- | --- | --- | --- | --- |
| Sodium (mg/l) | 0.58 | 0.85 | 4.31 | 3.5 | 4.9 | 50 |
| Calcium (mg/l) | 10.8 | 33.4 | 56.82 | 96 | 120 | 400 |
| Magnesium (mg/l) | 0.39 | 5.5 | 4 | 29 | 40 | 23 |
| Potassium (mg/l) | 0.45 | 0.41 | 0.42 | 1.2 | 1.7 | 48 |
| Silica (mg/l) | 3.5 | 4.8 | 5.8 | nd | nd | 87 |
| Bicarbonate (mg/l) | 32.3 | 127 | 182 | 313 | 303 | 1500 |
| Sulfate (mg/l) | 1.5 | 2.5 | 7.56 | 106 | 222 | 4.5 |
| Nitrate (mg/l) | 1 | 0.96 | 1.34 | 2.9 | 5.6 | 4 |
| Chloride (mg/l) | 0.55 | 0.95 | 8.09 | nd | 4.9 | 16.9 |
| Fixed residue at 180°C (mg/l) | 34.9 | 112.1 | 178.4 | 422 | 566 | 1300 |
| Conductivity at 20°C (μS/cm) | 52 | 171 | 303.9 | 582 | 767 | 1780 |
| pH | 7 | 7.4 | 7.58 | 7.4 | 7.2 | 6.2 |

Table S3. Summary of the experimental conditions explored during DoE analysis and corresponding resolution and overall analysis time.

| N° exp. | Column width (mm) | Column length (mm) | HNO3 conc. (M) | Sample flow rate (mL min-1) | Resolution | Analysis time (s) |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | 2 | 50 | 0.5 | 1.39 | 0.73 | 95.8 |
| 2 | 2 | 50 | 0.5 | 1.14 | 0.53 | 112.9 |
| 3 | 2 | 50 | 0.25 | 1.39 | 0.78 | 181.0 |
| 4 | 2 | 50 | 0.25 | 1.14 | 0.99 | 198.8 |
| 5 | 2 | 25 | 0.5 | 1.39 | 0.38 | 59.1 |
| 6 | 2 | 25 | 0.5 | 1.14 | 0.37 | 70.3 |
| 7 | 2 | 25 | 0.25 | 1.39 | 0.52 | 82.6 |
| 8 | 2 | 25 | 0.25 | 1.14 | 0.51 | 97.2 |
| 9 | 3 | 50 | 0.5 | 1.39 | 1.05 | 147.9 |
| 10 | 3 | 50 | 0.5 | 1.14 | 1.00 | 172.6 |
| 11 | 3 | 50 | 0.25 | 1.39 | 1.77 | 238.5 |
| 12 | 3 | 50 | 0.25 | 1.14 | 1.71 | 289.8 |
| 13 | 3 | 25 | 0.5 | 1.39 | 0.52 | 91.3 |
| 14 | 3 | 25 | 0.5 | 1.14 | 0.54 | 107.4 |
| 15 | 3 | 25 | 0.25 | 1.39 | 0.89 | 143.0 |
| 16 | 3 | 25 | 0.25 | 1.14 | 0.87 | 171.6 |

Figure S3. Surface responses depicting the effect of (a,b) sample flow rate and nitric acid concentration and (c,d) column geometry on the resolution and analysis time. Chromatographic conditions are detailed in their respective panels.

Figure S4. Scatter plots depicting Experimental versus Fitted resolution and analysis time data resulting from multivariate regression. Each data point is labeled with a number corresponding to the specific experimental condition (refer to Table S3 for details).

Figure S5. Frontal chromatogram obtained for a Te(IV)-Te(VI) mixture solution (both species at a concentration equal to 1 μg kg−1) under the following conditions: column internal diameter = 3 mm, column length = 50 mm, HNO3 concentration = 0.5 M, and sample flow rate = 1.39 mL min-1.

Table S4. Formulas used to correct Xe interferences.

| Isotope | Correction formula |
| --- | --- |
| 125Te | No corrections |
| 126Te | 126Te - 0,00340909 × 129Xe |
| 128Te | 128Te - 0,0723485 × 129Xe |

Figure S6. Calibrations curves obtained under optimized conditions for Te(IV) and Te(VI) using 125Te (red), 126Te (green) and 128Te (blue) isotopes.

Figure S7. Frontal chromatograms of 5 ng kg−1 Te(IV) and Te(VI) solutions analyzed for the estimation of LOD under optimized conditions (ten replicates).