RESEARCH ARTICLE

Next-Generation Digital Polymerase Chain Reaction Assay for BCR::ABL1 in Chronic Myeloid Leukemia: Comprehensive Validation Enabling Standardized Measurable Residual Disease Detection

Yong-Hwi Kim1, Min-Sik Jang1, Shin-Hyun Kim1, In-Young Roh1, Bong-Suk Kim1, Kyung-Mi Kee2, Soo-Hyun Kim2, Dong-Wook Kim3, and Jae-Hoon Kim1*

1BioTNS Co., Ltd., Daejeon 34183, Republic of Korea

2Leukemia Omics Research Institute, Eulji University, Uijeongbu 11759, Republic of Korea

3Department of Hematology, Hematology Center, Eulji Medical Center, Eulji University, Uijeongbu 11759, Republic of Korea

* Correspondence to: Jae-Hoon Kim, Ph.D.

BioTNS Co., Ltd., Daejeon 34183, Republic of Korea

E-mail: jhkim@biotns.co.kr
Supplementary Table 1. Limit of blank (LoB) analysis results based on 60 replicate measurements using eight HL‑60 negative cell line samples (BCR::ABL1 fusion–negative)

| No. | Total droplet | BCR::ABL11 |  | BCR2 |  |  | No. | Total droplet | BCR::ABL1 |  | BCR |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | Positive droplet | Conc.3 | Positive droplet | Conc. |  |  |  | Positive droplet | Conc.* | Positive droplet | Conc. |
| 1 | 20,084 | 0.00 | 0.00 | 19,252 | 8,019.76 |  | 31 | 21,445 | 0 | 0.00 | 19,972 | 6,746.07 |
| 2 | 18,596 | 0.00 | 0.00 | 17,519 | 7,175.74 |  | 32 | 22,385 | 0 | 0.00 | 21,334 | 7,704.41 |
| 3 | 19,493 | 0.00 | 0.00 | 18,448 | 7,370.37 |  | 33 | 22,099 | 0 | 0.00 | 20,656 | 6,873.57 |
| 4 | 17,943 | 0.00 | 0.00 | 17,123 | 7,772.42 |  | 34 | 22,104 | 0 | 0.00 | 20,003 | 5,927.82 |
| 5 | 19,539 | 0.00 | 0.00 | 18,591 | 7,621.70 |  | 35 | 18,217 | 0 | 0.00 | 17,305 | 7,542.75 |
| 6 | 19,152 | 0.00 | 0.00 | 18,440 | 8,292.40 |  | 36 | 19,345 | 0 | 0.00 | 18,133 | 6,977.74 |
| 7 | 20,377 | 0.00 | 0.00 | 19,481 | 7,869.58 |  | 37 | 17,842 | 0 | 0.00 | 16,884 | 7,366.40 |
| 8 | 19,625 | 0.00 | 0.00 | 18,424 | 7,036.90 |  | 38 | 19,559 | 0 | 0.00 | 18,391 | 7,098.60 |
| 9 | 17,870 | 0.00 | 0.00 | 17,432 | 9,341.71 |  | 39 | 19,051 | 0 | 0.00 | 17,514 | 6,340.77 |
| 10 | 19,407 | 0.00 | 0.00 | 18,590 | 7,979.22 |  | 40 | 24,214 | 0 | 0.00 | 23,113 | 7,785.17 |
| 11 | 16,612 | 0.00 | 0.00 | 15,806 | 7,621.65 |  | 41 | 23,749 | 0 | 0.00 | 21,218 | 5,639.61 |
| 12 | 19,021 | 0.00 | 0.00 | 18,217 | 7,969.02 |  | 42 | 19,997 | 0 | 0.00 | 18,885 | 7,278.14 |
| 13 | 22,795 | 0.00 | 0.00 | 21,411 | 7,056.83 |  | 43 | 24,271 | 0 | 0.00 | 22,885 | 7,211.24 |
| 14 | 19,101 | 0.00 | 0.00 | 18,455 | 8,530.72 |  | 44 | 19,552 | 0 | 0.00 | 18,736 | 8,001.05 |
| 15 | 17,697 | 0.00 | 0.00 | 17,170 | 8,851.26 |  | 45 | 21,944 | 0 | 0.00 | 21,229 | 8,624.60 |
| 16 | 17,719 | 0.00 | 0.00 | 16,943 | 7,879.70 |  | 46 | 16,053 | 0 | 0.00 | 15,347 | 7,869.11 |
| 17 | 19,338 | 0.00 | 0.00 | 18,329 | 7,438.57 |  | 47 | 14,992 | 0 | 0.00 | 13,858 | 6,503.19 |
| 18 | 19,068 | 0.00 | 0.00 | 18,636 | 9,539.90 |  | 48 | 24,959 | 0 | 0.00 | 22,313 | 5,652.86 |
| 19 | 20,467 | 0.00 | 0.00 | 19,449 | 7,559.13 |  | 49 | 18,627 | 0 | 0.00 | 17,542 | 7,161.29 |
| 20 | 19,649 | 0.00 | 0.00 | 18,632 | 7,458.86 |  | 50 | 17,974 | 0 | 0.00 | 17,294 | 8,248.33 |
| 21 | 18,471 | 0.00 | 0.00 | 17,686 | 7,955.35 |  | 51 | 18,675 | 0 | 0.00 | 16,908 | 5,939.30 |
| 22 | 23,572 | 0.00 | 0.00 | 20,994 | 5,574.42 |  | 52 | 18,752 | 0 | 0.00 | 17,742 | 7,358.56 |
| 23 | 24,383 | 0.00 | 0.00 | 23,332 | 7,919.76 |  | 53 | 23,132 | 1 | 0.11 | 21,667 | 6,950.53 |
| 24 | 20,422 | 0.00 | 0.00 | 19,325 | 7,365.32 |  | 54 | 20,896 | 1 | 0.12 | 19,865 | 7,579.42 |
| 25 | 17,321 | 0.00 | 0.00 | 16,191 | 6,875.82 |  | 55 | 19,034 | 1 | 0.13 | 18,055 | 7,474.69 |
| 26 | 21,893 | 0.00 | 0.00 | 20,058 | 6,244.64 |  | 56 | 18,201 | 1 | 0.14 | 16,145 | 5,492.98 |
| 27 | 21,491 | 0.00 | 0.00 | 19,167 | 5,602.88 |  | 57 | 17,853 | 1 | 0.14 | 16,549 | 6,591.27 |
| 28 | 19,301 | 0.00 | 0.00 | 18,416 | 7,764.04 |  | 58 | 16,202 | 1 | 0.16 | 15,656 | 8,539.73 |
| 29 | 24,543 | 0.00 | 0.00 | 22,920 | 6,841.69 |  | 59 | 16,096 | 1 | 0.16 | 15,124 | 7,070.45 |
| 30 | 18,593 | 0.00 | 0.00 | 17,498 | 7,133.58 |  | 60 | 15,589 | 1 | 0.16 | 14,680 | 7,158.63 |

1, FAM fluorescence; 2, HEX fluorescence; 3, copies/μL.

Supplementary Table 2. Limit of detection (LoD) and Limit of quantitation (LoQ) results for each BCR::ABL1 transcript type (e13a2, e14a2) based on cell line–derived serial dilution panels (K‑562 and KCL‑22 spiked into HL‑60), determined using a non-parametric approach

| Lot | Transcript identity | Total specimens | Positive specimens | Positive detection rate (%) | CV (%) |  | Molecular indicator |  | Results |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  | %IS | MR | %IS | MR |  |
| 1 | e13a2 | 24 | 24 | 100.00 | 34.35 | 3.54 | 0.0032 | 4.49 | Pass |
| 1 | e14a2 | 24 | 24 | 100.00 | 48.83 | 4.42 | 0.0032 | 4.49 | Pass |
| 1 | e13a2 | 24 | 23 | 95.83 | 42.21 | 5.13 | 0.0025 | 4.60 | Pass |
| 1 | e14a2 | 24 | 23 | 95.83 | 37.83 | 4.17 | 0.0025 | 4.60 | Pass |
| 1 | e13a2 | 24 | 17 | 70.83 | 68.47 | 5.55 | 0.0020 | 4.70 | Fail |
| 1 | e14a2 | 24 | 17 | 70.83 | 72.83 | 4.36 | 0.0020 | 4.70 | Fail |
| 1 | e13a2 | 24 | 10 | 41.67 | 93.46 | 6.89 | 0.0016 | 4.80 | Fail |
| 1 | e14a2 | 24 | 13 | 54.17 | 98.70 | 5.46 | 0.0016 | 4.80 | Fail |
| 1 | e13a2 | 24 | 10 | 41.67 | 107.98 | 7.99 | 0.0013 | 4.89 | Fail |
| 1 | e14a2 | 24 | 10 | 41.67 | 111.98 | 6.16 | 0.0013 | 4.89 | Fail |
| 1 | e13a2 | 24 | 8 | 33.33 | 122.37 | 7.31 | 0.0010 | 5.00 | Fail |
| 1 | e14a2 | 24 | 7 | 29.17 | 159.74 | 5.38 | 0.0010 | 5.00 | Fail |
| 2 | e13a2 | 24 | 24 | 100.00 | 39.03 | 4.21 | 0.0032 | 4.49 | Pass |
| 2 | e14a2 | 24 | 24 | 100.00 | 36.14 | 4.06 | 0.0032 | 4.49 | Pass |
| 2 | e13a2 | 24 | 23 | 95.83 | 41.60 | 4.37 | 0.0025 | 4.60 | Pass |
| 2 | e14a2 | 24 | 23 | 95.83 | 52.06 | 4.30 | 0.0025 | 4.60 | Pass |
| 2 | e13a2 | 24 | 18 | 75.00 | 79.65 | 7.69 | 0.0020 | 4.70 | Fail |
| 2 | e14a2 | 24 | 17 | 70.83 | 75.81 | 5.27 | 0.0020 | 4.70 | Fail |
| 2 | e13a2 | 24 | 10 | 41.67 | 124.49 | 6.60 | 0.0016 | 4.80 | Fail |
| 2 | e14a2 | 24 | 14 | 58.33 | 82.79 | 7.49 | 0.0016 | 4.80 | Fail |
| 2 | e13a2 | 24 | 9 | 37.50 | 127.51 | 9.18 | 0.0013 | 4.89 | Fail |
| 2 | e14a2 | 24 | 9 | 37.50 | 131.76 | 9.55 | 0.0013 | 4.89 | Fail |
| 2 | e13a2 | 24 | 11 | 45.83 | 111.84 | 6.76 | 0.0010 | 5.00 | Fail |
| 2 | e14a2 | 24 | 10 | 41.67 | 116.57 | 7.92 | 0.0010 | 5.00 | Fail |

CV, coefficient of variation; %IS, % international scale; MR, molecular response. Bold indicates the LoD and LoQ ranges.

Supplementary Table 3. Linearity assessment for the major BCR::ABL1 fusion transcript (e13a2)

| Lot 1 |  |  |  |  |  | Lot 2 |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Samples | %IS | Mean | MR | Mean |  | Samples | %IS | Mean | MR | Mean |
| 10.4301 %IS (MR0.9817) | 8.6624 | 8.2355 | 1.0624 | 1.0854 |  | 10.4301 %IS (MR0.9817) | 8.6068 | 8.8236 | 1.0652 | 1.0561 |
|  | 7.4337 |  | 1.1288 |  |  |  | 9.8999 |  | 1.0044 |  |
|  | 8.6103 |  | 1.0650 |  |  |  | 7.9640 |  | 1.0989 |  |
| 1.1549 %IS (MR1.9375) | 1.0332 | 1.1629 | 1.9858 | 1.9380 |  | 1.1549 %IS (MR1.9375) | 1.1323 | 1.1468 | 1.9460 | 1.9438 |
|  | 1.3765 |  | 1.8612 |  |  |  | 1.3266 |  | 1.8772 |  |
|  | 1.0791 |  | 1.9669 |  |  |  | 0.9813 |  | 2.0082 |  |
| 0.1282 %IS (MR2.8921) | 0.0979 | 0.0968 | 3.0092 | 3.0143 |  | 0.1282 %IS (MR2.8921) | 0.1485 | 0.1166 | 2.8284 | 2.9414 |
|  | 0.0915 |  | 3.0388 |  |  |  | 0.0941 |  | 3.0265 |  |
|  | 0.1012 |  | 2.9950 |  |  |  | 0.1073 |  | 2.9692 |  |
| 0.0101 %IS (MR3.9966) | 0.0129 | 0.0134 | 3.8908 | 3.8799 |  | 0.0101 %IS (MR3.9966) | 0.0176 | 0.0168 | 3.7557 | 3.7752 |
|  | 0.0109 |  | 3.9619 |  |  |  | 0.0163 |  | 3.7888 |  |
|  | 0.0163 |  | 3.7869 |  |  |  | 0.0166 |  | 3.7812 |  |
| 0.0023 %IS (MR4.6399) | 0.0042 | 0.0047 | 4.3782 | 4.3313 |  | 0.0023 %IS (MR4.6399) | 0.0051 | 0.0055 | 4.2944 | 4.2638 |
|  | 0.0048 |  | 4.3224 |  |  |  | 0.0060 |  | 4.2214 |  |
|  | 0.0051 |  | 4.2933 |  |  |  | 0.0053 |  | 4.2755 |  |
| 0.0006 %IS (MR5.2596) | 0.0022 | 0.0024 | 4.6557 | 4.6277 |  | 0.0006 %IS (MR5.2596) | 0.0018 | 0.0039 | 4.7546 | 4.4569 |
|  | 0.0028 |  | 4.5577 |  |  |  | 0.0060 |  | 4.2236 |  |
|  | 0.0021 |  | 4.6697 |  |  |  | 0.0040 |  | 4.3926 |  |
| Linearity (R2) | 0.9928 |  |  |  |  | Linearity (R2) | 0.9939 |  |  |  |
| CF | 0.8001 |  |  |  |  | CF | 1.0836 |  |  |  |

*, Negative. Evaluated with two independent lots (not the same as analytical accuracy lots).

Supplementary Table 4. Linearity assessment for the major BCR::ABL1 fusion transcript (e14a2)

| Lot 1 |  |  |  |  |  | Lot 2 |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Samples | %IS | Mean | MR | Mean |  | Samples | %IS | Mean | MR | Mean |
| 10.6813 %IS (MR0.9714) | 10.4742 | 10.4713 | 0.9799 | 0.9800 |  | 10.6813 %IS (MR0.9714) | 8.8772 | 9.3911 | 1.0517 | 1.0280 |
|  | 10.4792 |  | 0.9797 |  |  |  | 9.1499 |  | 1.0386 |  |
|  | 10.4606 |  | 0.9804 |  |  |  | 10.1463 |  | 0.9937 |  |
| 1.0500 %IS (MR1.9788) | 0.8242 | 0.8347 | 2.0840 | 2.0795 |  | 1.0500 %IS (MR1.9788) | 1.0390 | 1.0164 | 1.9834 | 1.9950 |
|  | 0.7689 |  | 2.1141 |  |  |  | 0.8866 |  | 2.0523 |  |
|  | 0.9110 |  | 2.0405 |  |  |  | 1.1235 |  | 1.9494 |  |
| 0.0990 %IS (MR3.0045) | 0.1356 | 0.1242 | 2.8679 | 2.9113 |  | 0.0990 %IS (MR3.0045) | 0.1262 | 0.1299 | 2.8988 | 2.8899 |
|  | 0.0977 |  | 3.0103 |  |  |  | 0.1117 |  | 2.9521 |  |
|  | 0.1394 |  | 2.8557 |  |  |  | 0.1517 |  | 2.8190 |  |
| 0.0082 %IS (MR4.0838) | 0.0119 | 0.0112 | 3.9232 | 3.9530 |  | 0.0082 %IS (MR4.0838) | 0.0136 | 0.0143 | 3.8675 | 3.8459 |
|  | 0.0093 |  | 4.0333 |  |  |  | 0.0156 |  | 3.8057 |  |
|  | 0.0125 |  | 3.9026 |  |  |  | 0.0137 |  | 3.8644 |  |
| 0.0022 %IS (MR4.6484) | 0.0059 | 0.0056 | 4.2294 | 4.2670 |  | 0.0022 %IS (MR4.6484) | 0.0064 | 0.0060 | 4.1966 | 4.2187 |
|  | 0.0037 |  | 4.4299 |  |  |  | 0.0059 |  | 4.2325 |  |
|  | 0.0072 |  | 4.1417 |  |  |  | 0.0059 |  | 4.2271 |  |
| 0.0010 %IS (MR4.9792) | 0.0007 | 0.0016 | 5.1493 | 4.8475 |  | 0.0010 %IS (MR4.9792) | 0.0052 | 0.0054 | 4.2841 | 4.2768 |
|  | 0.0022 |  | 4.6495 |  |  |  | 0.0064 |  | 4.1911 |  |
|  | 0.0018 |  | 4.7438 |  |  |  | 0.0044 |  | 4.3552 |  |
| Linearity (R2) | 0.9916 |  |  |  |  | Linearity (R2) | 0.9972 |  |  |  |
| CF | 0.8001 |  |  |  |  | CF | 1.0836 |  |  |  |

*, Negative. Evaluated with two independent lots (not the same as analytical accuracy lots).

Supplementary Table 5. Comparison of %IS and MR values across three lots of the LQ Check BCR::ABL1 Major %IS Kit, traceable to the WHO primary reference standard

| Lot 1 |  |  |  | Lot 2 |  |  |  | Lot 3 |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Samples |  | %IS | MR | Samples |  | %IS | MR | Samples |  | %IS | MR |
| 08/198 (16.3129 %IS, MR0.7875) | Mean | 16.4105 | 0.7850 | 08/198 (16.3129 %IS, MR0.7875) | Mean | 16.2001 | 0.7911 | 08/198 (16.3129 %IS, MR0.7875) | Mean | 15.1063 | 0.8209 |
|  | Bias | -0.0976 | 0.0025 |  | Bias | -0.1128 | -0.0036 |  | Bias | 1.2066 | -0.0335 |
|  | SD | 0.4566 | 0.0121 |  | SD | 1.0470 | 0.0276 |  | SD | 0.3510 | 0.0100 |
|  | CV (%) | 2.7827 | 1.5432 |  | CV (%) | 6.4632 | 3.4920 |  | CV (%) | 2.3233 | 1.2229 |
| 08/196 (1.6627 %IS, MR1.7792) | Mean | 1.9564 | 1.7096 | 08/196 (1.6627 %IS, MR1.7792) | Mean | 1.8240 | 1.7407 | 08/196 (1.6627 %IS, MR1.7792) | Mean | 1.9066 | 1.7202 |
|  | Bias | -0.2937 | 0.0696 |  | Bias | -0.1613 | 0.0385 |  | Bias | -0.2439 | 0.0590 |
|  | SD | 0.1694 | 0.0379 |  | SD | 0.1937 | 0.0477 |  | SD | 0.1071 | 0.0247 |
|  | CV (%) | 8.6563 | 2.2148 |  | CV (%) | 10.6183 | 2.7412 |  | CV (%) | 5.6169 | 1.4378 |
| 08/194 (0.1753 %IS, MR2.7562) | Mean | 0.1675 | 2.7776 | 08/194 (0.1753 %IS, MR2.7562) | Mean | 0.1545 | 2.8152 | 08/194 (0.1753 %IS, MR2.7562) | Mean | 0.2115 | 2.6747 |
|  | Bias | 0.0078 | -0.0213 |  | Bias | 0.0208 | -0.0590 |  | Bias | -0.0362 | 0.0815 |
|  | SD | 0.0177 | 0.0468 |  | SD | 0.0250 | 0.0735 |  | SD | 0.0028 | 0.0057 |
|  | CV (%) | 10.5500 | 1.6866 |  | CV (%) | 16.2156 | 2.6107 |  | CV (%) | 1.3127 | 0.2137 |
| 08/192 (0.0195 %IS, MR3.7100) | Mean | 0.0176 | 3.7552 | 08/192 (0.0195 %IS, MR3.7100) | Mean | 0.0211 | 3.6873 | 08/192 (0.0195 %IS, MR3.7100) | Mean | 0.0159 | 3.8082 |
|  | Bias | 0.0019 | -0.0453 |  | Bias | -0.0016 | 0.0227 |  | Bias | 0.0036 | -0.0982 |
|  | SD | 0.0017 | 0.0429 |  | SD | 0.0057 | 0.1182 |  | SD | 0.0043 | 0.1160 |
|  | CV (%) | 9.6825 | 1.1434 |  | CV (%) | 27.0494 | 3.2044 |  | CV (%) | 26.7040 | 3.0460 |
| CF | 1.28 |  |  | CF | 1.06 |  |  | CF | 1.07 |  |  |

Evaluated with three independent lots (different from linearity lots).