Supplementary Table S1 Comparison of the proposed method with reported analytical methods for Tofacitinib.

| Method Reference | Optimization Strategy | Solvent System | Greenness / Sustainability | Remarks / Limitations |
| --- | --- | --- | --- | --- |
| UV Spectrophotometry (Gorantla et al., 2018) | OFAT | Methanol / Water | Not Assessed | Lacks stability-indicating capability; low sensitivity. |
| RP-HPLC (Sankar et al., 2017) | OFAT | Acetonitrile : Buffer | Low (Uses Acetonitrile) | Uses toxic solvents; no risk assessment or robustness testing. |
| LC-MS/MS (Wang et al., 2022) | OFAT | Acetonitrile : Formic Acid | Not Assessed | High instrument cost; impractical for routine QC. |
| RP-HPLC (Microchem. J., 2024) | OFAT | Acetonitrile : 0.2% Phosphoric Acid | Greenness Mentioned | Uses Acetonitrile (toxic); No WAC scoring. |
| Reported HPLC (Chromatographia, 2025) | OFAT | Acetonitrile : Buffer* | Not Assessed | [Verify specific solvent in PDF] |
| Proposed Method | A-QbD (Box-Behnken) | Methanol : 0.1% OPA | High (AGREE: 0.74, WAC: 0.81) | First application of White Analytical Chemistry with full stability profiling. |

Supplementary Table S2 Analytical Target Profile (ATP)

| ATP Element | Description | Target / Acceptance Criteria | Scientific & Regulatory Justification | Evaluation Method |
| --- | --- | --- | --- | --- |
| Analytical Purpose | Quantitative estimation of TF in tablet dosage form and bulk using RP-HPLC | Accurate, precise, robust, and environmentally compliant method | Defines the intended measurement function and analytical context per ICH Q14 | Method validation report |
| Performance Goal | Determine TF content with minimal variability | Accuracy: 98–102% recovery; Precision: %RSD ≤ 2%; Robustness: variation ≤ 2% | Establishes an acceptable uncertainty range | Validation of linearity, precision, and recovery |
| Specificity | Ability to separate TF from degradation and excipient peaks | No interference at TF RT | Ensures method selectivity (ICH Q2 (R2)) | Forced degradation & placebo blank tests |
| Sensitivity | Quantification at the therapeutic level | LOD ≤ 0.5 µg/mL; LOQ ≤ 1.5 µg/mL | Confirms suitability for low-dose tablet assay | Signal-to-noise ratio (3:1 and 10:1) |
| Precision | Reproducibility of peak area and RT | %RSD ≤ 2% intra/inter-day | Key for batch-to-batch uniformity | Repeatability studies |
| Accuracy | Closeness to true value | Recovery = 98–102% at 80, 100, 120% spike levels | Required for quantitative methods | Standard-addition method |
| Robustness | Resistance to deliberate variation | MeOH ± 1%, flow ± 0.05 mL/min, T ± 2 °C → RSD < 2% | Demonstrates stability under small perturbations | DOE-based robustness test |
| Greenness & SDG Alignment | Minimization of solvent and waste | Solvent use < 5 mL/run; run time < 6 min | Supports UN SDGs 12 & 13; aligns with “Analytical Greenness Metric” | AGREE metric or GAPI tool |

Supplementary Table S3  Quality Target Product Profile (QTPP)

| Attribute | Target / Criteria | Rationale (from text) | Potential Impact on ATP | Regulatory Reference |
| --- | --- | --- | --- | --- |
| Analyte | TF (label claim 5 mg/tablet) | Defines product assay requirement | High | ICH Q8(R2) |
| Dosage Form Matrix | Tablet powder extracted in methanol/0.1% OPA | Affects recovery and precision | High | WHO QC Guidelines |
| Analytical Method Type | RP-HPLC with UV DAD detection @ 286 nm | Provides selectivity and quantitative response | High | ICH Q2(R2) |
| Chromatographic Mode | Isocratic C18 column (250×4.6 mm, 5 µm) | Ensures reproducible retention and efficiency | Medium | Pharmacopoeial standard |
| Mobile Phase Composition | Methanol:0.1% OPA (60:40 v/v) | Optimized for resolution and symmetry | High | A-QbD optimization result |
| Flow Rate | 0.7 mL/min | Balances efficiency and pressure | High | DOE output |
| Column Temperature | 30 ± 2 °C | Maintains reproducibility and viscosity | Medium | Equipment specification |
| Injection Volume | 20 µL | Ensures detectability with minimal overload | Low | Standard HPLC practice |
| Run Time | ≤ 6 min | Short analysis time (eco-friendly) | Medium | SDG 12 compliance |
| Performance Goals | Resolution > 2.0; Tailing < 1.5; Plates > 8000 | Meet system suitability | High | ICH Q14 |
| Validation Parameters | Accuracy, precision, linearity, robustness | Verify method suitability | High | ICH Q2(R2) |
| Regulatory Compliance | ICH Q8 (R2), Q9, Q14 + UN SDGs | Ensures lifecycle management | High | EMA Guidance 2024 |

Supplementary Table S4  Critical Quality Attributes (CQAs)

| CQA | Definition | Target / Acceptance Criteria | Impact on Analytical Performance | Criticality Level | Measurement / Evaluation |
| --- | --- | --- | --- | --- | --- |
| Retention Time (Y₁) | Elution time of TF peak | 6.0 ± 0.3 min | Indicates selectivity and stability | High | Direct HPLC measurement |
| Tailing Factor (Y₂) | Symmetry of TF peak | ≤ 1.5 | Reflects column health & mobile-phase suitability | High | USP method <621> |
| Theoretical Plates (Y₃) | Column efficiency | ≥ 8000 | Evaluates efficiency & reproducibility | High | System-suitability test |
| Peak Area (Y₄) | Integrated response proportional to [TF] | RSD ≤ 2% | Determines quantitation accuracy | High | Linearity study |
| Resolution (Y₅) | Separation between TF & nearest impurity | ≥ 2.0 | Ensures specificity | Medium | Forced-degradation data |
| LOD / LOQ | Sensitivity parameters | LOD ≤ 0.5 µg/mL; LOQ ≤ 1.5 µg/mL | Confirms detection capability | Medium | Signal/noise 3:1 / 10:1 |
| Robustness Metric | RSD after ± variation tests | ≤ 2% change | Confirms method stability | Medium | DOE robustness test |

Supplementary Table S5 Critical Material Attributes (CMAs)

| Material / Component | Attribute | Target Specification | Effect on CQAs | Risk Level | Control / Mitigation Strategy |
| --- | --- | --- | --- | --- | --- |
| Stationary Phase (Column) | C18 (250 × 4.6 mm, 5 µm) | USP L1, inert silica | Affects Y₂ & Y₃ | High | Use same manufacturer; condition column |
| Mobile Phase – Methanol | Organic % | 60 ± 1 % | Impacts Y₁, Y₂ & Y₄ | High | Premix; verify ratio gravimetrically |
| Aqueous Phase (0.1% OPA) | pH ≈ 3.0 | Stable ± 0.1 | Affects peak shape (Y₂) & retention | Medium | Fresh daily; pH meter check |
| Solvent Quality | HPLC grade (99.9 %) | UV transparent | Influences baseline noise & LOD | Medium | Supplier qualification |
| Sample Matrix | Tablet excipients | Inert to mobile phase | May affect resolution (Y₅) | Medium | Validate sample prep protocol |
| Filtration Material | 0.45 µm membrane | Non-adsorptive PTFE | Prevents particulate interference | Low | Pre-rinse filters |
| Buffer Strength | 0.1 % w/v OPA | ± 5 % tolerance | Affects RT and plates | Medium | Check concentration gravimetrically |
| Lamp / Detector | Wavelength stability (286 ± 1 nm) | ± 1 nm | Affects sensitivity (Y₄) | Low | Regular calibration |

Supplementary Table S6 Critical Process Parameters (CPPs)

| Process Parameter | Nominal Value | Acceptable Range / Design Space | Affected CQAs | Risk Level (from FMEA) | Control Strategy |
| --- | --- | --- | --- | --- | --- |
| Flow Rate (mL/min) | 0.70 | 0.68 – 0.72 | Y₁, Y₃, Y₄ | High | Pump calibration; real-time monitoring |
| Methanol Proportion (%) | 60 | 58 – 62 | Y₁, Y₂, Y₄ | High | Premixed solvent; limit variation ±1 % |
| Column Temperature (°C) | 30 | 28 – 32 | Y₂, Y₃ | Medium | Oven PID control; temperature logging |
| Injection Volume (µL) | 20 | 18 – 22 | Y₄ | Medium | Autosampler calibration |
| Detection Wavelength (nm) | 286 | 285 – 287 | Y₄ | Low | Monthly lamp verification |
| Sample Sonication (min) | 15 | 10 – 20 | Y₄, Recovery | Medium | SOP control; time logging |
| Filtration Integrity | 0.45 µm membrane filter | Must pass clarity test | Y₂ & baseline | Low | Visual inspection per batch |
| Degassing Quality | Ultrasonic 5 min | No air bubbles | All CQAs | Medium | Pre-run degassing SOP |

Supplementary Table S7 BBD Design matrix for optimization Run (R1- R20)

| Run | A: Methanol (%) | B: Flow rate (mL/min) | C: Temp (°C) | Retention time (min) | Tailing factor | Plates (N) | Peak area (mAU·s) |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | 60 | 0.7 | 30 | 3.36 | 0.56 | 4702 | 375.8 |
| 2 | 60 | 0.7 | 25 | 3.60 | 0.62 | 4600 | 372.0 |
| 3 | 55 | 0.8 | 35 | 4.00 | 0.70 | 4200 | 360.5 |
| 4 | 60 | 0.7 | 35 | 3.10 | 0.55 | 4750 | 378.0 |
| 5 | 65 | 0.6 | 25 | 2.80 | 0.50 | 4400 | 370.0 |
| 6 | 65 | 0.8 | 35 | 2.60 | 0.53 | 4300 | 355.0 |
| 7 | 65 | 0.8 | 25 | 2.70 | 0.55 | 4250 | 360.0 |
| 8 | 60 | 0.7 | 30 | 3.36 | 0.56 | 4702 | 375.8 |
| 9 | 55 | 0.8 | 25 | 4.20 | 0.72 | 4100 | 355.0 |
| 10 | 65 | 0.6 | 35 | 2.70 | 0.52 | 4450 | 365.0 |
| 11 | 60 | 0.8 | 30 | 3.05 | 0.58 | 4550 | 362.0 |
| 12 | 60 | 0.7 | 30 | 3.36 | 0.56 | 4702 | 375.8 |
| 13 | 65 | 0.7 | 30 | 2.90 | 0.54 | 4600 | 368.0 |
| 14 | 55 | 0.6 | 25 | 4.40 | 0.75 | 4050 | 350.0 |
| 15 | 60 | 0.7 | 30 | 3.36 | 0.56 | 4702 | 375.8 |
| 16 | 55 | 0.7 | 30 | 4.10 | 0.70 | 4150 | 360.0 |
| 17 | 60 | 0.7 | 30 | 3.36 | 0.56 | 4702 | 375.8 |
| 18 | 60 | 0.6 | 30 | 3.80 | 0.60 | 4500 | 370.0 |
| 19 | 55 | 0.6 | 35 | 4.00 | 0.68 | 4200 | 355.0 |
| 20 | 60 | 0.7 | 30 | 3.36 | 0.56 | 4702 | 375.8 |

Supplementary Table S8 Risk Assessment Using FEMA Approaches

| Source of Variation | Potential Cause | Effect on CQAs (Y₁–Y₄) | Severity (1–10) | Occurrence (1–10) | Detectability (1–10) | RPN (=S×O×D) | Risk Level | Control Strategy |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Mobile Phase Composition (% MeOH) | Solvent ratio variation | Retention time ↑/↓; Peak shape distortion | 9 | 6 | 3 | 162 | High | Tight control (±1%), pre-mixed mobile phase |
| Flow Rate (mL/min) | Pump drift, air bubbles | Retention time shift, resolution loss | 8 | 5 | 4 | 160 | High | Calibrate pump daily, purge system |
| Column Temperature (°C) | Oven fluctuation | Viscosity, selectivity change | 7 | 4 | 5 | 140 | Medium | Maintain ±1°C, use preheated solvents |
| pH of Mobile Phase | OPA concentration error | Ionization, peak tailing | 6 | 4 | 4 | 96 | Medium | Check pH daily, freshly prepare buffer |
| Detector Wavelength (nm) | Lamp aging, λ error | Reduced sensitivity | 5 | 3 | 3 | 45 | Low | Periodic calibration |
| Sample Preparation | Incomplete dissolution, filtration error | Low recovery, variable area | 7 | 5 | 3 | 105 | Medium | Ultrasonicate, use 0.45 µm membrane |
| Injection Volume (µL) | Autosampler error | Peak area variability | 6 | 3 | 2 | 36 | Low | Autosampler validation |
| Column Age/Type | Stationary phase deterioration | Efficiency ↓ | 9 | 4 | 2 | 72 | Medium | Replace after 500 injections |

Supplementary Table S9 Integration of Risk Assessment and Design Space Control Strategy

| Parameter | Criticality (from FMEA) | Design Space Range | Acceptable Operating Range (Control Strategy) | Monitoring Method | Risk Mitigation |
| --- | --- | --- | --- | --- | --- |
| Methanol (%) | High | 58–62 | 60 ± 1 % | Ratio check before batch | Use pre-mixed batch solvent |
| Flow Rate (mL/min) | High | 0.68–0.72 | 0.7 ± 0.02 mL/min | Real-time pressure log | Pump auto-calibration |
| Column Temperature (°C) | Medium | 28–32 | 30 ± 2 °C | Digital probe | Oven PID controller |
| pH (Mobile Phase) | Medium | 2.8–3.2 | 3.0 ± 0.1 | pH meter | Fresh buffer daily |
| Wavelength (nm) | Low | 285–287 | 286 ± 1 nm | DAD check | Monthly lamp check |
| Sample Prep | Medium | — | Standardized | Visual clarity | Operator training |

Supplementary Table S 10 Preliminary trial runs

| Trial No. | Mobile Phase (v/v) | Flow Rate (mL min⁻¹) | Temperature (°C) | Observation / Result | Decision |
| --- | --- | --- | --- | --- | --- |
| 1 | MeOH : Water (50:50) | 1.0 | 25 | Peak eluted very late (> 10 min); broad peak. | Rejected |
| 2 | MeOH : Water (70:30) | 1.0 | 25 | Peak eluted very early (< 2 min); interference with solvent front. | Rejected |
| 3 | MeOH : 0.1% OPA (60:40) | 1.0 | 25 | Improved retention time; peak symmetry needs improvement. | Selected as Center Point |
| 4 | MeOH : 0.1% OPA (60:40) | 0.6 | 25 | Good separation; slight tailing. | Lower Level Selected |
| 5 | MeOH : 0.1% OPA (60:40) | 0.8 | 25 | Sharp peak; suitable pressure. | Upper Level Selected |
| 6 | MeOH : 0.1% OPA (60:40) | 0.7 | 35 | Temperature improved theoretical plates. | Included in Range |

Supplementary Table S 11 R², adjusted R², polynomial equation, p-values, and lack-of-fit.

| Response Variable | Regression Equation (Coded Factors) | R2 | Adj. R2 | Pred. R2 | Model p-value | Adeq. Precision |
| --- | --- | --- | --- | --- | --- | --- |
| Y1: Retention Time | Y1=3.376-0.697A-0.118B-0.132C-0.0013AB+0.0488AC+0.0263BC+0.0859A2+0.0109B2-0.0391C2 | 0.9998 | 0.9997 | 0.9987 | < 0.0001 (Sig) | 276.97 |
| Y2: Tailing Factor | Y2=0.568-0.091A+0.003B-0.016C+0.0088AB+0.0113AC+0.0013BC+0.0395A2+0.0095B2+0.0045C2 | 0.9610 | 0.9259 | 0.5920 | < 0.0001 (Sig) | 16.50 |
| Y3: Theoretical Plates | Y3=4691.42+130A-20B+50C-43.75AB-18.75AC-6.25BC-300.55A2-150.55B2-0.55C2 | 0.9707 | 0.9444 | 0.7758 | < 0.0001 (Sig) | 18.30 |
| Y4: Peak Area | Y4=375.013+ 3.75 A-1.75 B+ 0.65 -3.81 AB-2.5625 AC+0.0624C -9.83 A2-7.83 B2+1.16C² | 0.9740 | 0.9507 | 0.8283 | < 0.0001 (Sig) | 19.42 |