Supplementary Information

Calibration method for N2O5, ClNO2, and HONO with Iodide Chemical Ionization Mass Spectrometry (I-CIMS)

Jeongin Songa, Yoonseo Kima, Dabeen Junga, Jae-Bum Leeb, Yonghee Leeb, Seung-Hee Eunb, Song-Chul Hongb,*, Taehyoung Leea,*

a National Department of Environmental Science, Hankuk University of Foreign Studies, Yongin 17035, Republic of Korea

b Air Quality Forecasting Center, National Institute of Environmental Research, Incheon 22689, Republic of Korea

# S1. Atmospheric Pathways and Reaction Mechanisms of N₂O₅, ClNO₂, and HONO

Dinitrogen pentoxide (N₂O₅) is one of the key gas-phase precursors involved in nighttime nitrate formation. It is produced via the reaction of nitrogen dioxide (NO₂) with the nitrate radical (NO₃·) (R1, R2). Both NO₃· and N₂O₅ are rapidly photolyzed by sunlight, resulting in negligible concentrations during the daytime, however, they can persist in the atmosphere for short periods during the night in the absence of solar radiation. N₂O₅ reacts heterogeneously on particle surfaces to form either HNO₃ or ClNO₂. The rate of heterogeneous reactions is influenced by both relative humidity (RH) and the available particle surface area (Tham et al., 2018; Zhou et al., 2018; Wang et al., 2020). While increasing RH may enhance uptake by promoting aqueous layer formation, the relationship is nonlinear and can saturate or even decline under high-RH conditions. Similarly, larger aerosol surface area generally increases uptake and product formation. The HNO₃ formed via nighttime N₂O₅ hydrolysis may subsequently react with ammonia (NH₃) to produce nitrate (Chang et al., 2011; Yang et al., 2021; Zang et al., 2022).

| NO2+ O3 →NO3 + O2 | R1 |
| --- | --- |
| NO3+ NO2 ↔ N2O5 | R2 |
| N2O5(g)+ H2O(aq) → 2HNO3 | R3 |

Nitryl chloride (ClNO₂) is known to act as a reservoir for nitrogen oxides (NOₓ) and chlorine radicals (Cl·), thereby influencing ozone (O₃) formation and related photochemical processes (Li et al., 2021; Chen et al., 2024). In the atmosphere, ClNO₂ is formed through the heterogeneous reaction of N₂O₅ with particulate chloride (Cl⁻) on aerosol surfaces. Its production is affected by the concentration of N₂O₅, N₂O₅ uptake coefficient (γ(N₂O₅)), and concentration of Cl⁻. ClNO₂ is rapidly photolyzed after sunrise under solar irradiation, thereby leading to enhanced levels of NO₂ and Cl· during the early morning hours. Previous studies have reported that Cl· released from the photolysis of ClNO₂ significantly impact the atmospheric oxidation capacity, with consequential effects on O₃ formation and secondary organic aerosol (SOA) production (Ma et al., 2021; Dai et al., 2025; Wang et al., 2025).

| N2O5g+ Cl-aq→ ClNO2 + NO3-(aq) | R4 |
| --- | --- |
| ClNO2 +hv → NO2+ Cl∙ | R5 |
| Cl∙+ VOCs → SOA | R6 |

Nitrous acid (HONO) is a major precursor of the hydroxyl radical (OH·), which is a key oxidant in the chemical reactions of organic and inorganic compounds, and it reacts with volatile organic compounds (VOCs) to affect the formation of O₃ and PM₂.₅ (Chen et al., 2023). HONO can be formed in the atmosphere via the reaction between nitric oxide (NO) and OH·, but it is also directly emitted from a variety of sources including vehicle exhaust, soil emissions, and the photolysis of nitrate-containing compounds. Due to its importance, the formation mechanisms of HONO remain an active area of research (Liu et al., 2017; Xue, 2022; Andersen et al., 2023).

S2. Iodide Chemical Ionization Mass Spectrometry (I-CIMS)

A CH₃I permeation tube (CH₃I + C₆H₆) was used to generate iodide ions (I⁻), and the permeation tube was heated to approximately 80 °C. Approximately 250 sccm of high-purity nitrogen gas (N₂, 99.99%) was passed through the permeation tube to generate CH₃I + C₆H₆ gas, which was then introduced into the reactor equipped with a vacuum ultraviolet (VUV) lamp. A krypton lamp, which emits at 116.486 nm and 123.582 nm, was used as the VUV source. When the benzene containing CH₃I absorbs VUV light, CH₃I reacts with photoelectrons to produce I⁻. The sample gas was introduced into the reactor at a flow rate of approximately 2000 sccm, and the pressure inside the reactor was maintained at approximately 50–55 mbar. Acetonitrile (ACN) was introduced (10~15 sccm) at the front of the reactor as a dopant to promote the formation of I(ACN)⁻ reagent ions instead of I(H₂O)ₙ⁻ clusters. This minimizes the humidity dependence of the ion–molecule reactions and improves signal reproducibility (Riva et al., 2024). The acetonitrile was mixed with the sample and introduced into the reactor, where it reacted with I⁻ ions generated by the VUV lamp and was subsequently transferred to a standard time-of-flight (ToF) mass spectrometer chamber (Tofwerk AG, Switzerland) for ion detection and analysis.

S3. The quality control and quality assurance (QC/QA) of I-CIMS

The quality control and quality assurance (QC/QA) of I-CIMS measurements were conducted to determine the minimum detection limit (MDL), relative standard deviation (RSD)(EPA, 2017, App. B; Skoog et al., 2017). The MDL was calculated using blank sample data and the t-value at a 95% confidence level (Eq. S1). The blank samples used for MDL were obtained by introducing high-purity N2​ gas (99.99%).

In Eq. S1, N₁ denotes the number of independent blank runs, and N₂ denotes the number of replicate blank analyses per run (e.g., pre- and post-blank); within each replicate, time-series points were averaged to a single value to preserve independence. In Eq. S2, Sb is defined as the effective standard deviation of the blanks obtained by combining independent variance components.

|  | MDL ≥ t × Sb×N1+N2N1×N2 |  | Eq. S1 |
| --- | --- | --- | --- |
|  |  | t = value of t for 95% confidence level Sb = Standard deviation of blanks N1 = The number of independent blank runs N2 = The number of replicate blank analyses per run |  |

|  | Sb=S12+S22+⋯ | Eq. S2 |  |
| --- | --- | --- | --- |
|  | Si = Standard deviation of each independent variance component |  |

|  | S=(xi-x)2N-1 |  |  | Eq. S3 |  |
| --- | --- | --- | --- | --- | --- |
|  |  | N = The number of data points xi = Each of the values of the data x = The mean of xi |  |  |

To evaluate the analytical precision of I-CIMS, the relative standard deviation (RSD) was calculated using Eq. S3.

|  | RSD= xi-x2N-1x ×100% |  | Eq. S4 |
| --- | --- | --- | --- |
|  |  | N = The number of data points xi = Each of the values of the data x = The mean of xi |  |

Since no certified standard gases exist for N₂O₅, ClNO₂, or HONO, the absolute calibration accuracy cannot be directly evaluated against a reference standard. Therefore, in this study, the calibration uncertainty was estimated by considering the accuracy of the NO₂ (T500U) or NOₓ (T200) analyzers used to generate the calibration precursors, together with the dilution uncertainty of the mass flow controllers (MFCs). The combined input uncertainty, including both factors, was determined to be 4.5~4.7% (1σ) and used to derive the expanded uncertainties of the reported sensitivities (k = 2, 95% confidence level). Wall losses were not explicitly accounted for; to minimize their impact, the sampling line was kept as short and straight as possible. Consequently, the reported uncertainties primarily reflect the contributions from dilution and reference monitor accuracy, along with the regression statistics of the sensitivity fits.

|  | δinput= δdil2+δanalyzer2 u(S)= ​(δinput)2+SE(S)2 U=k*u(S) |  | Eq. S4 |
| --- | --- | --- | --- |
|  |  | δdil = uncertainty of the dilution ratio δanalyzer = uncertainty of the Gas monitor u (S)= Standard uncertainty SES= the standard error of the slope in calibration result U = Expanded uncertainty ( k=2, 95% confidence level) |  |

The Table S1 summarizes the QC/QA results for each compound measured by I-CIMS.

Table S1. Quality control and quality assurance (QC/QA) for I-CIMS

|  | N2O5* | ClNO2* | HONO** |
| --- | --- | --- | --- |
| MDL | 4.0 | 2.6 | 0.3 |
| RSD (%) | 4.2 | 4.5 | 0.9 |
| Expanded Uncertainty (%) | 9.4 | 10.7 | 10.2 |
| * : ppt, **: ppb |  |  |  |

## S4. Gas monitors

A NO₂ analyzer (Teledyne T500U) was used to monitor changes in NO₂ concentration during the calibration of N₂O₅ and ClNO₂. In addition, an O₃ analyzer (Thermo 49C) was installed to observe variations in O₃ concentration. The NO₂ analyzer operates based on the cavity-attenuated phase-shift (CAPS) method and measures NO₂ concentration using the absorbance of 450 nm light-corresponding to the NO₂ absorption range-based on the Beer-Lambert law and the optical path length. The NO₂ analyzer was calibrated by diluting NO₂ standard gas (9.72 ppm, gas cylinder) with zero air. The O₃ analyzer (Thermo 49C) is an ultraviolet (UV) photometric analyzer that determines O₃ concentration by measuring the absorption of UV light at 254 nm by O₃ in the sample gas. Calibration of the O₃ analyzer was performed using a 146C Dynamic Gas Calibrator (Thermo Scientific, U.S.) to generate O₃.

A NOₓ analyzer (Teledyne T200, Molybdenum converter) was employed to track variations in NOₓ concentration during the calibration of HONO. The NOₓ analyzer measures NO, NO₂, and NOₓ based on the chemiluminescence method. For the calibration of NO and NO₂ measured by the NOₓ analyzer, NO standard gas (10.1 ppm, gas cylinder) and NO₂ standard gas (9.72 ppm, gas cylinder) were diluted with zero air.

S5. N2O5 Calibration

In this study, to quantify N₂O₅ measured by I-CIMS, N₂O₅ was generated by reacting NO₂ and O₃, following the method described by Liu et al. (2024). Generally, N₂O₅ crystals can be produced by reacting excess NOₓ and O₃, or by reacting P₂O₅ with fuming nitric acid to obtain high-concentration N₂O₅ crystals (Bertram et al., 2009; Liu et al., 2024). However, N₂O₅ crystals obtained through these methods require multiple purification steps to achieve high purity, thereby rendering the process operationally complex, especially when targeting ppt-level concentrations through dilution. Furthermore, the dilution process from high-concentration N₂O₅ increases the uncertainty in generating N₂O₅ at specific ppt levels (Wang et al., 2016).

To verify the stoichiometric relationship between NO₂ consumption and N₂O₅ formation, experiments were conducted under O₃ levels (47.1, 49.1, 59.4, 90.1, 137.6 and 202.5 ppb) while maintaining a constant NO₂ concentration (9.72 ppm) and a residence time of approximately 74 s in the PFA reactor. As summarized in Table S2 and Table S3, Δ[NO₂] and [N₂O₅] exhibited an linear correlation (slope = 0.5, R² = 0.99), consistent with the stoichiometric conversion 2NO₂ → N₂O₅ (R1–R2). This result confirms that the contribution of the NO₃ self-reaction (R3) to the overall NO₂ loss was negligible under the excess-NO₂ conditions used in this study.

| O3+ NO2 → NO3+O2 | R1 |
| --- | --- |
| NO3+ NO2 → N2O5 | R2 |
| NO3+ NO3 → 2NO2+O2 | R3 |

Table S2. Experimental conditions for N₂O₅ generation under different O₃ concentrations

| O3 before mixing (ppb) | O3 after mixing (ppb) | NO₂ after mixing (ppb) | [O₃]/[NO₂] ratio | Residence time (s) |  |
| --- | --- | --- | --- | --- | --- |
| 53.3 | 49.1 | 773 | 0.06 | 74 |  |
| 97.8 | 90.1 | 773 | 0.12 | 74 |  |
| 220 | 202.5 | 773 | 0.26 | 74 |  |
| Note: NO₂ concentration was held constant (9.72 ppm), and only O₃ was varied to change the [O₃]/[NO₂] ratio. |  |  |  |  |

Table S3. Relationship between the decrease in NO₂ (ΔNO₂) and the produced N₂O₅ concentration

|  | Δ[NO₂] (ppb). | [N2O5] (ppb) | Slope ([N2O5]/ Δ[NO₂]) |
| --- | --- | --- | --- |
| 1 | 19.4 | 9.35 | 0.48 |
| 2 | 126.5 | 66.1 | 0.52 |
| 3 | 283.8 | 142.5 | 0.50 |
| Note: The slope (~ 0.5) and R² ≈ 0.99 confirm a 2 : 1 relation between ΔNO₂ and N₂O₅ formation. |  |  |  |

Fig.S1. Relationship between the decrease in NO₂ (ΔNO₂) and the produced N₂O₅. Red circles - experimental data; red dashed line - theoretical 0.5 × ΔNO₂ relation; black line - linear fit (slope = 0.50, R² = 0.99).

Table S4. Summary of N2O5 concentration in previous studies

| Location | Site type | Period | N2O5 (ppt) | Reference |
| --- | --- | --- | --- | --- |
| Seoul, Korea | Urban | 2015.05 | 4100 ± 1200 2600 ± 1600 | Brown et al., 2017 |
| Jiangsu, China | Suburban | 2018.05 | 26.0 ± 35.7 | Li et al., 2020 |
| Jinan, China | Urban | 2014.09 | 22 ± 12 (278*) | Wang et al., 2017 |
| South China Sea | Island | 2021.11 | 119.5 ± 128.6** | Wang et al., 2024 |
| Shenzhen, China | Coastal Background | 2019.09 | 55.6 ± 89 (1420*) | Niu et al., 2022 |
| Changzhou, China | Suburban | 2019.05 | 61.0 ± 63.1 (477.2*) | Zhai et al., 2023 |
| *: The highest concentration of N2O5, **: The nighttime average concentration of N2O5 |  |  |  |  |

S6. ClNO2 Calibration

Two representative methods for ClNO₂ generation have been established in previous studies: (1) the second relies on the reaction between N₂O₅ and a NaCl slurry and, (2) the first utilizes Cl₂ gas introduced into a NaCl and NaNO₂ mixture to initiate reaction (R4), (Wang et al., 2016; Xia et al., 2020; Lou et al., 2022; Chen et al., 2024).  In this study, we adopted approach (1), in which N₂O₅ was reacted with NaCl slurry.

In the Cl₂ based method, a slurry prepared by mixing NaCl and NaNO₂ at a molar ratio of approximately 10:1 is loaded along the entire length of the reaction chamber, and Cl₂ gas is then introduced. Through R4, the NO₂⁻ generated in the NaCl–NaNO₂ slurry reacts with Cl2 to synthesize ClNO₂ (Thaler et al., 2011).

| Cl2g+ NO2-aq→ClNO2g+Cl-(aq) | R4 |
| --- | --- |

Table S5. Summary of ClNO2 concentration in previous studies

| Location | Site type | Period | ClNO2 (ppt) | Reference |
| --- | --- | --- | --- | --- |
| Nanjing, China | Suburban | 2018.04 | 3700* | Xia et al., 2020 |
| Jinan, China | Urban | 2014.09 | 94 ± 58 | Wang et al., 2017 |
| Beijing, China | Urban | 2017.06 | 174.3 ± 262 | Zhou et al., 2018 |
| Beijing, China | Semi-rural | 2016.05 | 487** (2900*) | Breton et al., 2018 |
| Shenzhen, China | Background | 2019.10 | 247 ± 264.3** | Niu et al., 2022 |
| *: The highest concentration of ClNO2, **: The nighttime average concentration of ClNO2 |  |  |  |  |

S7. Sensitivity

Table S6. Sensitivitiy for I-CIMS in this study

| Compound | Sensitivity (cps/ppt) | R2 |
| --- | --- | --- |
| N2O5 | 1.1 | 0.99 |
| ClNO2 | 0.31 | 0.98 |
| HONO | 0.665 | 0.99 |

While previous studies have reported that N₂O₅ may exhibit a sensitivity up to an order of magnitude higher than ClNO₂ under ideal collision-limit conditions (Lee et al., 2014; Lopez-Hilfiker et al., 2016), the smaller difference observed in this study (~3.5×) is likely due to instrument-specific factors such as ion-transfer efficiency, and declustering voltages, as discussed by Aggarwal et al. (2025). These parameters can redistribute the detected ion signal without contradicting the kinetic limit.

References

Aggarwal, S., Bansal, P., Wang, Y., Jorga, S., Macgregor, G., Rohner, U., Bannan, T., Salter, M., Zieger, P., Mohr, C., Lopez-Hilfiker, F., 2025. Identifying key parameters that affect sensitivity of flow tube chemical ionization mass spectrometers. Atmos. Meas. Tech. 18, 4227-4247.

Andersen, S.T., Carpenter, L.J., Reed, C., Lee, J.D., Chance, R., Sherwen, T., Vaughan, A.R., Stewart, J., Edwards, P.M., Bloss, W.J., Sommariva, R., Crilley, L.R., Nott, G.J., Neves, L., Read, K., Heard, D.E., Seakins, P.W., Whalley, L.K., Boustead, G.A., Fleming, L.T., Stone, D. and Fomba, K.W. (2023). Extensive Field Evidence for the Release of Hono from the Photolysis of Nitrate Aerosols. Science Advances 9: eadd6266.

Bertram, T.H., Thornton, J.A., Riedel, T.P., Middlebrook, A.M., Bahreini, R., Bates, T.S., Quinn, P.K. and Coffman, D.J. (2009). Direct Observations of N2o5 Reactivity on Ambient Aerosol Particles. Geophysical Research Letters 36.

Chang, W.L., Bhave, P.V., Brown, S.S., Riemer, N., Stutz, J. and Dabdub, D. (2011). Heterogeneous Atmospheric Chemistry, Ambient Measurements, and Model Calculations of N2o5: A Review. Aerosol Science and Technology 45: 665-695.

Chen, D., Zhou, L., Liu, S., Lian, C., Wang, W., Liu, H., Li, C., Liu, Y., Luo, L., Xiao, K., Chen, Y., Qiu, Y., Tan, Q., Ge, M. and Yang, F. (2023). Primary Sources of Hono Vary During the Daytime: Insights Based on a Field Campaign. Science of The Total Environment 903: 166605.

Chen, G., Xu, L., Yu, S., Xue, L., Lin, Z., Yang, C., Ji, X., Fan, X., Tham, Y.J., Wang, H., Hong, Y., Li, M., Seinfeld, J.H. and Chen, J. (2024). Increasing Contribution of Chlorine Chemistry to Wintertime Ozone Formation Promoted by Enhanced Nitrogen Chemistry. Environmental Science & Technology 58: 22714-22721.

Chen, X., Xia, M., Wang, W., Yun, H., Yue, D. and Wang, T. (2023). Fast near-Surface Clno2 Production and Its Impact on O3 Formation During a Heavy Pollution Event in South China. Science of The Total Environment 858: 159998.

Dai, J., Wang, T., Shen, H., Xia, M., Sun, W. and Brasseur, G.P. (2025). Significant Impact of a Daytime Halogen Oxidant on Coastal Air Quality. Environmental Science & Technology 59: 2169-2180.

U.S. Environmental Protection Agency (U.S. EPA). Definition and Procedure for the Determination of the Method Detection Limit (MDL); 40 CFR Part 136, Appendix B (Revision 2); 2017.

Lao, M., Crilley, L.R., Salehpoor, L., Furlani, T.C., Bourgeois, I., Neuman, J.A., Rollins, A.W., Veres, P.R., Washenfelder, R.A., Womack, C.C., Young, C.J. and VandenBoer, T.C. (2020). A Portable, Robust, Stable, and Tunable Calibration Source for Gas-Phase Nitrous Acid (Hono). Atmos. Meas. Tech. 13: 5873-5890.

Le Breton, M., Hallquist, Å.M., Pathak, R.K., Simpson, D., Wang, Y., Johansson, J., Zheng, J., Yang, Y., Shang, D., Wang, H., Liu, Q., Chan, C., Wang, T., Bannan, T.J., Priestley, M., Percival, C.J., Shallcross, D.E., Lu, K., Guo, S., Hu, M. and Hallquist, M. (2018). Chlorine Oxidation of Vocs at a Semi-Rural Site in Beijing: Significant Chlorine Liberation from Clno2 and Subsequent Gas- and Particle-Phase Cl–Voc Production. Atmos. Chem. Phys. 18: 13013-13030.

Lee, B.H., Lopez-Hilfiker, F.D., Mohr, C., Kurtén, T., Worsnop, D.R. and Thornton, J.A. (2014). An Iodide-Adduct High-Resolution Time-of-Flight Chemical-Ionization Mass Spectrometer: Application to Atmospheric Inorganic and Organic Compounds. Environmental Science & Technology 48: 6309-6317.

Li, Q., Fu, X., Peng, X., Wang, W., Badia, A., Fernandez, R.P., Cuevas, C.A., Mu, Y., Chen, J., Jimenez, J.L., Wang, T. and Saiz-Lopez, A. (2021). Halogens Enhance Haze Pollution in China. Environmental Science & Technology 55: 13625-13637.

Li, Z., Xie, P., Hu, R., Wang, D., Jin, H., Chen, H., Lin, C. and Liu, W. (2020). Observations of N2o5 and No3 at a Suburban Environment in Yangtze River Delta in China: Estimating Heterogeneous N2o5 Uptake Coefficients. Journal of Environmental Sciences 95: 248-255.

Lindsay, A.J. and Wood, E.C. (2022). Comparison of Two Photolytic Calibration Methods for Nitrous Acid. Atmos. Meas. Tech. 15: 5455-5464.

Liu, Y., Jia, Y., Chu, B., Li, S., Cao, Q., Liu, J., Ma, W., Li, Y., Wang, L., Nie, W., Ma, Q. and He, H. (2024). An Alternative Calibration Method for Measuring N2o5 with an Iodide-Chemical Ionization Mass Spectrometer and Influencing Factors. Analytical Chemistry 96: 4048-4056.

Liu, Y., Lu, K., Ma, Y., Yang, X., Zhang, W., Wu, Y., Peng, J., Shuai, S., Hu, M. and Zhang, Y. (2017). Direct Emission of Nitrous Acid (Hono) from Gasoline Cars in China Determined by Vehicle Chassis Dynamometer Experiments. Atmospheric Environment 169: 89-96.

Lopez-Hilfiker, F.D., Iyer, S., Mohr, C., Lee, B.H., D'Ambro, E.L., Kurtén, T. and Thornton, J.A. (2016). Constraining the Sensitivity of Iodide Adduct Chemical Ionization Mass Spectrometry to Multifunctional Organic Molecules Using the Collision Limit and Thermodynamic Stability of Iodide Ion Adducts. Atmos. Meas. Tech. 9: 1505-1512.

Lou, S., Tan, Z., Gan, G., Chen, J., Wang, H., Gao, Y., Huang, D., Huang, C., Li, X., Song, R., Wang, H., Wang, M., Wang, Q., Wu, Y. and Huang, C. (2022). Observation Based Study on Atmospheric Oxidation Capacity in Shanghai During Late-Autumn: Contribution from Nitryl Chloride. Atmospheric Environment 271: 118902.

Niu, Y.-B., Zhu, B., He, L.-Y., Wang, Z., Lin, X.-Y., Tang, M.-X. and Huang, X.-F. (2022). Fast Nocturnal Heterogeneous Chemistry in a Coastal Background Atmosphere and Its Implications for Daytime Photochemistry. Journal of Geophysical Research: Atmospheres 127: e2022JD036716.

Riva, M., Pospisilova, V., Frege, C., Perrier, S., Bansal, P., Jorga, S., Sturm, P., Thornton, J.A., Rohner, U. and Lopez-Hilfiker, F. (2024). Evaluation of a Reduced-Pressure Chemical Ion Reactor Utilizing Adduct Ionization for the Detection of Gaseous Organic and Inorganic Species. Atmos. Meas. Tech. 17: 5887-5901.

Skoog, D.A.; Holler, F.J.; Crouch, S.R. Principles of Instrumental Analysis, 7th ed.; Cengage Learning: Boston, 2017; ISBN 978-1305577213.

Song, Y., Zhang, Y., Xue, C., Liu, P., He, X., Li, X. and Mu, Y. (2022). The Seasonal Variations and Potential Sources of Nitrous Acid (Hono) in the Rural North China Plain. Environmental Pollution 311: 119967.

Thaler, R.D., Mielke, L.H. and Osthoff, H.D. (2011). Quantification of Nitryl Chloride at Part Per Trillion Mixing Ratios by Thermal Dissociation Cavity Ring-Down Spectroscopy. Analytical Chemistry 83: 2761-2766.

Tham, Y.J., Wang, Z., Li, Q., Wang, W., Wang, X., Lu, K., Ma, N., Yan, C., Kecorius, S., Wiedensohler, A., Zhang, Y. and Wang, T. (2018). Heterogeneous N2o5 Uptake Coefficient and Production Yield of Clno2 in Polluted Northern China: Roles of Aerosol Water Content and Chemical Composition. Atmos. Chem. Phys. 18: 13155-13171.

ang, H., Chen, X., Lu, K., Tan, Z., Ma, X., Wu, Z., Li, X., Liu, Y., Shang, D., Wu, Y., Zeng, L., Hu, M., Schmitt, S., Kiendler-Scharr, A., Wahner, A. and Zhang, Y. (2020). Wintertime N2o5 Uptake Coefficients over the North China Plain. Science Bulletin 65: 765-774.

Wang, H., Yuan, B., Zhang, X., Wang, J., Chen, X., Wang, Y., Qin, Y., Li, X.-B., Zhang, C., Liu, A., Lu, K., Zheng, E., Li, L., Yang, L., Zhou, J., Song, X., Huangfu, Y., Wang, X. and Shao, M. (2025). Vertical Gradient of Nitryl Chloride and Implications for Atmospheric Photochemistry in Pearl River Delta, China, During Wintertime. Environmental Science & Technology 59: 2144-2156.

Wang, J., Wang, H., Tham, Y.J., Ming, L., Zheng, Z., Fang, G., Sun, C., Ling, Z., Zhao, J. and Fan, S. (2024). Measurement Report: Atmospheric Nitrate Radical Chemistry in the South China Sea Influenced by the Urban Outflow of the Pearl River Delta. Atmos. Chem. Phys. 24: 977-992.

Wang, T., Tham, Y.J., Xue, L., Li, Q., Zha, Q., Wang, Z., Poon, S.C.N., Dubé, W.P., Blake, D.R., Louie, P.K.K., Luk, C.W.Y., Tsui, W. and Brown, S.S. (2016). Observations of Nitryl Chloride and Modeling Its Source and Effect on Ozone in the Planetary Boundary Layer of Southern China. Journal of Geophysical Research: Atmospheres 121: 2476-2489.

Wang, X., Wang, H., Xue, L., Wang, T., Wang, L., Gu, R., Wang, W., Tham, Y.J., Wang, Z., Yang, L., Chen, J. and Wang, W. (2017). Observations of N2o5 and Clno2 at a Polluted Urban Surface Site in North China: High N2o5 Uptake Coefficients and Low Clno2 Product Yields. Atmospheric Environment 156: 125-134.

Xia, M., Peng, X., Wang, W., Yu, C., Sun, P., Li, Y., Liu, Y., Xu, Z., Wang, Z., Xu, Z., Nie, W., Ding, A. and Wang, T. (2020). Significant Production of Clno2 and Possible Source of Cl2 from N2o5 Uptake at a Suburban Site in Eastern China. Atmos. Chem. Phys. 20: 6147-6158.

Xue, C. (2022). Substantially Growing Interest in the Chemistry of Nitrous Acid (Hono) in China: Current Achievements, Problems, and Future Directions. Environmental Science & Technology 56: 7375-7377.

Yang, S., Yuan, B., Peng, Y., Huang, S., Chen, W., Hu, W., Pei, C., Zhou, J., Parrish, D.D., Wang, W., He, X., Cheng, C., Li, X.B., Yang, X., Song, Y., Wang, H., Qi, J., Wang, B., Wang, C., Wang, C., Wang, Z., Li, T., Zheng, E., Wang, S., Wu, C., Cai, M., Ye, C., Song, W., Cheng, P., Chen, D., Wang, X., Zhang, Z., Wang, X., Zheng, J. and Shao, M. (2022). The Formation and Mitigation of Nitrate Pollution: Comparison between Urban and Suburban Environments. Atmos. Chem. Phys. 22: 4539-4556.

Zang, H., Zhao, Y., Huo, J., Zhao, Q., Fu, Q., Duan, Y., Shao, J., Huang, C., An, J., Xue, L., Li, Z., Li, C. and Xiao, H. (2022). High Atmospheric Oxidation Capacity Drives Wintertime Nitrate Pollution in the Eastern Yangtze River Delta of China. Atmos. Chem. Phys. 22: 4355-4374.

Zhai, T., Lu, K., Wang, H., Lou, S., Chen, X., Hu, R. and Zhang, Y. (2023). Elucidate the Formation Mechanism of Particulate Nitrate Based on Direct Radical Observations in the Yangtze River Delta Summer 2019. Atmos. Chem. Phys. 23: 2379-2391.

Zhou, W., Zhao, J., Ouyang, B., Mehra, A., Xu, W., Wang, Y., Bannan, T.J., Worrall, S.D., Priestley, M., Bacak, A., Chen, Q., Xie, C., Wang, Q., Wang, J., Du, W., Zhang, Y., Ge, X., Ye, P., Lee, J.D., Fu, P., Wang, Z., Worsnop, D., Jones, R., Percival, C.J., Coe, H. and Sun, Y. (2018). Production of N2o5 and Clno2 in Summer in Urban Beijing, China. Atmos. Chem. Phys. 18: 11581-11597.