Table 1S

MRM parameters for LC–MS/MS analysis of the 20 marker components in HYT

| Compound | Mode | MW (m/z) | Transition (m/z) | Collision energy (eV) | Cone voltage (V) |
| --- | --- | --- | --- | --- | --- |
| Gallic acid | Negative | 170.12 | 169 → 125 | 15 | 25 |
| Cimifugin | Positive | 306.31 | 307 → 259 | 30 | 40 |
| Geniposide | Negative | 388.37 | 387 → 123 | 20 | 55 |
| Paeoniflorin | Negative | 480.46 | 479 → 121 | 25 | 60 |
| Ferulic acid | Positive | 194.18 | 195 → 177 | 10 | 30 |
| Nodakenin | Positive | 408.40 | 409 → 247 | 10 | 30 |
| Narirutin | Negative | 580.54 | 579 → 271 | 30 | 45 |
| Naringin | Negative | 580.53 | 579 → 271 | 20 | 60 |
| Neohesperidin | Positive | 610.56 | 611 → 303 | 20 | 15 |
| Baicalin | Positive | 446.36 | 447 → 271 | 15 | 25 |
| Arctiin | Positive | 534.56 | 557 → 395 | 30 | 25 |
| Oxypeucedanin hydrate | Positive | 304.29 | 305 → 203 | 25 | 45 |
| Wogonoside | Positive | 460.39 | 461 → 285 | 20 | 30 |
| Baicalein | Positive | 270.24 | 271 → 123 | 30 | 40 |
| Arctigenin | Positive | 372.42 | 373 → 137 | 25 | 55 |
| Glycyrrhizin | Negative | 822.93 | 821 → 351 | 40 | 45 |
| Wogonin | Positive | 284.26 | 285 → 270 | 20 | 40 |
| Pulegone | Positive | 152.24 | 153 → 67 | 20 | 30 |
| Decursin | Positive | 328.36 | 329 → 229 | 20 | 35 |
| Decursinol angelate | Positive | 328.36 | 329 → 229 | 20 | 35 |

Table 2S

HPLC system suitability for the 20 marker compounds

| Compound | k |  | N | Rs | Tf |
| --- | --- | --- | --- | --- | --- |
| Gallic acid | 1.20 | 3.47 | 2124.64 | 33.70 | 1.00 |
| Cimifugin | 4.17 | 1.14 | 25661.64 | 8.93 | 1.10 |
| Geniposide | 4.76 | 1.18 | 34039.23 | 8.93 | 1.05 |
| Paeoniflorin | 5.62 | 1.19 | 20710.07 | 14.57 | 1.14 |
| Ferulic acid | 6.69 | 1.02 | 53930.39 | 2.28 | 1.03 |
| Nodakenin | 6.85 | 1.02 | 59577.87 | 2.28 | 1.04 |
| Narirutin | 7.00 | 1.04 | 73577.36 | 2.28 | 1.05 |
| Naringin | 7.30 | 1.07 | 59377.97 | 7.60 | 1.05 |
| Neohesperidin | 7.81 | 1.18 | 84119.77 | 20.10 | 1.04 |
| Arctiin | 9.19 | 1.03 | 89441.88 | 3.23 | 1.04 |
| Baicalin | 9.43 | 1.03 | 63495.60 | 3.23 | 1.04 |
| Oxypeucedanin hydrate | 9.75 | 1.13 | 77398.22 | 15.58 | 1.03 |
| Wogonoside | 11.02 | 1.08 | 84201.53 | 10.68 | 1.07 |
| Baicalein | 11.94 | 1.08 | 106802.23 | 10.37 | 1.09 |
| Arctigenin | 12.87 | 1.09 | 75025.60 | 12.73 | 1.03 |
| Glycyrrhizin | 14.01 | 1.02 | 120135.43 | 3.47 | 1.20 |
| Wogonin | 14.33 | 1.12 | 102533.81 | 3.47 | 1.02 |
| Pulegone | 16.04 | 1.14 | 92080.08 | 20.26 | 1.01 |
| Decursin | 18.22 | 1.01 | 247399.65 | 1.59 | 1.01 |
| Decursinol angelate | 18.37 | 1.01 | 263946.70 | 1.59 | 1.02 |

Table 3S

The linear range, regression equation, r2, LODs, and LOQs of the analytes from HYT (n = 3) obtained by LC–MS/MS analysis.

| Compound | Retention time (min) | Linear range (ng/mL) | Regression equationa | r2 | LODb (ng/mL) | LOQc (ng/mL) |
| --- | --- | --- | --- | --- | --- | --- |
| Gallic acid | 1.04 | 10.00–500.00 | y = 76.96x – 798.40 | 0.9967 | 0.266 | 0.797 |
| Cimifugin | 5.02 | 10.00–10000.00 | y = 4250.60x + 76625.09 | 0.9940 | 0.019 | 0.057 |
| Geniposide | 3.59 | 10.00–10000.00 | y = 8.53x – 154.55 | 0.9991 | 0.576 | 1.728 |
| Paeoniflorin | 4.22 | 10.00–10000.00 | y = 60.03x – 653.71 | 0.9981 | 0.259 | 0.778 |
| Ferulic acid | 4.65 | 10.00–10000.00 | y = 408.37x + 11427.29 | 0.9934 | 1.135 | 3.404 |
| Nodakenin | 5.03 | 10.00–10000.00 | y = 13778.60x + 342858.23 | 0.9905 | 0.004 | 0.012 |
| Narirutin | 5.12 | 10.00–10000.00 | y = 404.66x – 13673.67 | 0.9929 | 0.501 | 1.502 |
| Naringin | 5.3 | 10.00–10000.00 | y = 75.10x – 1828.87 | 0.9977 | 0.192 | 0.577 |
| Neohesperidin | 5.61 | 10.00–10000.00 | y = 821.88x + 658.91 | 0.9995 | 0.019 | 0.057 |
| Arctiin | 5.78 | 10.00–10000.00 | y = 4208.78x + 7604.55 | 0.9998 | 0.034 | 0.101 |
| Baicalin | 6.45 | 10.00–10000.00 | y = 112.19x + 2325.43 | 0.9948 | 0.001 | 0.003 |
| Oxypeucedanin hydrate | 6.58 | 10.00–10000.00 | y = 12060.16x + 452546.31 | 0.9916 | 0.003 | 0.008 |
| Wogonoside | 6.68 | 10.00–10000.00 | y = 8798.83x + 68721.40 | 0.9991 | 0.007 | 0.020 |
| Baicalein | 7.84 | 10.00–10000.00 | y = 6100.64x – 120429.78 | 0.9987 | 0.029 | 0.087 |
| Arctigenin | 8.2 | 10.00–10000.00 | y = 23944.37x + 777382.03 | 0.9927 | 0.002 | 0.005 |
| Glycyrrhizin | 8.1 | 10.00–10000.00 | y = 8.88x + 103.33 | 0.9988 | 0.710 | 2.131 |
| Wogonin | 8.72 | 10.00–10000.00 | y = 15826.68x + 211206.40 | 0.9956 | 0.009 | 0.028 |
| Pulegone | 9.38 | 10.00–10000.00 | y = 112.61x + 1244.78 | 0.9997 | 0.940 | 2.819 |
| Decursin | 9.91 | 10.00–500.00 | y =70357.59x + 1982842.86 | 0.9910 | 0.001 | 0.003 |
| Decursinol angelate | 9.96 | 10.00–500.00 | y =15241.24x + 343092.33 | 0.9968 | 0.001 | 0.003 |

aRegression equation expressed y = ax + b, where y and x are peak area and concentration of compound, respectively

bLOD = 3.3  signal-to-noise ratio

cLOQ = 10  signal-to-noise ratio

Fig. 1S. Q1 (A) and Q3 (B) mass spectra of the 20 marker components. Gallic acid (1), cimifugin (2), geniposide (3), paeoniflorin (4), ferulic acid (5), nodakenin (6), narirutin (7), naringin (8), neohesperidin (9), arctiin (10), baicalin (11), oxypeucedanin hydrate (12), wogonoside (13), baicalein (14), arctigenin (15), glycyrrhizin (16), wogonin (17), pulegone (18), decursin (19), and decursinol angelate (20).