Supplementary Information

Blood mercury, lead, cadmium, manganese and selenium levels in pregnant women and their determinants: the Japan Environment and Children’s Study (JECS)

Authors

Shoji F. Nakayama1, Miyuki Iwai-Shimada1, Tomoko Oguri1, 2, Tomohiko Isobe1, Ayano Takeuchi1, 3, Yayoi Kobayashi1, Takehiro Michikawa1, 4, Shin Yamazaki1, Hiroshi Nitta1, Toshihiro Kawamoto1, and the Japan Environment and Children’s Study Group5

Affiliations

1 Japan Environment and Children’s Study Programme Office, National Institute for Environmental Studies, 16-2 Onogawa, Tsukuba, Ibaraki 305-8506, Japan

2 Research Institute of Science for Safety and Sustainability, National Institute of Advanced Industrial Science and Technology, 16-1 Onogawa, Tsukuba, Ibaraki 305-8569, Japan

3 Department of Preventive Medicine and Public Health, Keio University, 35 Shinanomachi, Shinjukuku, Tokyo 106-8582, Japan

4 Department of Environmental and Occupational Health, School of Medicine, Toho University, 5-21-16 Omori-nishi, Otaku, Tokyo 143-8540, Japan

5 The Study Group members are listed in the Appendix

Corresponding author:

Shoji F. Nakayama, MD, PhD

Japan Environment and Children’s Study Programme Office, National Institute for Environmental Studies

16-2 Onogawa, Tsukuba, Ibaraki 305-8506, Japan

Telephone: +81 (29) 850-2786

E-mail: fabre@nies.go.jp

- Materials and methods

  - Solution preparation

A standard stock solution containing all the target elements (except Hg) was prepared in 0.14 M nitric acid. A standard stock solution of Hg was generated by mixing the Hg standard solution with 0.056 M nitric acid, 0.5% (w/v) EDTA and 1% (v/v) TMAH. The stock solutions were mixed together to obtain a stock calibration solution in which the Hg, Pb, Cd, Mn and Se concentrations were 200, 200, 20, 600 and 2,000 ng g−1, respectively. An internal standard mixture containing Y, In and Tl (each at 250 ng g−1) was prepared in the same manner. A dilution solution containing 2% v/v butan-1-ol, 0.1% TMAH, 0.5 g l−1 POE and 0.5 g l−1 EDTA was prepared on each day of the analysis. A series of calibration solutions was made using the stock calibration solution, the internal standard mixture and the dilution solution to obtain the calibration curves shown in Table S1.

  - Calculation of element concentrations

Hg concentrations were calculated as total Hg. The intensity of Pb isotopes was the sum of the intensities of 206Pb, 207Pb and 208Pb. To correct for spectral overlap from molybdenum oxide (95Mo16O), 111Cd intensity was calculated using the following equation:

[Cd] = [m/z 111 intensity] − [m/z 95 intensity] × [95Mo16O generation rate].

The 95Mo16O generation rate was derived from the following equation:

[95Mo16O generation rate] = [the intensity of m/z 111 when 100 ppb 95Mo standard solution was analysed] / [the intensity of m/z 95 when 100 ppb 95Mo standard solution was analysed].

- Results

  - Performance characteristics

The performance characteristics of the ICP-MS method that we developed are shown in Table S1.

  - Method robustness

    - Method comparison

Total Hg was analysed by cold vapour atomic absorption spectroscopy (CVAAS) as a ‘gold-standard’ method. Pb, Cd, Mn and Se were measured by the acid digestion method using ICP-atomic emission spectroscopy (AES) as a ‘gold-standard’ method. Figure S1 shows Bland-Altman plots comparing the ‘gold-standard’ methods and the current method. The current method was comparable with the traditional methods.

    - Mercury

Analysis of Hg was hampered by two problems: 1) instability of Hg in solution and 2) carryover of Hg on instruments. Use of a high pH and EDTA helped to prevent Hg from vaporising. The second problem was solved using two rinse solutions. The first rinse solution was water containing 1 mg ml−1 L-cysteine and was used to wash the ICP-MS auto-sampler. The second rinse solution was water containing 2% butan-1-ol, 1% TMAH, 0.5 g l−1 POE and 0.5 g l−1 H4EDTA and was used to washout a sample prior to analysis of the next sample. This method reduced Hg carryover to negligible levels.

    - Selenium

ICP-MS is susceptible to polyatomic interference. Polyatomic ions such as 40Ar37Cl+, 38Ar40Ar+ and 40Ar40Ar+ can interfere with signals of the isotopes 77Se, 78Se and 80Se. This can be circumvented by introducing a diatomic gas (e.g., N2) or an extra carbon source (e.g., organic solvents) into the plasma [1, 2]. Delvis and Sieniawska reported that addition of butan-1-ol into a sample solution prevents polyatomic interference and blockage of the ICP torch injector tube for more than 6 hours of operation [3]. This procedure was incorporated into our method. Additionally, the Se signal gradually decreased over the course of less than 5 hours of analysis (Figure S2: Room temperature). This problem was solved by ensuring samples were not left on the auto-sampler tray at room temperature for more than 2 hours and by keeping the other samples refrigerated (Figure S2: Refrigerated).

    - Manganese

39K16O and 54Fe1H can interfere with measurement of 55Mn. Whole blood samples contain a significant amount of potassium, which may form 39K16O and overlap with the reading of 55Mn. We analysed a solution containing 100 µg g−1 potassium by ICP-MS and monitored the intensity of m/z 55. This intensity was 8,000-fold lower in the potassium solution than in whole blood samples, which was considered negligible. Iron is also abundant in blood samples. We measured Mn in SeronormTM Trace Elements Whole Blood L-1 by electrothermal atomic absorption spectrometry (ETAAS), which was considered to experience little or no interference by 54Fe1H, at four independent laboratories. The mean Mn concentration was 19.5 ng g−1 (95% confidence interval (CI): 17.7–21.4 ng g−1). The same sample was analysed by the current method, which yielded a concentration of 18.8 ng g−1 (95% CI: 17.9–19.7 ng g−1). The two measurements did not significantly differ; therefore, we concluded that there was little or no interference by 54Fe1H on Mn measurements using the current method.

    - Stability

Stability of the calibration solutions was tested by analysing a series of calibration solutions on day 0, 5, 8 and 14. The calibration solutions were stable for at least 14 days (data not shown). We decided to make the calibration solutions fresh every week. Thaw-and-freeze stability was examined by thawing and freezing the same samples multiple times. The results did not change in four trials over 7 days (data not shown). For QC of day-to-day analysis, we used the Shewhart control chart (X-Rm chart) in accordance with ISO 7870. Sample X-Rm charts are shown in Figures S3 and S4.

    - External quality control

QC sample measurements in three contract laboratories were in good agreement, with differences of less than 10% RSD between each measurement (Figure S5). All three laboratories also joined the external quality control programme (i.e., German External Quality Assessment Scheme (G-EQUAS)) and proved that their results were within the tolerance ranges for all measured elements.

References

- Evans EH, Ebdon L. Simple Approach to Reducing Polyatomic Ion Interferences on Arsenic and Selenium in Inductively Coupled Plasma Mass-Spectrometry. Journal of Analytical Atomic Spectrometry. 1989;4(3):299-300.

- Larsen EH, Sturup S. Carbon-Enhanced Inductively-Coupled Plasma-Mass Spectrometric Detection of Arsenic and Selenium and Its Application to Arsenic Speciation. Journal of Analytical Atomic Spectrometry. 1994;9(10):1099-105.

- Delves HT, Sieniawska CE. Simple method for the accurate determination of selenium in serum by using inductively coupled plasma mass spectrometry. Journal of Analytical Atomic Spectrometry. 1997;12(3):387-9.

Figure Caption

Figure S1: Bland-Altman plots comparing gold-standard methods and the current method.

(A) CVAAS (gold-standard) vs. ICP-MS (current) for total mercury. Acid digestion (gold-standard) vs. alkaline dilution (current) for (B) lead, (C) cadmium, (D) manganese and (E) selenium.

Figure S2: Se stability over time after sample preparation.

The intensity of Se at time 0 was used as a reference for each condition.

Abbreviations: Se, selenium; Y, yttrium; Se/Y, the intensity of selenium divided by that of yttrium.

Figure S3: X charts for (A) mercury, (B) lead, (C) cadmium, (D) manganese and (E) selenium.

Figure S4: Rm charts for (A) mercury, (B) lead, (C) cadmium, (D) manganese and (E) selenium.

Figure S5: Boxplots of quality control sample measurements in three contract laboratories (A, B and C).

(A) Mercury, (B) lead, (C) cadmium, (D) manganese and (E) selenium.