Supplementary Information for

A low-voltage-driven MEMS ultrasonic phased-array transducer for fast 3D volumetric imaging

Yun Zhang1,2, Tong Jin1,2, Yining Deng1,2, Zijie Zhao1, *, Rui Wang3, Qiong He3, Jianwen Luo3, Jiawei Li4, Kang Du4, Tao Wu4, Chenfang Yan1,2, Hao Zhang1, Xinchao Lu1, Chengjun Huang1,2, *, and Hang Gao1, *

1 Institute of Microelectronics of the Chinese Academy of Sciences, Beijing, 100029, China

2 University of Chinese Academy of Sciences, Beijing, 100049, China

3 School of Biomedical Engineering, Tsinghua University, Beijing, 100084, China

4 School of Information Science and Technology, ShanghaiTech University, Shanghai, 201210, China

Single-cell equivalent circuit (EQC) model

A single cell in the array is developed based on a silicon-on-insulator (SOI) structure, which has a stack of several layers. A schematic diaphragm of the unimorph piezoelectric micromachined ultrasound transducer (pMUT) is shown in Fig. S1A. The pMUT contains a layer of piezoelectric material, which is sandwiched between two thin electrodes Since the pMUT relies on the strain mismatch between the active and passive layers to generate vibrations and to transmit or receive acoustic signals, we consider two main layers in the equivalent circuit model, the active layer (piezoelectric layer, PZT) and the passive layer (Si). The bottom face of the silicon layer is at z =0 as shown in Fig. S1B and the out-of-plane displacement of the plate is in the negative direction along the z-axis, which is denoted by w(r,t). Since the diameter of a circular pMUT is much larger than its thickness, the equivalent circuit (EQC) model can be settled based on classic plate theory39.

For such a multi-layer structure, it has a neutral layer:

zm=i=1nYi'zihii=1nYi'hi# AUTONUM  \* Arabic

Yi'=Yi1-vi2#2

where Yi, vi, hi and zi are Young’s modulus, Poisson’s ratio, thickness and the middle position of the ith layer.

The vibration equation of the circular diaphragm is39:

D∇2∇2w=q-ρs∂2w∂t#3

where q is the external pressure acting on the surface of the plate, and the mass density per unit area ρs is the density multiplied by the thickness. D is the plate flexural rigidity and can be presented as follows39:

D= I=1nYi'Ii+Zi2hi#4

The second moment of inertia and the distance from the neutral axis zm to the center of the ith layer are denoted as Ii and Zi:

Ii=hi312#5

Zi=zi-zm#6

In case of simple harmonic excitation, the displacement w (r, t) is given by39:

w=Wre-jωt#7

where ω is the angular frequency and W(r) is the shape function of the circular diaphragm when vibrating. Combining (4) and (7), the simplified equation of motion for the membrane vibration can be obtained:

∇2∇2W+β4W=qD#8

The general solution of (8) is as follows60:

Wr=-qω2ρs+Ai×J0βr+Bi×I0βr+Ci×Y0βr+Di×K0βr,i∈N#9

where β is the function of frequency, flexural rigidity and surface density and in the following form:

β4=ω2ρsD#10

and J0(x), I0(x), Y0(x) and K0(x) are the zero-order Bessel function of the first kind, the modified Bessel function of the first kind, the Bessel function of the second kind, and the modified Bessel function of the second kind, respectively. The coefficients Ai, Bi, Ci and Di are constants which are determined by boundary conditions.

The shape functions of region I and region II are the forms as follows:

WIr=-qω2ρs+A1J0βr+B1I0βr,  r∈(0,a)#11

WIIr=-qω2ρs+A2J0βr+B2I0βr+C2Y0βr+D2K0βr,  r∈(a,b)#12

As x tends to zero, the Bessel function Y0(x) and K0(x) approach infinity, so these two Bessel functions are discarded for the calculation of Region I. In order to solve (11) and (12), certain boundary conditions must be satisfied, i.e., continuity and the clamped boundary conditions.

At the junction of Region I and Region II, the displacements of the membrane, the first order derivatives of the displacements, the radial moment and the radial shear forces must be equal to ensure continuity:

WIa=WIIa#13

dWIdrr=a=dWIIdrr=a#14

Mrr,Ir=a=Mrr,IIr=a#15

Qr,Ir=a= Qr,IIr=a#16

where the radial moments are different in the two regions because only Region I is located directly below the top electrode39:

Mrr,I=-D∂2w∂r2+vr∂w∂r-Y0d31hV1-v#17a

Mrr,II=-D∂2w∂r2+vr∂w∂r#17b

and the radial shear force is defined as60:

Qr=-D1r∂∂rr∂w∂r#18

For a circular membrane with the clamped boundary conditions, the displacement of the membrane at the edge (r=b) must be zero, and the slope of the displacement along the radial direction is also zero at the edge39:

WIIb=0#19

dWIIdrr=b=0#20

On the basis of the shape function, the volumetric displacement of the membrane vibration is integrated by the displacement function over the area of the circular membrane27:

Wvol=2π0aWIrrdr+abWIIrrdr#21

In order to obtain the specific displacement shape function for the region I and region II, it is necessary to use the continuity conditions and the clamped boundary conditions to obtain A1, B1, A2, B2, C2 and D2. The boundary conditions (13) - (16) and (19) - (20) lead to six equations to solve for these constants, which are expressed in matrix form39:

&J0βaI0βa-J0βa-I0βa-Y0βa-K0βaJ1βa-I1βa-J1βaI1βa-Y1βa-K1βaJJβaIIβa-JJβa-IIβa-YYβa-KKβaJJJβaIIIβa-JJJβa-IIIβa-YYYβa-KKKβa00J0βbI0βbY0βbK0βb00-J1βbI1βb-Y1βb-K1βb#&×A1B1A2B2C2D2=00-Y0d31hVDβ21-v0qω2ρs0#(22)

where the functions II(x), JJ(x), YY(x), KK(x), JJJ(x), III(x), YYY(x) and KKK(x) are:

JJx=-0.5J0x-vxJ1x+0.5J2x#23

IIx=0.5I0x+vxI1x+0.5I2(x)#24

YYx=-0.5Y0x-vxY1x+0.5Y2x#25

KKx=0.5K0x-vxK1x+0.5K2x#26

JJJx=-2xJ0x+4+3x2J1x+2xJ2x-x2J3x#27

IIIx=2xI0x+-4+3x2I1x+2xI2x+x2I3x#28

YYYx=-2xY0x+4+3x2Y1x+2xY2x-x2Y3x#29

KKKx=2xK0x+4-3x2K1x+2xK2x-x2K3x#30

The expressions in the right-hand side matrix of (22) are the external force acting on the membrane, first of which is input voltage dependent and the second of which is caused by the incident sound pressure. Therefore, the constants we solve for in (22) are also related to both voltage (V) and acoustic pressure (q), so the final volumetric displacement (Wvol) can also be directly divided into two terms related to voltage and acoustic pressure27:

Wvol=Ymq+btV#31

where Ym is the volumetric displacement generated by per unit incident pressure; bt is the electromechanical transduction coefficient and is defined as the volumetric displacement by per unit input voltage.

Ym=2πq0aWIrrdr+abWIIrrdrV=0 #32

bt=2πV0aWIrrdr+abWIIrrdrq=0 #33

Therefore, the mechanical impedance and the electromechanical transduction ratio shown in Fig. S2 are derived as follows:

Zm=1jωYm#34

η=btYm#35

The Equivalent circuit model established for the single circular pMUT is shown in Fig. S2, where C0 is the blocked capacitance and it is calculated as follows61:

C0=εzAehp#36

where εz is the permittivity of piezoelectric material, Ae is the area of top electrode and hp is the thickness of the piezoelectric layer. Then the electrical impedance can be derived as:

Ze=1jωC0#37

And Za is the acoustic radiation impedance which in the form as follows40,41:

Za=ρc1-J12kbkb+iH12kbkbA#38

where k is the wavenumber, A is the effective area of the diaphragm, J1 is the first-order Bessel function of the first kind and H1 is the first-order Struve function.

Single cell characterization results

Both the EQC model and the finite element method (FEM) model calculated the volumetric displacement-frequency response in air and water, respectively and their results agree well with each other as shown in Fig. S3A. The resonant frequency in air is about 3.3 MHz, while in water it is about 2.7 MHz. And the calculation error of the EQC model is about 7%. To ensure that the device operates at the correct mode, we examined the mode at the resonant frequencies in water calculated by volumetric displacement-frequency response. As shown in Fig. S3B, pMUT operates in mode I at the frequency of 3.3 MHz in water. Then the effect of the coverage of the top electrode on the membrane vibration displacement was investigated. The electrode we chose is a circular inner electrode, and the displacement of pMUT vibration increases with the increase of electrode radius, and the vibration displacement of pMUT center reaches the maximum at the top electrode coverage ratio of 0.67. When the upper electrode size increases further towards covering the whole piezoelectric layer, the membrane displacement decreases to zero.

Fig. S1 Scheme of the single piezoelectric micromachined ultrasound transducer (pMUT) cell. A. The SOI structure of the PZT pMUT. B. The pMUT cell before actuation and the adopted polar coordinate system. C. The piezoelectric layer actuated by an AC voltage and the pMUT divided by two regions. Region I is underneath the top electrode. Region II is the area beyond the top electrode.

Table S1. Piezoelectric transducers (conventional bulk PZT, 1-3 composite PZT, cMUTs, pMUTs) for in-air use, for monitoring or imaging the epidermis and deep tissues.

|  | Piezo. material | Freq. | Voltage | Application |
| --- | --- | --- | --- | --- |
| Conventional Transducer | 1-3 composite PZT | 2MHz | 25 V | Continuous monitoring of deep-tissue hemo-dynamics8 |
|  | 1-3 composite piezoelectric material | 3,7,10 MHz | 40V | Long-term continuous imaging of diverse organs2 |
|  | Bulk PZT | N/A | High-voltage | Epidermal patch：monitoring of hemo-dynamic and metabolic biomarkers36 |
|  | PZT-5H | 2-6MHz | High-voltage | Monitoring deep tissues in moving subjects10 |
|  | 1-3 composite PZT | 2.4MHz | 100 V | Photoacoustic patch：imaging of hemoglobin and core temperature37 |
| cMUT | N/A(cMUT) | 0.443-4.16 MHz | 40 V bias | Portable applications in air34 |
|  | N/A(cMUT) | 3.3-4.2 MHz | 7.4-25 V bias | N/A35 |
| pMUT | PZT | 4.8 MHz | 20 V | 2D intracardiac echo (ICE)31 |
|  | PMN-PZT | 5,10 MHz | 40V | Harmonic imaging and image-guided therapy38 |
|  | (K, Na) NbO3 | 104.5 kHz | 2V | In-air haptic feedback, loudspeaker, and AR/VR systems33 |
|  | PZT | 3 MHz | 5 V | Volumetric 3D imaging at 7cm depth (this work) |

Table S2.  Structural and material parameters of the single cell pMUT.

| Symbol | Value | Unit | Description |
| --- | --- | --- | --- |
| a | b×α | μm | Top electrode radius |
| b | 200 | μm | Diaphragm radius |
| h1 | 8 | μm | Si thickness |
| h2 | 5 | μm | PZT thickness |
| α | 0.67 | / | Top electrode coverage ratio |
| Y1 | 170 | GPa | Si Young’s modulus |
| v1 | 0.28 | / | Si Poisson’s ratio |
| ρ1 | 2329 | kg/m3 | Si density |
| Y2 | 83.26 | GPa | PZT Young’s modulus |
| v2 | 0.36 | / | PZT Poisson’s ration |
| ρ2 | 7500 | kg/m3 | PZT density |
| d31 | 123 | pm/V | PZT piezoelectric constant |
| ε | 950 | / | PZT Dielectric constant |

Fig. S2. Equivalent Circuit Model of the single piezoelectric micromachined ultrasound transducer (pMUT) cell.

Fig. S3. A. Volumetric displacement-frequency response of the single cell PZT piezoelectric micromachined ultrasound transducer (pMUT) in air and water. B. The model shape at resonance of the single cell pMUT in water. C. Effect of top electrode coverage(α) on membrane center displacement.

Fig. S4. Center frequency distribution characterization of the 8×8 piezoelectric micromachined ultrasound transducer (pMUT) phased array transducer with 64 elements (channels).

Fig. S5. Optical microscope photos of different layout elements. A. 3×3 layout. B. 4×4 layout.

Fig. S6. Depths underneath the skin and dimensions of representative tissues and organs of interest8.