Supplementary Information for

High-Uniformity Miniaturized PMUT Array with Broadband and High-Sensitivity for Wearable Ultrasound Imaging

Xingli Xu1, Wanli Yang1, Zhenzhen Wang2, Yongquan Ma1, Yuewu Gong3, Zhuochen Wang4*, Xiaohua Jian5, Xiaochun Wang6, Sheng Zhou6, Wei Pang1, Pengfei Niu1*

1. State Key Laboratory of Precision Measuring Technology and Instruments, Tianjin University, Tianjin 300072, China.

2. Department of Ultrasound, Tianjin Chest Hospital, Tianjin 300072, China

3. Precision Measurement Research and Development Center, Sun Yat-sen University Nanchang Research Institute, Nanchang, 330200, China.

4. College of Mechanical and Electrical Engineering, Beijing University of Chemical Technology, Beijing 100029, China.

5. Institute of Functional Materials and Intelligent Manufacturing, Nanjing University, Suzhou 215163, China.

6. Institute of Biomedical engineering, Chinese Academy of Medical Sciences and Peking Union Medical college, Tianjin 300052, China.

*Corresponding authors: zc.wang@buct.edu.cn; niupengfei2018@tju.edu.cn.

Contents

- Supplementary Notes

Note S1. Design of individual PMUT cell

Note S2. Equivalent circuit model for PMUT array (ECMA)

Note S3. Design of surface encapsulation layer for PMUT array

Note S4. Fabrication process for the PMUT array

Note S5. Calculation of blood pressure with ultrasound imaging data

- Supplementary Figures

Fig. S1. Schematic drawing of an individual PMUT cell

Fig. S2. Equivalent circuit models for PMUT array (ECMA)

Fig. S3. The acoustic radiation impedance of PMUT cell

Fig. S4. Comparison of the 2D acoustic field between the equivalent circuit model (ECM) and finite element method (FEM)

Fig. S5. Crosstalk analysis model for large PMUT arrays

Fig. S6. Effects of inter-cell spacing (d) on vibration amplitude and phase for all cells in a 20×20-element array

Fig. S7. Effects of inter-cell spacing (d) on the acoustic field distribution of 20×20-element PMUT array

Fig. S8. Ultrasound pressure and bandwidth characteristics of 20×20-element PMUT array

Fig. S9. Effects of inter-cell spacing (d) on vibration amplitude and phase for all cells in a 2.2mm x 2.2 mm array

Fig. S10. Effects of inter-cell spacing (d) on the acoustic field distribution of a 2.2mm x 2.2 mm array

Fig. S11. Ultrasound pressure and bandwidth characteristics of a 2.2mm x 2.2 mm array

Fig. S12. Pulse-echo spectra of a 2.2mm x 2.2 mm array (cell diameter = 80μm and a center frequency being about 3MHz) with various inter-cell spacing d

Fig. S13. The impact of PMUT size on wearability

Fig. S14. The influence of beam-forming on the ultrasound pressure fields generated by a ~7 MHz, 64-element linear ultra-narrow PMUT array

Fig. S15. The influence of phased-array beam-forming on the imaging characteristics of the ~7 MHz, 64-element linear ultra-narrow PMUT array

Fig. S16. Surface encapsulation layer for PMUT array

Fig. S17. Reduction of mechanical crosstalk through isolation trenches

Fig. S18. Curve of the complex polynomial coefficient A(ka) as a function of ka.

Fig. S19. Fabrication flow of our PMUT device

Fig. S20. SEM image of the vibration diaphragm's cross section of an individual PMUT.

Fig. S21. The ultrasound signal variation of the PMUT array at various temperatures.

Fig. S22. Carotid artery ultrasound images acquired at 1-second intervals from continuous imaging

Fig. S23. Original ultrasound images before post-processing of a. carotid artery, b. thyroid gland, c. dorsalis pedis artery.

Fig. S24. Comparison of ultrasound images from our PMUT array versus those acquired with a clinical linear transducer and reference PMUTs.

- Supplementary Tables

Table S1. Polynomial coefficients for the function A(ka) in Fig. S18 and Eq. (35)

Table S2. The structural and material parameters in our individual PMUT cell

Table S3. Design specifications of our developed linear AlN-PMUT array

Table S4. Comparison about the ultrasound bandwidth and signal noise ratio between reference PMUTs and piezoelectric bulk ultrasound transducer

Table S5. Comparison of transducer size and resolution between this work and representative reported studies

Table S6. Comparison of the transducers' surface temperature and power consumption between this work and representative reported studies

- Supplementary Movies about the real time ultrasound imaging of the different human tissues.

Movie S1. The long-axis of the carotid artery

Movie S2. The short-axis of the carotid artery

Movie S3. The carotid sinus

Movie S4. The internal and external carotid arteries

Movie S5. Transverse section of the thyroid gland

Movie S6. Longitudinal section of the thyroid gland

Movie S7. The dorsalis pedis artery

Note S1. Design of individual PMUT cell

Fig. S1 shows a schematic cross-sectional view of the PMUT (CP) employed in this study. From top to bottom, the diaphragm comprises a silicon dioxide (SiO₂) elastic layer, a molybdenum (Mo) top electrode, an aluminum nitride (AlN) piezoelectric layer, and a Mo bottom electrode. Beneath the diaphragm lies a silicon (Si) substrate featuring an etched cavity. The cavity boundary clamps the diaphragm and provides mechanical support.

The elastic layer functions to shift the neutral plane position of the diaphragm by utilizing its stiffness, thereby creating a moment arm between the neutral plane and the piezoelectric layer. When a voltage is applied, this moment arm generates a bending moment, driving the flexural vibration of the diaphragm. Among various materials for elastic layers such as Si [1], Si₃N₄ [2], SiO₂ [3], AlN [4,5], and Mo [6]. Since SiO₂ has a relatively low Young’s modulus and small residual stress during deposition, which is beneficial for fabricating PMUT arrays with high uniformity, thus it was selected as the elastic layer herein and positioned at the outermost surface of the diaphragm.

Designing an individual PMUT cell involves optimizing its frequency response. The resonance frequency f of a circular CP is related to its radius, as described by Eq. (1):

(1)

where k0 and m0 are the equivalent stiffness and equivalent mass of the vibration diaphragm. t is the thickness and a is the radius of the diaphragm of CP. From Eq. (1), the f is proportional to t and inversely proportional to a2. Thus a thicker diaphragm and smaller radius lead to a higher resonance frequency.

The transmit acoustic pressure P of a single CP cell can be expressed by Eq. (2):

|  | (2) |
| --- | --- |

where ω is the angular frequency, ρ is the medium density, v is the vibration velocity of the CP, r is the distance from the CP surface to the measurement point, dc is the static displacement amplitude of the CP, and Q is the quality factor. The effective vibrating radius and effective vibrating area of the CP are denoted by aeff and Aeff, respectively, and A is the physical area of the CP. The effective vibrating area Aeff describes the performance difference between a PMUT and an ideal piston transducer. For equal center displacements, the CP's output is roughly one-third of the value produced by a piston transducer.

When the CP operates in water, most of the energy is dissipated into the medium as acoustic energy [1, 2] while the quality factor Q is small [3]. From Eq. (2), the transmit pressure P of the CP is approximately proportional to the static displacement dc as follows:

(3)

According to the vibration theory of a d31 mode bending plate, the static displacement dcis inversely proportional to the equivalent stiffness k0 and proportional to the driving moment arm as per Eq. (4).

(4)

Here the driving moment arm  is the distance between the piezoelectric layer and the neutral plane of the diaphragm and is proportional to the diaphragm thickness t. The equivalent stiffness k0 of the CP is related to the diaphragm dimensions

() [4], and thus dc can be written as Eq. (5):

(5)

Substituting Eq. (1) into Eq. (5) yields Eqs. (6):

(6)

Thus for PMUT devices designed to operate at the same resonance frequency, smaller radius and thinner diaphragms generally lead to better performance. However, reducing the device size increases fabrication difficulty, and making the piezoelectric layer too thin can degrade breakdown strength and reliability. Considering these trade-offs, we set the CP diameter to 40 µm in this work. From the vibration mode of the CP, the electromechanical coupling efficiency is maximized when the top electrode diameter is about 70% of the CP diameter. Therefore, the top electrode diameter is chosen as 28 µm.

To reduce parasitic resistance and energy loss in the electrodes, the thickness of the Mo electrodes is set to 0.1 µm. The AlN piezoelectric layer thickness is set to 0.5 µm. Electrical tests show that at this thickness the PMUT can withstand excitation voltages above 100Vpp for long periods without damage, ensuring safety and reliability in human imaging applications. The SiO₂ elastic layer thickness is set to 1.2 µm, which provides high sensitivity while keeping the resonance frequency near 7 MHz, within the 7~15 MHz range required for high-resolution imaging of superficial organs such as the carotid artery.

Detailed structural parameters of the clamped PMUT are listed in Table S2.

Note S2. Equivalent circuit model for PMUT array (ECMA)

For the PMUT, the electrical, mechanical, and acoustic domains are coupled. An applied voltage induces strain in the piezoelectric layer via the inverse piezoelectric effect, drives membrane motion and emits an acoustic wave into the surrounding medium. Conversely, an incoming acoustic wave deforms the membrane and, through the piezoelectric effect, generates charge and voltage. This interaction is modeled using an equivalent circuit model (ECM), which maps geometry, material and fluid parameters onto the system’s electrical, mechanical and acoustic responses, thereby predicting PMUT performance and guiding design.

Each PMUT cell’s ECM contains electrical, mechanical and acoustic domains coupled by the electromechanical conversion coefficient η and PMUT effective vibrating area Aeff. The Aeff and η are defined by Eqs. (7) and (8):

(7)

(8)

where e31,f is the effective piezoelectric coefficient of piezoelectric material. ZP is the distance from the middle of the piezoelectric layer ZP to the neutral layer position ZNA. Ipizeo is the strain integral associated with the modal vibrational mode φ. aeff and  are the effective radius and the real cavity radius of the PMUT, respectively. ,  and φ can be expressed as Eqs. (9), (10), (11):

(9)

(10)

(11)

where  is the thickness of each layer of the PMUT.  is the height from the center of each layer to the bottom surface.  is the area of each layer.  denotes the plate modulus, where  and  denote the Young’s modulus and Poisson, s ratio of the material, respectively.  denotes the ratio of the radius of the top electrode to the real radius of the cavity.

In the electrical domain,  represents is the input voltage and  represents the static capacitance of the PMUT,  can be calculated by Eq. 12.

(12)

where  and  are the vacuum and relative dielectric constants,  is the thickness of the piezoelectric layer.

In the mechanical domain, the force F represents the effective driving force of the piezoelectric material.  and  are the equivalent stiffness and equivalent mass of the multilayer plate structure of the PMUT in the ECMC, respectively. They can be obtained as follows:

(13)

(14)

(15)

where  is the material density of each layer of the PMUT. In the mechanical domain, the current  is the vibration velocity of the PMUT membrane.

The vibration of the PMUT produces a corresponding acoustic wave Pout in the acoustic domain, where the self-acoustic radiating impedance  is used to quantify the resistance of the acoustic medium to the diaphragm motion. It arises from the radiating effect of the PMUT film itself.  can be expressed as Eq. (16):

(16)

where  represents the density of the medium and  represents the acoustic velocity of the medium.  and  represent the real and imaginary parts of the self-acoustic radiating impedance, respectively.  is the first-order Bessel function of the first class and k is the wave number.

When multiple PMUT cells are formed as an array, the equivalent circuit model of array (ECMA) is shown in Fig. S2. The array consists of N independent cells Pi  (i=1, 2, ..., N). Each cell has an intrinsic capacitance , an electromechanical conversion coefficient , an effective stiffness , and an effective mass . Under the synchronous operation of the array, the acoustic radiation fields of the different cells are coupled to each other. The coupling between any two cells Pi and Pj can be described by the mutual-acoustic radiation impedance Zij in the form of Eq. (17).

(17)

where d is the inter-cell distance between adjacent PMUT cells. A is a function of the product of wave number and radii (ka). The analytical expression and the curves of A(ka) are shown in Fig. S3. The polynomial coefficients for the function A(ka) are listed in Table S1.

In a PMUT array, all active cells introduce crosstalk to the target cell Pi. So the total crosstalk suffered by Pi can be written as Eq. (18).

(18)

where  and  denote the volume velocities of Pi and Pj.

Crosstalk changes the vibrational velocity of each PMUT cell. Thus the complete acoustic radiating impedance network must be constructed to resolve the overall acoustic response of the array. According to Kirchhoff’s voltage law (KVL), the self-consistent equation of the ECMA can be described as Eqs. (19) and (20).

(19)

(20)

Eq. (19) can be written as Eq. (21) [7].

(21)

where vector  denotes the vibration velocity of all cells in the array,  denotes the electromechanical coupling coefficient vector,  represents the input voltage vector,  refers to the Hadamard (element-wise) product.  is the mechanical impedance matrix,  is the self-radiation impedance matrix, and  is the mutual radiation impedance matrix. Their expressions are given as follows by Eqs. (22), (23) and (24).

(22)

(23)

(24)

The mechanical impedance of cell Pi in Eq. (22) can be written as Eq. (25).

(25)

By solving Eq. (21), the cell velocity matrix  within the PMUT array can be obtained. And the complete acoustic field generated by the PMUT array can then be solved by the following mechanical-acoustic coupling Eq. (26).

(26)

Note S3. Design of surface encapsulation layer for PMUT array

Surface encapsulation layer plays a key role in the wearable application of PMUT. It provides comfort contacting with human skin, also is better to enhance the acoustic performance of PMUT array. Also it safeguards the device, particularly during in vivo human imaging procedures. Silicone elastomers are now the preferred material for PMUT acoustic impedance matching layers, thanks to their biocompatibility, closely matching the acoustic impedance of water, and the excellent processability [7]. PDMS, dragon skin, and ecoflex are among the most commonly used materials. Fig.S16 illustrates the structure and operating mechanism of our proposed encapsulation layer. When the PMUT film vibrates, its mechanical velocity drives the encapsulation layer medium to contract and expand periodically, creating a propagating acoustic wave Tx inside the encapsulation layer. Since silicone materials have a compressive velocity of ~1000 m/s, the wavelength inside the matching layer is ~150 µm which is around one-seventh the thickness of the encapsulation layer generally applied (>1 mm). Therefore the effect of multiple reflected wave P2 inside the encapsulation layer can be neglected. The final effective acoustic wave entering water is denoted as P1. To evaluate the contribution of the encapsulation layer to the acoustic pressure transmission efficiency, we define the transmission coefficient T as the ratio of the acoustic pressure amplitude with / without the encapsulation layer, which can be calculated from Eq. (29). The parameters in Eq.(29) can be obtained from Eqs. (27), (28) and (30).

(27)

(28)

(29)

(30)

where t is the matching layer thickness;  and  are the attenuation coefficients of the silicone encapsulation layer and the water respectively.  and  are the surface pressure magnitude of PMUT in encapsulation layer and water.  represents the transmission coefficient of acoustic wave propagating from the encapsulation layer to water.  is the magnitude of incident acoustic pressure reaching the matched layer/water interface, and  is the magnitude of incident acoustic pressure measured at the same depth in the absence of the encapsulation layer (the region is replaced by water). Both  and  are calculated from ECMC. The acoustic impedances of the water and the encapsulation layer were  = 1.5 MRayl and  = 1.0 MRayl, respectively. The attenuation coefficients of the silicone encapsulation layer and the water were ~2 dB/mm @5MHz and 0.0055 dB/mm @5MHz, respectively [7, 8]. Substituting the above parameters into Eq. (29) yields the acoustic pressure transmission efficiency T = 58.7% for the 1mm silicone encapsulation layer.

In order to avoid liquid silicone penetrating into the cavity and affecting the performance of the device, a bilayer silicone encapsulation structure was used in this study, including an inner layer of PDMS and an outer layer of ecoflex. The process proceeded as follows: A 15 µm-thick PDMS layer was covalently bonded to the SiO₂ surface to achieve initial cavity sealing. Subsequently, Ecoflex precursors were mixed evenly and cast over the entire device to provide secondary encapsulation and a soft skin-contact interface after curing. The resulting structure exhibits reliable sealing, excellent acoustic wave transmission properties, and good biocompatibility, thereby enabling its use in long-term wearable ultrasound imaging applications.

Note S4. Fabrication process for the PMUT array

The PMUT employs AlN as the piezoelectric film and it was fabricated through sacrifical layer process. Fig. S19 schematically show the fabrication flow. Firstly 1.5μm deep cavity was etched on silicon wafer and PSG was deposited on the cavity wafer. Chemical mechanical polishing was then carried out to flatten the wafer and a slightly over polishing was implemented, making the final cavity depth to be around 1.2μm. Subsequently, the bottom Mo electrode layer was grown and patterned, followed by the deposition of the AlN piezoelectric film and top Mo electrode layers, with the top electrode then patterned. Then the etch hold was opened, and the PSG was removed in buffered oxide etch (BOE) to release the cavity. Next 1.2μm thick SiO2 elastic layer was deposited on the structured layer. This thickness is equal to the cavity depth and it can partly fill the release hole on silicon substrate. Then we removed the SiO2 over the electrical pads and pathways, and deposited gold layer for electrical connection. After wafer dicing and device bonding to a printed circuit board, the silicone elastomer encapsulation layer was formed on the device. The gold patterns are not shown in the drawings. Subsequently, the SiO₂ over the electrical pads and pathways was removed, and a gold layer was deposited for electrical connection. The gold patterns are not shown in the drawings. After wafer dicing and device bonding to a printed circuit board, a silicone elastomer encapsulation layer was formed on the device.

Note S5. Calculation of blood pressure with ultrasound imaging data

Based on the dynamic changes in arterial diameter extracted from sequences of continuous ultrasound images, the corresponding blood pressure waveform can be further calculated [9]. The variation of blood pressure with time 𝑡 can be expressed as Eq. (31)

(31)

where  is the diastolic pressure.  is the diastolic arterial cross-sectional area.  represents the cross-sectional area of the artery at time t.  represents the stiffness coefficient of the artery. Normally the cross-section of an artery is circular in shape, so the cross-sectional area A(t) of an artery can be expressed by Eq. (32).

(32)

where  is the diameter of the artery at time t. Human blood vessels have good elasticity. The hysteresis of the pressure-diameter relationship is usually less than 0.2%. Therefore Eq. (32) can be substituted into Eq. (31) to reconstruct the blood pressure waveform using the vessel diameter waveform (Eq. (33)).

(33)

where  is the cross-sectional diameter of the artery in diastole. Therefore,  is denoted as Eq. (34).

(34)

where  is the cross-sectional diameter of the artery during systole and  denotes systolic blood pressure.  and  can be measured by commercially available blood pressure measuring devices, such as an upper arm electronic sphygmomanometer. Once the arterial stiffness coefficient  is obtained, the change in blood pressure waveform over time can be detected in real time by using the vessel diameter measured from real-time ultrasound images.

Fig. S1. Schematic drawing of an individual PMUT cell

Fig. S2. Equivalent circuit models for PMUT array (ECMA)

Fig. S3. The acoustic radiation impedance of PMUT cell, a. for a single PMUT cell, there is no crosstalk, and only the self-acoustic radiation impedance from the medium acts upon it. The real (Rself) and imaginary (Xself) parts of the self-acoustic radiation impedance vary with ka (k × a). b. when two PMUT cells in an array operate simultaneously, they experience mutual crosstalk. The mutual-acoustic radiation impedance components Rij and Xij are 90° out of phase. This makes the crosstalk to decay in an oscillatory manner as the inter-cell spacing d increases.

Fig. S4. Comparison of the 2D acoustic field between the equivalent circuit model (ECM) and finite element method (FEM). A linear PMUT array consists of 20 cells connected in parallel, with a circular cell diameter of 40 µm and an inter-cell spacing of 90 µm, is used in this simulation to compare the acoustic field distribution at the same frequency. a. FEM simulation of the acoustic field, using a computation time of ~30min. b. The acoustic field computed by PMUT array’s ECM (ECMA), with a computation time of ~3s. c. The acoustic pressure along the z-axis over the surface of PMUT array extracted from the calculated acoustic field by FEM and ECM, respectively.

The results demonstrate that the two curves exhibit high consistency in tendency, focus position, and acoustic pressure values, confirming that the ECM method offers significant computational accuracy. The computing time for ECM is significantly shorter than that of FEM, demonstrating a substantial advantage in computational efficiency.

Fig. S5. Crosstalk analysis model for large PMUT arrays. a. For large PMUT array, calculating the mutual-acoustic radiation impedance for every cell is impractical. When the array contains numerous cells, most cells exhibit nearly identical amplitude and phase profiles. In this regime, mutual impedance becomes weakly dependent on individual cell velocities, allowing the originally nonlinear coupling to be approximated as linear. Consequently, a representative element within the array can be isolated and its crosstalk can be analyzed separately. b. Reduced equivalent-circuit model (RECM) for a single array cell. Zmulti denotes the total mutual-acoustic radiation impedance (i.e., acoustic crosstalk) acting on the PMUT cell.

Fig. S6. Effects of inter-cell spacing (d) on vibration amplitude and phase for all cells in a 20×20-element array (cell diameter = 40μm). a. d = 50 μm, b. d = 70 μm, c. d = 90 μm, d. d = 110 μm, e. d = 130 μm, f. d = 150 μm.

Fig. S7. Effects of inter-cell spacing (d) on the acoustic field distribution of 20×20-element PMUT array (cell diameter = 40μm). a. d = 50 μm, b. d = 70 μm, c. d = 90 μm, d. d = 110 μm, e. d = 130 μm, f. d = 150 μm.

Fig. S8. Ultrasound pressure and bandwidth characteristics of 20×20-element PMUT array (cell diameter = 40μm). a. normalized sound pressure distribution from -10° to 10° relative to the sound pressure at 0°. The ultrasound pressure is obtained at 2cm away from the PMUT surface. b. frequency-dependent normalized sound pressure relative to its maximum value at the natural focus point. c. effect of cell spacing d on the acoustic sensitivity at the natural focus point and −6 dB ultrasound transmission bandwidth. They were normalized to the maximum value in this group.

Fig. S9. Effects of inter-cell spacing (d) on vibration amplitude and phase for all cells in a 2.2mm x 2.2 mm array (cell diameter = 40μm). a. d = 50 μm, b. d = 70 μm, c. d = 90 μm, d. d = 110 μm, e. d = 130 μm, f. d = 150 μm.

Fig. S10. Effects of inter-cell spacing (d) on the acoustic field distribution of a 2.2mm x 2.2 mm array (cell diameter = 40μm). a. d = 50 μm; b. d = 70 μm; c. d = 90 μm; d. d = 110 μm; e. d = 130 μm; f. d = 150 μm.

Fig. S11. Ultrasound pressure and bandwidth characteristics of a 2.2mm x 2.2 mm array (cell diameter = 40μm) with various inter-cell spacing (d). a. normalized sound pressure distribution from -10° to 10° relative to the sound pressure at 0°. The ultrasound pressure is obtained at 2cm away from the PMUT surface. b. frequency-dependent normalized sound pressure relative to its maximum value at the natural focus point. c. effect of cell spacing d on the acoustic sensitivity at the natural focus point and −6 dB ultrasound transmission bandwidth. They were normalized to the maximum value in this group.

Fig. S12. Pulse-echo spectra of a 2.2mm x 2.2 mm array (cell diameter = 80μm and a center frequency being about 3MHz) with various inter-cell spacing d. a. d = 105 μm, b. d = 100 μm, c. d = 90 μm; d. comparison of sensitivity and −6 dB bandwidth, showing decreasing performance as the increase of inter-cell spacing.

Fig. S13. The impact of PMUT size on wearability. a and b are the schematics of a wide and ultra-narrow PMUT array for wearable applications respectively.

Fig. S14. The influence of beam-forming on the ultrasound pressure fields generated by a ~7 MHz, 64-element linear ultra-narrow PMUT array (1 cm length × 0.15 cm width). a. ultrasound pressure field distributions without beam-forming and with the main lobe focused at 20 mm depth via beam-forming, b. normalized pressure values relative to the reference pressure at the focal point, measured along the horizontal axis and the z-axis, both passing through the focal position.

Fig. S15. The influence of phased-array beam-forming on the imaging characteristics of the ~7 MHz, 64-element linear ultra-narrow PMUT array (1 cm length × 0.15 cm width). a. Ultrasound images of three human hairs (~60 μm in diameter) obtained with and without beam-forming techniques. b. lateral (x-direction) and axial (z-direction) resolutions obtained with and without beam-forming techniques.

Fig. S16. Surface encapsulation layer for PMUT array. a. schematics of the PMUT array coated with a surface encapsulation layer, b. acoustic-transmission principle through the surface encapsulation layer, c. structure of the bi-layer PDMS/Ecoflex silicone surface encapsulation layer used in this work.

Fig. S17. Reduction of mechanical crosstalk through isolation trenches. a. a finite element simulated example of stress distribution on silicon substrate with trench under two-cell excitation. b. the effect of trench depth on mechanical crosstalk, demonstrated by normalizing the center displacement of the cross-coupled PMUT cells to the values when no trench exists.

Fig. S18. Curve of the complex polynomial coefficient A(ka) as a function of ka.

A(ka) is a fundamental parameter in mutual acoustic radiation impedance and can be expressed as a tenth-order polynomial by Eq. (35) [1]. The coefficients pi are listed in Table S1.

A(ka)=p1 ka 10+ p2 ka 9 + p3 ka 8 +⋯+ p10 ka + p11			(35)

Fig. S19. Fabrication flow of our PMUT device.

Fig. S20. SEM image of the vibration diaphragm's cross section of an individual PMUT.

Fig. S21. The ultrasound signal variation of the PMUT array at various temperatures.

Fig. S22. Carotid artery ultrasound images acquired at 1-second intervals from continuous imaging.

Fig. S23. Original ultrasound images before post-processing of a. carotid artery, b. thyroid gland, c. dorsalis pedis artery.

Fig. S24. Comparison of ultrasound images from our PMUT array (g, h, i) versus those acquired with a clinical 160-channel linear transducer (L12-3ERGO) (a, b) and reference PMUTs based on PZT (c, d) [10, 11] or AlN (e, f) [12, 13] thin films using a research-grade imaging system. Images of a. the carotid artery, b. dorsalis pedis artery, c. carotid artery, d. carotid artery, e. 6 mm metal block, f. 0.4 mm copper pin, g. carotid artery, h. dorsalis pedis artery, i. 0.06 mm human hair.

Table S1. Polynomial coefficients for the function A(ka) in Fig. S18 and Eq. (35)

| pn |  | Real | Imagery |
| --- | --- | --- | --- |
| n | 1 | -5.34237e-07 | -2.15428e-07 |
|  | 2 | 9.60487e-06 | 6.64439e-06 |
|  | 3 | -1.76278e-05 | -8.03963e-05 |
|  | 4 | -6.79795e-04 | 4.61167e-04 |
|  | 5 | 5.09681e-03 | -1.14002e-03 |
|  | 6 | -7.47336e-03 | 7.16523e-04 |
|  | 7 | -2.50852e-02 | -2.72073e-03 |
|  | 8 | -9.71827e-03 | 1.47321e-02 |
|  | 9 | 2.87292e-01 | -9.01486e-04 |
|  | 10 | -1.09336e-03 | 1.92266e-04 |
|  | 11 | 8.62694e-05 | -1.42685e-05 |

Table S2. The structural and material parameters in our individual PMUT cell

| Parameter | Material | Radius (μm) | Thickness (μm) | Young's modulus (GPa) | Density (kg/m3) |
| --- | --- | --- | --- | --- | --- |
| Elastic layer | SiO2 | 20 | 1.2 | 70 | 2200 |
| Top electrode | Mo | 14 | 0.1 | 312 | 10200 |
| Piezoelectric layer | AlN | 20 | 0.5 | 345 | 3300 |
| Bottom electrode | Mo | 20 | 0.1 | 312 | 10200 |
| Substrate | Si | - | 600 | 169 | 2329 |

Table S3. Design specifications of our developed linear AlN-PMUT array

| Parameter | Value |
| --- | --- |
| Center Frequency | ~7 MHz |
| Bandwidth | ~60% |
| Number of channels | 64 (64×1) |
| Element pitch | 167 μm |
| Spacing between PMUT | 10 μm |
| Azimuth aperture | 1 cm |
| Elevation aperture | 0.15 cm |
| Focal depth | 2 cm |
| Cable length | 1 m |
| Array area | 0.15 cm2 |
| Cell number in the PMUT array | 5760 |
| Encapsulation layer | Silicone material |
| Fill factor | 50.24 % |

Table S4. Comparison about the ultrasound bandwidth and signal noise ratio between reference PMUTs and piezoelectric bulk ultrasound transducer

| Reference | Transducer type | Frequency (MHz) | Pulse-echo bandwidth (%) | Pulse-echo signal noise ratio (dB) |
| --- | --- | --- | --- | --- |
| Ref[14] | Piezoelectric Bulk Ultrasound Transducer | 2 | 24.6 | 38@4cm |
| Ref[15] | PZT-PMUT | 6 | 18 | 16.6@1.4cm |
| Ref[10] | PZT-PMUT | 11 | 68 | 14.1@6.7mm |
| Ref[13] | AlN-PMUT | 7 | 14.3 | 21.4@5mm |
| This Work | AlN-PMUT | 7 | 63 | 28.2@1cm |

Table S5. Comparison of transducer size and resolution between this work and representative reported studies

|  | Transducer type | Frequency (MHz) | Area (cm2) | Aperture size (cm) | AR* | Lateral resolution (mm) | Axial resolution (mm) | Human imaging | Human imaging depth (cm) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Ref[16] | Piezoelectric Bulk Ultrasound Transducer | 7 | 2 | 1 | 2:1 | 0.38 | 0.23 | Yes | 6 |
| Ref[17] | Piezoelectric Bulk Ultrasound Transducer | 3 | 1 | 1 | 1:1 | 0.56 | 0.5 | Yes | 4 |
| Ref[18] | Piezoelectric Bulk Ultrasound Transducer | 5 | 2.3 | 0.5 | 9:1 | 1.2 | 0.3 | Yes | 2 |
| Ref[19] | PZT-PMUT | 3 | 0.04 | 0.2 | 1:1 | 1.5 | - | No | - |
| Ref[20] | PZT-PMUT | 8.5 | 2.1 | 0.8 | 3.3:1 | 1.8 | 0.61 | No | - |
| Ref[21] | PZT-PMUT | 5 | 0.22 | 0.2 | 6:1 | 1.1 | 0.5 | No | - |
| Ref[22] | AlN-PMUT | 7 | 0.11 | 0.34 | 1:1 | ~1 | ~1 | No | - |
| Ref[23] | AlN-PMUT | 2.7 | 1.4 | 0.55 | 4.7:1 | 5 | - | No | - |
| This work | AlN-PMUT | 7 | 0.15 | 0.15 | 7:1 | 0.26 | 0.25 | Yes | 3 |

*AR: aspect ratio of the transducer array.

Table S6. Comparison of the transducers' surface temperature and power consumption between this work and representative reported studies.

|  | Transducer type | Operating temperature (℃) | Array power (mW) | Continuous monitoring time (h) |
| --- | --- | --- | --- | --- |
| Ref[16] | Piezoelectric Bulk Ultrasound Transducer | ~31.5 | - | 48 |
| Ref[14] | Piezoelectric Bulk Ultrasound Transducer | ~30 | 720 | 1 |
| Ref[24] | Piezoelectric Bulk Ultrasound Transducer | ~36 | - | - |
| Ref[25] | Piezoelectric Bulk Ultrasound Transducer | ~32.5 | 59.1 | - |
| This work | AlN-PMUT | 31.4±0.5 | 15.6 | >5 |

References

- Maadi, M. Zemp, R. J. Self and Mutual Radiation Impedances for Modeling of Multi-Frequency CMUT-Arrays. IEEE T. Ultrason. Ferr. 63, 1441-1454 (2016).

- Lu, Y. Horsley, D. A. Modeling, Fabrication, and Characterization of Piezoelectric Micromachined Ultrasonic Transducer Arrays Based on Cavity SOI Wafers. J. Microelectromech. S. 24, 1142-1149 (2015).

- Ledesma, E., Zamora, I., Yanez, J., Uranga, A. Barniol, N. Single-cell System Using Monolithic PMUTs-on-CMOS to Monitor Fluid Hydrodynamic Properties. Microsyst. Nanoeng. 8, 76 (2022).

- Luo, G.. Fung, S. Wang, Q. et al. High Fill Factor Piezoelectric Micromachined Ultrasonic Transducers on Transparent Substrates, in Proceedings of 2017 19th International Conference on Solid-State Sensors, Actuators and Microsystems (TRANSDUCERS), Kaohsiung, Taiwan, 18-22 June 2017, 1053-1056.

- Wang, Q. Luo, G. L. Kusano, Y. et al. Surface Micromachined  Piezoelectric Micromachined Ultrasonic Transducers with High Fill-factor and High Performance, in Hilton Head 2018 Solid-State Sensors, Actuators and Microsystems Workshop, June 2018.

- Fu, Y. Sun, S. Wang, Z. et al. Piezoelectric Micromachined Ultrasonic Transducer With Superior Acoustic Outputs for Pulse-Echo Imaging Application. IEEE Electr. Device. L. 41, 1572-1575 (2020).

- Xu, X. Ma, Y. Gong, Y. et al. Acoustic Matching Strategy of Piezoelectric Micromachined Ultrasound Transducer Towards Medical Imaging Application. J. Micromech. Microeng. 35, 045010, (2025).

- Chang, C. Firouzi, K. Park, K. K. et al. Acoustic Lens for Capacitive Micromachined Ultrasonic Transducers. J. Micromech. Microeng. 24, 085007, (2014).

- Meinders, J. M. Hoeks, A. P. G. Simultaneous Assessment of Diameter and Pressure Waveforms in the Carotid Artery. Ultrasound Med. Biol. 30, 147-154 (2004).

- Zhao, L. Bao, A. Yang, C. et al. Miniature and Wireless Ultrasonic Imaging System for Wearable Human Carotid Artery Health Monitoring. in Proceedings of the 2025 IEEE 38th International Conference on Microelectromechanical Systems (MEMS), Kaohsiung, Taiwan, 19-23 January 2025, 480-483.

- Yang, H. Liu, H. Chen, J. et al. A PMUT-based Wearable Ultrasound System for Central Blood Pressure Monitoring. in Proceedings of the 2024 IEEE Ultrasonics, Ferroelectrics, and Frequency Control Joint Symposium, Taipei, Taiwan, 22-26 September 2024.

- Luo, G. Kusano, Y. Horsley, D. A. Airborne Piezoelectric Micromachined Ultrasonic Transducers for Long Range Detection. J. Microelectromech. S. 30, 81-89 (2021).

- Zamora, I. Ledesma, E. Uranga, A. et al. Phased Array Based on AlScN Piezoelectric Micromachined Ultrasound Transducers Monolithically Integrated on CMOS. IEEE Electr. Device. L. 43, 1113-1116 (2022).

- Wang, C. Qi, B. Lin, M. et al. Continuous Monitoring of Deep-tissue Haemodynamics with Stretchable Ultrasonic Phased Arrays. Nat. Biomed. Eng. 5, 749-758 (2021).

- Sadeghpour, S. Joshi, S. V. Wang, C. et al. Novel Phased Array Piezoelectric Micromachined Ultrasound Transducers (pMUTs) for Medical Imaging. IEEE OPEN J. Ultrason. 2, 194-202 (2022).

- Wang, C. Chen, X. Wang, L et al. Bioadhesive Ultrasound for Long-Term Continuous Imaging of Diverse Organs. Science 377, 517-523 (2022).

- Hu, H., Ma, Y., Gao, X et al. Stretchable Ultrasonic Arrays for the Three-dimensional Mapping of the Modulus of Deep Tissue. Nat. Biomed. Eng. 7, 1321-1334(2023).

- Pashaei, V. Dehghanzadeh, P. Enwia, G. et al. Flexible Body-Conformal Ultrasound Patches for Image-Guided Neuromodulation. IEEE T. Biomed. Circ. S. 14, 305-318 (2019).

- Zhang, Y., Jin, T., Deng, Y. et al. A Low-Voltage-Driven MEMS Ultrasonic Phased Array Transducer for Fast 3D Volumetric Imaging. Microsyst. Nanoeng. 10, 128 (2024).

- Nistorica, C. Latev, D. Gardner, D. et al. Characterization of a 3D-MEMS Piezoelectric Transducer for Portable Imaging Systems. In Proceedings 2015 IEEE International Ultrasonics Symposium (IUS), Taipei, Taiwan, 21-24 October 2015.

- Dausch, D. E. Gilchrist, K. H. Carlson, J. B. et al. In Vivo Real-time 3-D Intracardiac Echo Using PMUT Arrays. IEEE T. Ultrason. Ferr. 61, 1754-1764 (2014).

- Billen, M. Gao, H. Tabruyn, D. et al. Multi-Frequency Imaging with a CMOS Compatible Scandium Doped Aluminum Nitride PMUT Array. in Proceedings of the 2021 IEEE 34th International Conference on Micro Electro Mechanical Systems (MEMS), Gainesville, FL, USA, 25-29 January 2021.

- Sarafianou, M. Choong, D. S. W. Chen. D. S. H. et al. Long-Range High-Resolution Imaging With Silicon-on-Nothing ScAlN pMUTs. IEEE Sens. J. 23, 24254-24263 (2023).

- Gao, X., Chen, X., Hu, H. et al. A Photoacoustic Patch for Three Dimensional Imaging of Hemoglobin and Core Temperature. Nat. Commun. 13, 7757 (2022).

- Zhou, S., Gao, X., Park, G. et al. Transcranial Volumetric Imaging Using a Conformal Ultrasound Patch. Nature 629, 810-818 (2024).