We provide 8 Matlab scripts that run on Matlab 2015a version and above with the image processing and parallel computing toolboxes. 4 scripts are used for joint deconvolution of dual-view or quad-view data acquired with different diSPIM configurations.

- JointDecon_A.m: for joint deconvolution of conventional diSPIM imaging on glass coverslips (0.8/0.8 NA and 1.1/0.71NA). Data can be visualized from either objective perspective or from the coverslip perspective.

- JointDecon_B.m: for joint deconvolution of symmetric diSPIM on reflective coverslips (0.8/0.8 NA). Data are visualized from the perspective of one of the objective lenses.

- JointDecon_C.m: For joint deconvolution of asymmetric diSPIM on reflective coverslips (e.g. 1.1/0.71 NA). Data are visualized from the perspective of reflective coverslips.

- JointDecon_D.m: For joint deconvolution of quad-view asymmetric diSPIM on reflective coverslips (1.1/0.71 NA). Data are visualized from the perspective of reflective coverslips.

There are 4 accessory scripts that are called after running the main software:

- ReadTifStack.m: for loading tif stacks

- WriteTifStack.m: for writing tif stacks

- ConvFFT3_S.m: for performing 3D convolutions in the Fourier domain using fftn and ifftn functions.

- IlluminationPattern.m: for calculating Gaussian light sheet patterns

Before running the code, users should organize all the input data as tif stacks (Fig. 1). Time-series (4D) diSPIM data are grouped into a single folder:

Time-series view U1:  SPIMA_0, SPIMA_1, SPIMA_2, …. etc, where the number indicates the time points;

Time-series view U2:    SPIMB_0, SPIMB_1, SPIMB_2, …. etc, where the number indicates the time points;

For quad-view data (U1, U2, U3, and U4), users need to set U3 and U4 views as SPIMC_0, and SPIMD_0, etc in the same folder. All these views must be registered and have the same dimensions.

PSF images can be placed in the same folder as U1-U4 data or grouped into another folder, named PSFA for view U1, PSFB for view U2, PSFC for view U3 and PSFD for view U4 (Fig. 2). Users need to generate PSFs (according to the Online Methods and Supplementary Notes 1-2) with image size (in all 3 dimensions) equal to or larger than the size of the data, i.e. U1,U2 etc. The PSFs should be in oriented from the same perspective as their corresponding views, e.g. PSFA with U1, etc.

While running any of the four main code scripts, users need to (1) upload the raw data (see Fig. 1); (2) upload the matching PSF files (see Fig. 2); (3) set parameters (see Fig.3) to

- produce the Gaussian excitation sheet by setting the pixel size of the images (in microns), the beam waist of the light sheet thickness (FWHM value in microns), and the angle of the light sheet in degrees relative to coverslip (e.g., 45 degrees for 0.8/0.8 NA diSPIM, 33 degrees for 1.1/0.71 NA diSPIM).

- define the scanning increment δ in units of pixels.

- define the iteration number (the typical number of iterations we used in this work range from 20-40).

- Set the number of time points (in range form, e.g. 1-10) to be processed.

Deconvolved data will be saved in the same folder that contains the input data U1-U4. For time-series data, they are saved as Decon_0, Decon_1, Decon_2, …. etc, where number indicates the number of time points. Some intermediate files (e.g., excitation light sheet pattern and normalization images) will be saved in the same folder that contains the PSF images. Upon request, we can provide test data, simulated PSFs, and deconvolved output.

Fig. 1. Time series data are grouped into a single folder, and named as SPIMA_0, SPIMA_1, SPIMA_2, …. etc for the first view; SPIMB_0, SPIMB_1, SPIMB_2, …. etc, for the second view. Here the number indicates the number of time points. Upon running one of the main scripts, this dialog box will pop up to guide users as to where to locate the files.

Fig. 2. PSF images are grouped into a single folder, and named as PSFA, PSFB, PSFC and PSFD, corresponding to views U1-U4.  Upon running one of the main scripts, this dialog box will pop up to guide users as to where to locate the folders (choose any one of the PSF images).

Fig. 3. The dialog box for parameter setting, including the pixel size of the images (in microns), the beam waist of the light sheet thickness (FWHM values in microns), the angle of the light sheet relative to the coverslip, the increment δ in pixels that was used to scan samples in the x direction, the iteration number and time points to be jointly deconvolved.