## Supplementary Information for: Fortin et.al.

## A novel art of continuous noninvasive blood pressure measurement

# Supplementary Method 1

## Pseudocode of the filter cascade:

fc = [fc1, fc2, …, fcN]		// Vector of cut-off frequencies from max to min

DSW = [DSW1, DSW2, …, DSWN]		// Vector of “Down Sampling Window sizes”

UC = [UC1, UC2, …, UCN|		// Vector of update coefficients (IIR-parameters)

// whereas UC[n] = 1-e-2πfc[n]fs1nDSW[n]

N: 		Number of cascaded filter stages – here 6

Nrhythm:		Index to cascade stage with fc[Nrhythm] for vRhythm(t) – here 3

vSum [N]: 	Vector of summed v(t)

vN [N]:	Vector of down sampling counters

vFilt[n]:	Vector of filtered signals

CascadeFilter (int n, float v, int N)

vSum[n] = vSum]n] + v

vN[n] = vN[n| + 1

if (vN[n] == DSW[n])

vFilt[n] = vFilt[n] + (vSum[n]/vN[n] – vFilt[n])*UC[n] // *

vSum[n] = vN[n] = 0

if (n<N)

CascadeFilter (n+1, vFilt[n], N)

return

* NOTE: vFilt[n] on the right side of the equation contains the old (n-1) value and is then overwritten with the new value on the left side of the equation

NewData (float v)			// New sample v has been created

// called every 4ms -> 250Hz sampling frequency

CascadeFilter (0, v, N)

vPulse 	= v – vFilt[Nrhythm]

vVCT 	= v – vFilt[N]

vRhythm 	= vFilt[Nrhythm] – vFilt[N]

vdRhythm 	= vFilt[Nrhythm+1] – vFilt[Nrhythm-1]

The cut off frequencies are set to: 			[0.05,  0.05,  0.05,  0.01,  0.01,  0.01]

Down Sampling Window sizes are set to: 		[5,  1,  1,  5,  1,  1]

resulting in sampling frequencies for cascade 0-2 of 50Hz and for cascade 3-5 of 10Hz.

## Supplementary Method 2

## Pseudocode of the basic antiresonance filter elements:

AdaptiveNotch (float x)		// called every 4ms -> 250Hz sampling frequency

// Downsampling from fs=250Hz to fds=5Hz -> DSW = 50

xdsSum += x; DScount++;

if (DScount == DSW)

xds[i] = xdsSum / DScount		XdsSum = DScount = 0

// Select the total frequency area, where resonance can occur:

xbp[i] = FIR-Butterworth Band-pass-Filter

// A high power indicates possible resonance

Pow[i] = UCPow * xbp[i]^2 + (1-UCPow) * Pow[i-1]

// FIR Notch Filter with estimated resonance

yfir[i] = xbp[i] - 2*[i]*Xbp[i-1] + xbp[i-2]

// Adapt the resonance frequency indicator  with Steepest Descent

if (Pow[i] > Threshold)

[i] = [i-1] + µ/2* Yfir[i]* xbp[i-1]

// Calculate periodic length Tr of the resonance

Tr[i] = 2* / (fds * arccos([i]))

// Calculate adaptive gain factor

[i] = Anorm^(Aexp/Tr[i])

// Calculate IIR-coefficients Aiir & Biir

Biir[i] = {1   -2[i]   1} * (1+[i])/2

Aiir[i] = {1   -[i]*(1+[i])   [i]}

yiir[i] = IIR-Notch with Biir[i] / Aiir[i]

return x - xds[i] + yiir[i]		// Replace average with filtered signal

## Supplementary Method 3

## Pseudocode for beat detection:

bool BeatDetected (float v)	// called every 4ms -> 250Hz sampling frequency

- each state-of-the-art beat detector for pulse detection can be used

- we use a modified version of the beat detector concept described in Supplementary Reference 1.

- This open-source algorithm was adapted to the nature of an inverted PPG-signal

returns TRUE, if a beat has been detected, otherwise FALSE

## Supplementary Method 4

## Pseudocode for the open loop phase:

The software algorithm for the open loop phase processes the following steps:

- Continuously measure v(t) as well as pc(t) at a sampling frequency (fs) of 250 Hz.

- Apply different contact pressures pc(t) to the finger. Although a pressure ramp would be suitable, we use the pressure steps implemented in the CNAP HD software.

- Send v(t) to the filter cascade and afterwards to the beat detector.

- Measure the vPulse(t) amplitude after each detected beat and remember those amplitudes as well as their corresponding pc(t).

- At the maximum amplitude, take the corresponding pc(t) as the starting mBP and remember this value as P0.

- Calculate coefficients cBI, cP, cD and cI for VCT by using the maximum vPulse. As v(t) is a dimensionless digital signal, these coefficients will highly depend on the PPG-system and its amplifiers.

- Sets control variables yI, Vn, PI, sumVn to zero and Pn to P0.

## Supplementary Method 5

## Pseudocode of the main control function:

VCT_Controller (float v)		// called every 4ms -> 250Hz sampling frequency

NewData (v)			// Call to the filter cascades and create

// vPulse, vVCT, vRhythm and vdRhythm

// Continuously tracking the physiological rhythms

yPD=cP∙vRhythm+ cD∙vdRhythm 	// proportional and differential part

yI+=cI∙yPD			// the summation builds the integral part

// Reconstruct very slow changing BP changes

Vn += vVCT			// Builds the beat-to-beat integral controller

PI += 4ms

if BeatDetected (vPulse)

Vn = Vn / PI

sumVn += Vn

Pn = P0 - cBI * sumVn - cBP * Vn

Vn = PI = 0

// Calculate contact pressure

pCt=Pn-yI-yPD

// Call antiresonance

pc(t) = AdaptiveNotch (pc(t))

Apply pc(t) to the finger

Supplementary Reference 1

Zong, W., Heldt, T., Moody, G. B., & Mark, R. G. An open-source algorithm to detect onset of arterial blood pressure pulses. Computers in Cardiology 30, 259–262 (2003).