## Supplementary Data 1A
Supplementary Data 1A. The detailed description of 58 intra-experiment metrics. The arrow in the Optimal column indicates the direction for improved performance. The metrics highlighted in blue font are newly added metrics in MSCohort. According to steps in the LC-MS workflow, the 58 intra-experiment metrics can be assigned to seven categories: sample, chromatography, ion source, DIA windows, dynamic sampling, MS1 and MS2 signal, and identification result. Among them, the metrics in DIA windows category are applicable only to data-independent acquisition (DIA) experiments, and the metrics in dynamic sampling category are applicable only to data-dependent acquisition (DDA) experiments, the others are applicable to both DDA and DIA experiments.
(1) 58 intra-experiment metrics
Category | Metrics | Description | Optimal | Purpose/use
Sample | S1. Missed cleavages (n=0) of peptides | The percentage of all identified precursors without missed cleavage sites. MSCohort reads this value from the identification result of the search engine (e.g., Spectronaut) report. | ↑ | Evaluates the enzymatic digestion effect during sample preparation
| S2. Median peptide length | The median of all identified peptide sequence length. MSCohort reads the identified peptide sequence from the identification result of the search engine (e.g., Spectronaut) report and calculates the number of amino acids for each peptide sequence. | ↓ | Evaluates the enzymatic digestion effect during sample preparation
| S3. Median total ion current (TIC) | The median of total ion chromatography (TIC) from raw file. MSCohort sums the peak intensities from ms1 and ms2 file. | ↑ | Evaluates whether the sample loading is sufficient
| S4. Specificity of digestion | The percentage of peptides with C-term and N-term enzymatic specificities. MSCohort reads this value from the identification result of the open search engine (e.g., pFind) report. | ↑ | Evaluates the enzymatic digestion effect during sample preparation
| S5. Frequency of cysteine modification | The percentage of occurrences of the modification Carbamidomethyl[C]. MSCohort reads this value from the identification result of the open search engine (e.g., pFind) report. | ↑ | Evaluates the alkylation effect during sample preparation
| S6. Semi/full trptic peptides | The ratio of the number of peptides with C-term only or N-term only enzymatic specificities to peptides with C-term and N-term enzymatic specificities. MSCohort reads this value from the identification result of the open search engine (e.g., pFind) report. | ↓ | Evaluates the enzymatic digestion effect during sample preparation
| S7. Frequency of contaminants | The summed peak intensity corresponding to contaminant proteins divided by the total summed intensity for all quantified proteins. MSCohort calculates this value from the quantification result of the search engine (e.g., Spectronaut) report. | ↓ | Evaluates the presence/quantity of contaminants during sample harvest or sample preparation
Chromatography | C1. Chromatographic invalid acquiring time | The retention time length from the start to the number of identified MS2 scans per minute more than 100 (user adjustable parameter). MSCohort counts the number of identified MS2 scan per minute based on ms2 file and the identification result of the search engine (e.g., Spectronaut) report. | ↓ | Shorter chromatographic invalid acquiring time indicate better efficient sampling and identification
| C2. Median of full width at half maximum (FWHM) | The median full width at half maximum(FWHM) in retention time (RT) dimension for all peptide precursors identified. The FWHM is calculated on the sum intensity XlC of all fragments associated with certain peptide precursor. MSCohort reads and calculates this value from the result of the search engine (e.g., Spectronaut). | ↓ | Shorter full width at half maximum (FWHM) indicate better chromatographic resolution
| C3. Median of peak width | The median peak width in RT dimension for all peptide precursors identified. MSCohort reads and calculates this value from the result of the search engine (e.g., Spectronaut). | ↓ | Shorter peak widths indicate better chromatographic resolution
| C4. Proportion of precursors with long eluting width | The percentage of peptide precursors with eluting width more than 1 minutes (user adjustable parameter). MSCohort reads the peptide precursors peak width from the result of the search engine (e.g., Spectronaut) and counts the number of precursors with peak width more than 1 minutes. | ↓ | Evaluates the ratio of tailing peptides, with fewer tailing peptides indicating better column separation performance
| C5. Median precursors identified over RT | The median number of precursors identified per minute. MSCohort reads the identified precursors and corresponding RT from the result of the search engine (e.g., Spectronaut) and divides it by per minute. | ↑ | More precursors identified indicate efficient sampling and identification
| C6. Maximum identification rate over RT | The maximum identification rate per minute. The identification rate is calculated by dividing the number of identified MS2 scans by the number of acquired MS2 scans per minute. MSCohort reads the number of acquired MS2 scans from ms2 file and counts the number of identified MS2 scan per minute based on ms2 file and the result of the search engine (e.g., Spectronaut). | ↑ | Higher rates indicate efficient sampling and identification
DIA Windows (for DIA) | W1. Acquired MS2 scans in one cycle | The MS2 DIA windows number in one cycle. MSCohort counts the number of MS2 DIA windows from ms2 file. | ↑ | Morer MS2 scans indicates more sampling
| W2. Median of window size (Da) | The median number of isolation width of MS2 DIA windows. MSCohort reads the isolation width of MS2 DIA windows from ms2 file. | ↓ | Smaller acquisition window size indicates lower spectral complexity of MS2 scans
| W3. Redundant identified precursors/ Identified scan rate | The mean identified precursors number per MS2 scan, which calculated by dividing the number of total identified precursors(including redundant identified precursors) by the number of identified MS2 scans. MSCohort counts the number of identified MS2 scan and identified precursors based on ms2 file and the result of the search engine (e.g., Spectronaut). This metric is used to represent the spectral complexity, that is, how many peptide precursors can be identified per MS2 scan, and the more identified the higher the complexity of the spectrum. | ↑ | Higher redundant identified precursors/ identified scan rate indicates more identified precursors per MS2 scan, and higher spectral complexity of MS2 scans
| W4. Redundant identified precursors/ Identified precursors rate | The mean redundant identification number per precursor, which calculated by dividing the number of total identified precursors(including redundant identified precursors) by the number of unique identified precursors. MSCohort counts the number of identified MS2 scan and identified precursors based on ms2 file and the result of the search engine (e.g., Spectronaut). | ↓ | Lower redundant identified precursors/ identified precursor rate indicates lower duplicate identification rate
| W5. Identified precursors/ identified scan rate | The mean unique identified precursors number per MS2 scan, which calculated by dividing the number of unique identified precursors by the number of identified MS2 scans. MSCohort counts the number of identified MS2 scan and identified precursors based on ms2 file and the result of the search engine (e.g., Spectronaut). | ↑ | Higher identified precursors/ identified scan rate indicates higher utilization rate of MS2 scans
| W6. MS1 data point per peak | The median MS1 data points per peak in RT dimension for all peptide precursors identified in the calibration. MSCohort reads and counts this value from the result of the search engine (e.g., Spectronaut). | ↑ | Evaluates MS1 scan number for quantification
| W7. MS2 data point per peak | The median MS2 data points per peak in RT dimension for all peptide precursors identified in the calibration. MSCohort reads and counts this value from the result of the search engine (e.g., Spectronaut). | ↓ | Evaluates MS2 scan number for quantification
Ion Source | IS1. 1+/2+ ions detected | The number of 1+ peptide features over 2+ peptide features. MSCohort extracts isotope patterns from ms1 file and records the charge information to get this value. | ↓ | High ratios of 1+/2+ peptides may indicate inefficient ionization
| IS2. 3+/2+ ions detected | The number of 3+ peptide features over 2+ peptide features. MSCohort extracts isotope patterns from ms1 file and records the charge information to get this value. | ↓ | High ratios of 3+/2+ peptides may preferentially favor longer peptides
| IS3. 4+/2+ ions detected | The number of 4+ peptide features over 2+ peptide features. MSCohort extracts isotope patterns from ms1 file and records the charge information to get this value. | ↓ | High ratios of 3+/2+ peptides may preferentially favor longer peptides
| IS4. Median precursor m/z | The median m/z value for all identified precursors. MSCohort reads the identified precursors m/z value from the result of the search engine (e.g., Spectronaut). | ↓ | Higher median m/z may correlate with inefficient or partial ionization
Dynamic Sampling (for DDA) | D1. Scan/peptide rate | The average number of MS2 scans per peptide. MSCohort gets the number of identified MS2 scan, the number of identified unique peptides from the identification result of the search engine (e.g., pFind) and takes the ratio of them. | ↓ | Evaluates oversampling
| D2. Oversampling once/twice | The ratio of peptide precursors identified by one MS2 scan divided by number identified by two MS2 scans. MSCohort counts the number of MS2 scan corresponding to each unique peptide based on the identification result of the search engine (e.g., pFind) and gets this value | ↑ | Evaluates oversampling
| D3. Oversampling twice/thrice | The ratio of peptide precursors identified by two MS2 scans divided by number identified by three MS2 scans. MSCohort counts the number of MS2 scan corresponding to each unique peptide based on the identification result of the search engine (e.g., pFind) and gets this value. | ↑ | Evaluates oversampling
| D4. Maximum redundant scans of one peptide | The maximum times a peptide precursor identified by MS2 scan. MSCohort counts the number of MS2 scan corresponding to each unique peptide based on the identification result of the search engine (e.g., pFind) and get the maximum. | ↓ | Evaluates oversampling
| D5. MS1 max/sampled for IDs | The ratio of chorogram profile maximum point to MS1 value at sampling point (1 = sampled at peak maxima). MSCohort constructs chorogram profile for identifications and determine the intensity of the center and sampling points | ↓ | Evaluates position on peak where sampled for peptides of all abundances
MS1 and MS2 Signal | M1. Proportion of MS1 time | The proportion of time taken to acquire MS1 scan. MSCohort sums the scan time of MS1 scan based on ms1 and ms2 files and get this value. | ↓ | Fewer MS1 scan time indicates more sampling
| M2. Median MS1 ion injection time | The median of ion injection time of MS1 scan. MSCohort reads the ion injection time of each acquired MS1 scan from the ms1 file. | ↓ | Lower times indicate a higher abundance of parent ions
| M3. Median MS2 ion injection time | The median of ion injection time of MS2 scan. MSCohort reads the ion injection time of each acquired MS1 scan from the ms2 file. | ↑ | Higher times indicate a higher AGC and abundance of fragment ions
| M4. MS1 cycle time | The median of MS1 cycle time. MSCohort reads the scan time from ms1 file and calculated the cycle time of each acquired MS1 scan. | ↓ | Evaluates cycle time and MS1 scan number for quantification
| M5. MS2 cycle time | The median of MS2 cycle time. MSCohort reads the scan time from ms1 file and calculated the cycle time of each acquired MS2 scan. | ↑ | Evaluates cycle time and MS2 scan number for quantification
| M6. Median peaks intensity of MS1 | The median of sum peak intensities (log10) per MS1 scan. MSCohort reads the peak intensities of each MS1 scan from ms1 file. | ↑ | Higher MS1 peaks intensity may correlate with more signal
| M7. Median peak counts of MS1 | The median number of peaks in an MS1 scan. MSCohort reads the peak number of each MS1 scan from ms1 file. | ↑ | Higher MS1 peak counts may correlate with more signal
| M8. Median peak counts of MS2 | The median number of peaks in an MS2 scan. MSCohort reads the peak counts of each MS2 scan from ms2 file. | ↑ | Higher MS2 peak counts may correlate with more signal
| M9. Median peaks intensity of MS2 | The median of sum peak intensities (log10) per MS2 scan. MSCohort reads the peak intensities of each MS2 scan from ms2 file. | ↑ | Higher MS2 peaks intensity may correlate with more signal
| M10. Identified / detected features | The percentage of identified isotope patterns (features) divided by detected isotope patterns (features). MSCohort extracts detected isotope patterns from ms1 file, identified isotope patterns corresponding to identified precursors m/z, and make audgment whether features are identified or not and then get this value. | ↑ | Higher identified / detected features indicates efficient sampling and identification
| M11. Signal to noise（S/N）for precursors | The signal to noise ratio of the LC-MS peak for a given peptide precursor.The signal is calculated as the maximum intensity of the fragment sum X1C within the peak boundaries.The noise is calculated as the average fragment sum XIC intensity outside the peak boundaries (EG.Noise). MSCohort reads the values from the result of the search engine (e.g., Spectronaut). | ↑ | Higher S/N may correlate with higher signal discrimination
| M12. Median of MS1 raw mass accuracy | The median of delta mass between the monoisotopic theoretical and the measured m/z of precursors, ppm = ((theoretical - measured) / theoretical)*106. MSCohort reads the values from the result of the search engine (e.g., Spectronaut). | ↓ | Measures the accuracy of the identificatons
| M13. Median of MS2 raw mass accuracy | The median of delta mass between the monoisotopic theoretical and the measured m/z of fragments, ppm = ((theoretical - measured) / theoretical)*106. MSCohort reads the values from the result of the search engine (e.g., Spectronaut). | ↓ | Measures the accuracy of the identificatons
| M14. Median intensities for precursors | The median of precursors quantity depending on the setting (MS1 or MS2). MSCohort reads the values from the result of the search engine (e.g., Spectronaut). | ↑ | Evaluates the intensity for quantified precursors
| M15. Median intensities for peptides | The median of peptide group quantity in this run calculated as defined in the quantification section of the analysis settings. MSCohort reads the values from the result of the search engine (e.g., Spectronaut). | ↑ | Evaluates the intensity of quantified peptides
| M16. Median intensities for protein groups | The median of protein group quantity in this run calculated as defined in the quantification section of the analysis settings. MSCohort reads the values from the result of the search engine (e.g., Spectronaut). | ↑ | Evaluates the intensity of quantified protein groups
| M17. Acquired MS1 scans | The mean acquired MS1 scans per minute. MSCohort reads the acquired MS1 scans from ms1 file. | ↓ | Fewer MS1 scans indicates more sampling
| M18. Acquired MS2 scans | The mean acquired MS2 scans per minute. MSCohort reads the acquired MS2 scans from ms2 file. | ↑ | Morer MS2 scans indicates more sampling
| M19. MS2 identification rate | The proportion of identified MS2 scans divided by acquired MS2 scans. MSCohort reads the number of acquired MS2 scans from ms2 file and counts the number of identified MS2 scan based on ms2 file and the result of the search engine (e.g., Spectronaut). | ↑ | Higher rate indicates efficient sampling and identification
| M20. Maximum identification rate over M/Z range | The identification rate over each MS2 DIA window. MSCohort counts the number of acquired MS2 scans for each MS2 DIA window from ms2 file, and counts the number of identified MS2 scan based on ms2 file and the result of the search engine (e.g., Spectronaut). | ↑ | Higher rate indicates efficient sampling and identification
| M21. Mean number of tags | The average number of 5-mer sequence tags extracted per MS2 scan. Based on ms2 file, MSCohort extracts 5-mer sequences for each MS2 scan. | ↑ | Higher number of tags indicate better MS2 scan for identification
| M22. Median PIF | Short for “Median of Parent lon Fraction", indicates the fraction that target peak makes up of the total intensity in the inclusion window. Based on ms1 files, MSCohort calculates PIF for identifications and gets median value. | ↑ | Higher PIF indicate better MS1 signal for identification
Identification Result | ID1. Number of Identified MS2 scans | The number of identified MS2 scans. MSCohort reads the values from ms2 file and the result of the search engine (e.g., Spectronaut). | ↑ | Total identified MS2 scans correlate with high levels of peptide signals, performance
| ID2: Number of precursors | The number of identified precursors. MSCohort reads the values from the result of the search engine (e.g., Spectronaut). | ↑ | A good overall performance measure
| ID3: Number of peptides | The number of identified peptides. MSCohort reads the values from the result of the search engine (e.g., Spectronaut). | ↑ | A good overall performance measure
| ID4. Number of protein groups | The number of identified protein groups. MSCohort reads the values from the result of the search engine (e.g., Spectronaut). | ↑ | A good overall performance measure
| ID5. Number of proteins | The number of identified proteins. MSCohort reads the values from the result of the search engine (e.g., Spectronaut). | ↑ | A good overall performance measure
| ID6. AVG peptides per protein group | The average peptides per protein group. MSCohort reads and calculates the values from the result of the search engine (e.g., Spectronaut). | ↑ | A good overall performance measure
| ID7. AVG precursors per protein group | The average precursors per protein group. MSCohort reads and calculates the values from the result of the search engine (e.g., Spectronaut). | ↑ | A good overall performance measure

## Supplementary Data 1B
Supplementary Data 1B. The detailed description of 23 inter-experiment metrics.
23 inter-experiment metrics
Category | Metrics | Description
Sample | SP1. Contaminants——erythrocytes (user-defined) | The summed intensity of erythrocytes contaminant proteins is divided by the total summed intensity for all quantified proteins in each run, yielding the ratio of contaminant proteins. The erythrocytes contaminants occurs by a different extent of hematuria or hemolysis caused by kidney function or systemic disorders during urine harvesting. The erythrocyte quality marker panel was reported from previous study (EMBO Mol Med. 2019 Nov 7;11(11):e10427). The contaminant protein panel was user-defined, users can change the contaminant protein list according to the actual sample type. MSCohort extracts and calculates the ratio for each run,and compares the consistency of the ratios across all runs. A run was flagged for potential contamination if the ratio differed more than two standard deviations from the median of all runs within the cohort.
| SP2. Contaminants——cellular debris (user-defined) | The summed intensity of cellular debris contaminant proteins is divided by the total summed intensity for all quantified proteins in each run, yielding the ratio of contaminant proteins. Insufficient removal of cells and cellular debris from urine will lead to increased detection of intracellular proteins with a high sample-to-sample variability compared to regularly secreted urinary proteins. The cellular debris quality marker panel was reported from previous study (EMBO Mol Med. 2019 Nov 7;11(11):e10427). The contaminant protein panel was user-defined, users can change the contaminant protein list according to the actual sample type. MSCohort extracts and calculates the ratio for each run,and compares the consistency of the ratios across all runs. A run was flagged for potential contamination if the ratio differed more than two standard deviations from the median of all runs within the cohort.
| SP3. Contaminants——serum (user-defined) | The summed intensity of serum contaminant proteins is divided by the total summed intensity for all quantified proteins in each run, yielding the ratio of contaminant proteins. Proteinuria occurs due to abnormalities in kidney function or systemic disorders, resulting in the leakage of serum proteins into the urine. This can lead to increased detection of serum proteins with high sample-to-sample variability. The serum quality marker panel was reported from previous study (Mol Syst Biol. 2017 Sep 26;13(9):942). The contaminant protein panel was user-defined, users can change the contaminant protein list according to the actual sample type. MSCohort extracts and calculates the ratio for each run,and compares the consistency of the ratios across all runs. A run was flagged for potential contamination if the ratio differed more than two standard deviations from the median of all runs within the cohort.
Chromatography | LC1. DeltaRT (predicted-measured) | The difference between the predicted retention time (RT) and the measured mean apex RT is computed for each precursor. MSCohort reads this value from the result of search engine (e.g., Spectronaut). MSCohort extracts the DeltaRT value for each run,and compares the consistency of the DeltaRT value across all runs. A run was flagged for potential abnormal if the DeltaRT value differed more than two standard deviations from the median of all runs within the cohort.
| LC2. RT relative deviation | The RT relative deviation between each run and the first run for each precursor. MSCohort reads the mean apex RT for each precursor from the result of search engine (e.g., Spectronaut) and calculates the RT deviation between each run and the first run. A run was flagged for potential abnormal if the RT relative deviation value differed more than two standard deviations from the median of all runs within the cohort.
Precursor-level Results | MS1. The ratio of missing values of precursors | The missing values of precursors = (total number of identified precursors in all run - number of identified precursors in one run) / total number of identified precursors in all run. MSCohort reads the identified precursors number from the result of search engine (e.g., Spectronaut) and calculates the ratio of missing values for each run. A run was flagged for potential abnormal if the ratio differed more than two standard deviations from the median of all runs within the cohort.
| MS2. Median of precursor intensity | The median intensity of quantified precursors for each run. MSCohort reads the precursors intensity from the result of search engine (e.g., Spectronaut) and calculates the median value for each run. A run was flagged for potential abnormal if the median intensity differed more than two standard deviations from the median of all runs within the cohort.
| MS3. IQR of precursor intensity | The interquartile range (IQR) of the intensity of quantified precursors for each run. MSCohort reads the precursors intensity from the result of search engine (e.g., Spectronaut) and calculates the IQR (Q3-Q1) of the intensity for each run. A run was flagged for potential abnormal if the IQR differed more than two standard deviations from the median of all runs within the cohort.
| MS4. Robust Dev of precursor intensity | The median of robust standard deviation computed based on log2-transformed quantitation ratios between each run and all other runs. Robust standard deviation = (the 84.13% percentiles of log2-transformed quantitation ratios of two runs - 15.87% percentiles of log2-transformed quantitation ratios of two runs) / 2. MSCohort reads the precursors intensity from the result of search engine (e.g., Spectronaut) and calculates the median of the robust standard deviation for each run. A run was flagged for potential abnormal if the robust standard deviation differed more than two standard deviations from the median of all runs within the cohort.
| MS5. Pearson correlation of precursor intensity | The median of Pearson correlation coefficients between each run and all other runs. MSCohort reads the precursors intensity from the result of search engine (e.g., Spectronaut) and calculates the median pearson correlation coefficients for each run. A run was flagged for potential abnormal if the median pearson correlation coefficients differed more than two standard deviations from the median of all runs within the cohort.
| MS6. Normalization factor of precursor intensity | The normalization factor of precursor intensity for each run based on normalization algorithm. MSCohort reads the precursors intensity (setting) from the result of search engine (e.g., Spectronaut) and data normalization was performed for all runs. A run was flagged for potential abnormal if the normalization factor differed more than two standard deviations from the median of all runs within the cohort.
Peptide-level Results | MS7. The ratio of missing values of peptides | The missing values of peptides = (total number of identified peptides in all run - number of identified peptides in one run) / total number of identified peptides in all run. MSCohort reads the identified peptides number from the result of search engine (e.g., Spectronaut) and calculates the ratio of missing values for each run. A run was flagged for potential abnormal if the ratio differed more than two standard deviations from the median of all runs within the cohort.
| MS8. Median of peptide intensity | The median intensity of quantified peptides for each run. MSCohort reads the peptides intensity from the result of search engine (e.g., Spectronaut) and calculates the median value for each run. A run was flagged for potential abnormal if the median intensity differed more than two standard deviations from the median of all runs within the cohort.
| MS9. IQR of peptide intensity | The interquartile range (IQR) of the intensity of quantified peptides for each run. MSCohort reads the peptides intensity from the result of search engine (e.g., Spectronaut) and calculates the IQR (Q3-Q1) of the intensity for each run. A run was flagged for potential abnormal if the IQR differed more than two standard deviations from the median of all runs within the cohort.
| MS10. Robust Dev of peptide intensity | The median of robust standard deviation computed based on log2-transformed quantitation ratios between each run and all other runs. Robust standard deviation = (the 84.13% percentiles of log2-transformed quantitation ratios of two runs - 15.87% percentiles of log2-transformed quantitation ratios of two runs) / 2. MSCohort reads the peptides intensity from the result of search engine (e.g., Spectronaut) and calculates the median of the robust standard deviation for each run. A run was flagged for potential abnormal if the robust standard deviation differed more than two standard deviations from the median of all runs within the cohort.
| MS11. Pearson correlation of peptide intensity | The median of Pearson correlation coefficients between each run and the other runs. MSCohort reads the peptides intensity from the result of search engine (e.g., Spectronaut) and calculates median pearson correlation coefficients for each run. A run was flagged for potential abnormal if the median pearson correlation coefficients differed more than two standard deviations from the median of all runs within the cohort.
| MS12. Normalization factor of peptide intensity | The normalization factor of peptides intensity for each run based on normalization algorithm. MSCohort reads the peptides intensity from the result of search engine (e.g., Spectronaut) and data normalization was performed for all runs. A run was flagged for potential abnormal if the normalization factor differed more than two standard deviations from the median of all runs within the cohort.
Protein group-level Results | MS13. The ratio of missing values of proteins | The missing values of protein groups = (total number of identified protein groups in all run - number of identified protein groups in one run) / total number of identified protein groups in all run. MSCohort reads the identified protein groups number from the result of search engine (e.g., Spectronaut) and calculates the ratio of missing values for each run. A run was flagged for potential abnormal if the ratio differed more than two standard deviations from the median of all runs within the cohort.
| MS14. Median of protein intensity | The median intensity of quantified protein groups for each run. MSCohort reads the protein groups intensity (setting) from the result of search engine (e.g., Spectronaut) and calculates the median value for each run. A run was flagged for potential abnormal if the median intensity differed more than two standard deviations from the median of all runs within the cohort.
| MS15. IQR of protein intensity | The interquartile range (IQR) of the intensity of quantified protein groups for each run. MSCohort reads the protein groups intensity (setting) from the result of search engine (e.g., Spectronaut) and calculates the IQR (Q3-Q1) of the intensity for each run. A run was flagged for potential abnormal if the IQR differed more than two standard deviations from the median of all runs within the cohort.
| MS16. Robust Dev of protein intensity | The median of robust standard deviation computed based on log2-transformed quantitation ratios between each run and all other runs. Robust standard deviation = (the 84.13% percentiles of log2-transformed quantitation ratios of two runs - 15.87% percentiles of log2-transformed quantitation ratios of two runs) / 2. MSCohort reads the proteins intensity from the result of search engine (e.g., Spectronaut) and calculates the median of the robust standard deviation for each run. A run was flagged for potential abnormal if the robust standard deviation differed more than two standard deviations from the median of all runs within the cohort.
| MS17. Pearson correlation of protein intensity | The median of Pearson correlation coefficients between each run and the other runs. MSCohort reads the protein groups intensity (setting) from the result of search engine (e.g., Spectronaut) and calculates pearson correlation coefficients for each run. A run was flagged for potential abnormal if the median pearson correlation coefficients differed more than two standard deviations from the median of all runs within the cohort.
| MS18. Normalization factor of protein intensity | The normalization factor of protein groups intensity for each run based on normalization algorithm. MSCohort reads the protein groups intensity (setting) from the result of search engine (e.g., Spectronaut) and data normalization was performed for all runs. A run was flagged for potential abnormal if the normalization factor differed more than two standard deviations from the median of all runs within the cohort.

## Supplementary Data 1C
Supplementary Data 1C. The detailed information reported from Spectronaut software.
Run
R.Run Date
R.Gradient Length [min]
R.FileName
Protein Group
PG.ProteinGroups
PG.Qvalue
PG.Quantity
Peptide
PEP.NrOfMissedCleavages
PEP.StrippedSequence
PEP.Quantity
Elution Group
EG.IsDecoy
EG.ModifiedPeptide
EG.PrecursorId
EG.Qvalue
EG.ApexRT
EG.DatapointsPerPeak
EG.DatapointsPerPeak (MS1)
EG.DeltaRT
EG.EndRT
EG.FWHM
EG.PeakWidth
EG.StartRT
EG.SignalToNoise
EG.TotalQuantity (Settings)
Fragment Group
FG.Charge
FG.PrecMz
FG.PrecWindow
FG.PrecWindowNumber
FG.CalibratedMassAccuracy (PPM)
FG.RawMassAccuracy (PPM)
Fragment
F.RawMassAccuracy (PPM)
F.CalibratedMassAccuracy (PPM)