## Supplementary Data 2A
Supplementary Data 2A. The overview of the LC-MS conditions for 20 LC-MS platforms.
Abbreviation | LC-MS platform number | Mass spectrometer | Liquid Chromatography | Acquisition time | Loading Amount (μg) | Flow rate (nL/min) | Analytical column | DIA windows
Ten LC-MS platforms used multiple different LC and MS conditions (termed 10 platforms without LC-SOP for short)
M01-QEHF | M01 | Q Exactive HF | Easy-nanoLC 1200 | 30 minutes | 1 μg | 600 | 15cm×150μm, 1.9μm | 60
M02-tims Pro | M02 | timsTOF Pro | nanoElute | 30 minutes | 0.5 μg | 300 | 25cm×75μm, 1.6μm | 32
M03-E480F | M03 | Orbitrap Exploris 480 with FAIMS | Easy-nanoLC 1200 | 30 minutes | 1.5 μg | 350 | 15cm×75μm, 1.9μm | 72
M04-E480F | M04 | Orbitrap Exploris 480 with FAIMS | Easy-nanoLC 1200 | 30 minutes | 1.5 μg | 350 | 25cm×75μm, 2μm | 160
M05-Eclipse | M05 | Orbitrap Eclipse | Easy-nanoLC 1200 | 30 minutes | 1.5 μg | 400 | 20cm×75μm, 1.9μm | 24
M06-Eclipse | M06 | Orbitrap Eclipse | UltiMate 3000 | 30 minutes | 1 μg | 300 | 25cm×75μm, 2μm | 31
M07-Eclipse | M07 | Orbitrap Eclipse | EasyLC 1200 | 30 minutes | 1.5 μg | 300 | 25cm×75μm, 1.6μm | 36
M08-Eclipse | M08 | Orbitrap Eclipse | UltiMate 3000 | 30 minutes | 1 μg | 500 | 15cm×75μm, 2μm | 75
M09-E240 | M09 | Orbitrap Exploris 240 | Easy-nanoLC 1000 | 30 minutes | 2 μg | 300 | 25cm×75μm, 2μm | 40
M10-Zeno7600 | M10 | ZenoTOF 7600 | Waters M Class | 30 minutes | 2.5 μg | 5000 | 15cm×0.3mm, 2.1 μm | 80
Ten LC-MS platforms employed a uniform LC condition and consistent MS parameters for the same type of instrument (termed 10 platforms with LC-SOP for short)
U01-QEHF | U01 | Q Exactive HF | Easy-nanoLC 1000 | 30 minutes | 2 μg | 500 | 50cm×50μm, Monolithic Column | 80
U02-Lumos | U02 | Orbitrap Fusion Lumos | Easy-nanoLC 1000 | 30 minutes | 2 μg | 500 | 50cm×50μm, Monolithic Column | 80
U03-Lumos | U03 | Orbitrap Fusion Lumos | Easy-nanoLC 1000 | 30 minutes | 2 μg | 500 | 50cm×50μm, Monolithic Column | 80
U04-E480 | U04 | Orbitrap Exploris 480 | UltiMate 3000 | 30 minutes | 2 μg | 500 | 50cm×50μm, Monolithic Column | 80
U05-E480 | U05 | Orbitrap Exploris 480 | UltiMate 3000 | 30 minutes | 2 μg | 500 | 50cm×50μm, Monolithic Column | 80
U06-E480 | U06 | Orbitrap Exploris 480 | UltiMate 3000 | 30 minutes | 2 μg | 500 | 50cm×50μm, Monolithic Column | 80
U07-E240 | U07 | Orbitrap Exploris 240 | UltiMate 3000 | 30 minutes | 2 μg | 500 | 50cm×50μm, Monolithic Column | 80
U08-tims Pro2 | U08 | timsTOF Pro 2 | UltiMate 3000 | 30 minutes | 0.5 μg | 500 | 50cm×50μm, Monolithic Column | 50
U09-tims Pro2 | U09 | timsTOF Pro 2 | nanoElute | 30 minutes | 0.5 μg | 500 | 50cm×50μm, Monolithic Column | 50
U10-Zeno7600 | U10 | ZenoTOF 7600 | Waters M Class | 30 minutes | 2 μg | 500 | 50cm×50μm, Monolithic Column | 80

## Supplementary Data 2B
Supplementary Data 2B. LC-MS parameters for Thermo Orbitrap data.
| Lab name | M01-QEHF | M03-E480F | M04-E480F | M05-Eclipse | M06-Eclipse | M07-Eclipse | M08-Eclipse | M09-E240 | U01-QEHF | U02-Lumos | U03-Lumos | U04-E480 | U05-E480 | U06-E480 | U07-E240
| LC type | Easy-nanoLC 1200 | Easy-nanoLC 1200 | Easy-nanoLC 1200 | Easy-nanoLC 1200 | UltiMate 3000 | Easy-nanoLC 1200 | UltiMate 3000 | Easy-nanoLC 1000 | Easy-nanoLC 1000 | Easy-nanoLC 1000 | Easy-nanoLC 1000 | UltiMate 3000 | UltiMate 3000 | UltiMate 3000 | UltiMate 3000
| MS type | Q Exactive HF | Orbitrap Exploris 480 with FAIMS | Orbitrap Exploris 480 with FAIMS | Orbitrap Eclipse | Orbitrap Eclipse | Orbitrap Eclipse | Orbitrap Eclipse | Orbitrap Exploris 240 | Q Exactive HF | Orbitrap Fusion Lumos | Orbitrap Fusion Lumos | Orbitrap Exploris 480 | Orbitrap Exploris 480 | Orbitrap Exploris 480 | Orbitrap Exploris 240
LC method | Loading amount | 1 μg / 5 μL | 1.5 μg / 1.5 μL | 1.5 μg / 1.5 μL | 1.5 μg / 1.5 μL | 1 μg / 1 μL | 1.5 μg / 1.5 μL | 1 μg / 1 μL | 2 μg / 2 μL | 2 μg / 2 μL | 2 μg / 2 μL | 2 μg / 2 μL | 2 μg / 2 μL | 2 μg / 2 μL | 2 μg / 2 μL | 2 μg / 2 μL
| Precolumn | 2 cm × 150 μm, 3 μm (home-made) | No | Thermo 2 cm × 75 μm, 3 μm | No | Thermo 2 cm × 75 μm, 3 μm | No | No | Thermo 2 cm × 75 μm, 3 μm | no | no | no | no | no | no | no
| Analytical column | C18, 15 cm × 150 μm, 1.9 μm (home-made) | C18, 15 cm × 75 μm, 1.9μm (home-made) | C18, 25 cm × 75 μm, 2 μm, Thermo | C18, 20 cm × 75 μm, 1.9 μm (home-made) | Thermo C18, 25 cm × 75 μm, 2 μm | C18, 25 cm × 75 μm, 1.6 μm, ionopticks | C18, 15 cm × 75 μm, 2 μm, Thermo | C18, 15 cm × 75 μm, 2 μm, Thermo | Monolithic Column, 50 cm × 50 μm | Monolithic Column, 50 cm × 50 μm | Monolithic Column, 50 cm × 50 μm | Monolithic Column, 50 cm × 50 μm | Monolithic Column, 50 cm × 50 μm | Monolithic Column, 50 cm × 50 μm | Monolithic Column, 50 cm × 50 μm
| Mobile phase A | 100% H2O/0.1% FA | 2%ACN/98% H2O/0.1% FA | 100% H2O/0.1% FA | 2%ACN/98% H2O/0.1% FA | 100% H2O/0.1% FA | 100% H2O/0.1% FA | 100% H2O/0.1% FA | 100% H2O/0.1% FA | 100% H2O/0.1% FA | 100% H2O/0.1% FA | 100% H2O/0.1% FA | 100% H2O/0.1% FA | 100% H2O/0.1% FA | 100% H2O/0.1% FA | 100% H2O/0.1% FA
| Mobile phase B | 80% ACN/20% H2O/0.1% FA | 80% ACN/20% H2O/0.1% FA | 80% ACN/20% H2O/0.1% FA | 80% ACN/20% H2O/0.1% FA | 80% ACN/20% H2O/0.1% FA | 80% ACN/20% H2O/0.1% FA | 80% ACN/20% H2O/0.1% FA | 80% ACN/20% H2O/0.1% FA | 80% ACN/20% H2O/0.1% FA | 80% ACN/20% H2O/0.1% FA | 80% ACN/20% H2O/0.1% FA | 80% ACN/20% H2O/0.1% FA | 80% ACN/20% H2O/0.1% FA | 80% ACN/20% H2O/0.1% FA | 80% ACN/20% H2O/0.1% FA
| Flow rate（nL/min） | 600 | 350 | 350 | 400 | 500, 300 | 300 | 500 | 300 | 500 | 500 | 500 | 500 | 500 | 500 | 500
| Gradient | 40min | 30min | 30 min | 30min | 45min | 33min | 30min | 30min | 30min | 30min | 30min | 30min | 30min | 30min | 30min
| Mobile phase | 0min, 5%B; 4min, 12%B 22min, 28%B 24min, 42%B 27min, 95%B 30min, 95%B 32min, 5%B 33min, 5%B 35min, 95% 40min, 95% | 0min, 12%B; 20min, 30%B 23min, 38%B 25min, 95%B 30min, 95%B | 0min, 5%B; 23min, 25%B 26.3min, 30%B 26.4min, 90%B 27.4min, 90%B 27.5min, 2%B 30min, 2%B | 0min, 5%B; 22min, 28%B 25min, 46%B 26.3min, 95%B 30min, 95%B | 0min, 5%B, 500nL; 8min, 7%B, 300nL, Onduration 32min, 22%B, 300nL 34min, 34%B, 300nL 36.1min, 99%B, 300nL 38min, 99%B, 300nL 38.1min, 5%B, 500nL 45min, 5%B, 500nL | 0min, 5%B, 300nL; 23min, 28%B, 300nL 25min, 34%B, 300nL 26min, 45%B, 300nL 27min, 95%B, 305nL 33min, 95%B, 320nL | 0min, 1%B; 1min, 5%B 16min, 15%B 25min, 25%B 27min, 35%B 27.5min, 99%B 30min, 99%B | 0min, 3%B; 2min, 15%B 19min, 30%B 21min, 40%B 25min, 60%B 28min, 98%B 30min, 98%B | 0min, 5%B; 22min, 20%B 25.3min, 30%B 25.4min, 90%B 26.4min, 90%B 26.5min, 2%B 30min, 2%B | 0min, 5%B; 22min, 20%B 25.3min, 30%B 25.4min, 90%B 26.4min, 90%B 26.5min, 2%B 30min, 2%B | 0min, 5%B; 22min, 20%B 25.3min, 30%B 25.4min, 90%B 26.4min, 90%B 26.5min, 2%B 30min, 2%B | 0min, 5%B; 22min, 20%B 25.3min, 30%B 25.4min, 90%B 26.4min, 90%B 26.5min, 2%B 30min, 2%B | 0min, 5%B; 22min, 20%B 25.3min, 30%B 25.4min, 90%B 26.4min, 90%B 26.5min, 2%B 30min, 2%B | 0min, 5%B; 22min, 20%B 25.3min, 30%B 25.4min, 90%B 26.4min, 90%B 26.5min, 2%B 30min, 2%B | 0min, 5%B; 22min, 20%B 25.3min, 30%B 25.4min, 90%B 26.4min, 90%B 26.5min, 2%B 30min, 2%B
DDA method | Acquisition time | 30min | 30min | 30min | 30min | 30min | 30min | 30min | 30min | 30min | 30min | 30min | 30min | 30min | 30min | 30min
| Acquisition mode | DDA | DDA | DDA | DDA | DDA | DDA | DDA | DDA | DDA | DDA | DDA | DDA | DDA | DDA | DDA
| Advanced Peak Determination | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE
| FAIMS Mode | Not installed | Standard Resolution(-45, -65) | Standard Resolution(-45, -65) | Not installed | Not installed | Not installed | Not installed | Not installed | Not installed | Not installed | Not installed | Not installed | Not installed | Not installed | Not installed
| Detection type for full MS | Orbitrap | Orbitrap | Orbitrap | Orbitrap | Orbitrap | Orbitrap | Orbitrap | Orbitrap | Orbitrap | Orbitrap | Orbitrap | Orbitrap | Orbitrap | Orbitrap | Orbitrap
| Resolution for full MS | 120000 | 60000 | 60000 | 60000 | 60000 | 60000 | 60000 | 120000 | 60000 | 60000 | 60000 | 60000 | 60000 | 60000 | 60000
| Scan range for full MS | 348-1202 | 400-1400 | 350 - 1500 | 400-1200 | 350 - 1200 | 350 - 1500 | 375 - 1500 | 350 - 1500 | 350 - 1200 | 350 - 1200 | 350 - 1200 | 350 - 1200 | 350 - 1200 | 350 - 1200 | 350 - 1200
| Maximum injection time for full MS | 80ms | 20ms | 20ms | 50ms | 20ms | 50ms | 50ms | 45ms | 20ms | 20ms | 20ms | 20ms | 20ms | 20ms | 20ms
| AGC target for full MS | 3000000 | custom | custom | 400000 | 400000 | 1000000 | 400000 | custom | custom | custom | custom | custom | custom | custom | custom
| Normalized AGC target for full MS |  | 300% | 300% | 100% | 100% | 250% | 100% | 300% | 250% | 250% | 250% | 250% | 250% | 250% | 250%
| Minimum intensity | 45000 | 25000 | 5000 | 50000 | 5000 | 5000 | 25000 | 20000 | 5000 | 5000 | 5000 | 5000 | 5000 | 5000 | 5000
| Top N | 20
| cycle time |  | 1s | 1s | 2s | 1s | 0.8s | 2s | 3s | 1s | 1s | 1s | 1s | 1s | 1s | 1s
| Include charge state | 2-6 | 2-7 | 2-6 | 2-7 | 2-7 | 2-5 | 2-7 | 2-7 | 2-6 | 2-6 | 2-6 | 2-6 | 2-6 | 2-6 | 2-6
| Exclusion duration | 12s | 30s | 45s | 45s | 15s | 15s | 45s | 50s | 15s | 15s | 15s | 15s | 15s | 15s | 15s
| Data dependent mode | TopN | Cycle time | Cycle time | Cycle time | Cycle time | Cycle time | Cycle time | Cycle time | Cycle time | Cycle time | Cycle time | Cycle time | Cycle time | Cycle time | Cycle time
| Activation type | HCD | HCD | HCD | HCD | HCD | HCD | HCD | HCD | HCD | HCD | HCD | HCD | HCD | HCD | HCD
| Isolation window(m/z) | 1.6 | 1.6 | 1.6 | 1.6 | 1.6 | 1.6 | 1.6 | 1.6 | 1.6 | 1.6 | 1.6 | 1.6 | 1.6 | 1.6 | 1.6
| Normalized collision energy (%) | 28 | 30 | 30 | 30 | 30 | 32 | 30 | 30 | 32 | 32 | 32 | 32 | 32 | 32 | 32
| Detection type for MS2 | Orbitrap | Orbitrap | Orbitrap | Orbitrap | Orbitrap | Orbitrap | Orbitrap | Orbitrap | Orbitrap | Orbitrap | Orbitrap | Orbitrap | Orbitrap | Orbitrap | Orbitrap
| Resolution for MS2 | 15000 | 15000 | 15000 | 15000 | 30000 | 15000 | 15000 | 15000 | 15000 | 15000 | 15000 | 15000 | 15000 | 15000 | 15000
| Fixed first mass for MS2 | 100 | 110 | 110 | 110 | 110 | 150 | 110 | 120 | 110 | 110 | 110 | 110 | 110 | 110 | 110
| AGC target for MS2 | 20000 | custom | custom | 50000 | 50000 | 37500 | 50000 | custom | custom | custom | custom | custom | custom | custom | custom
| Normalized AGC target for MS2 |  | 75% | 75% | 100% | 100% | 75% | 100% | 100% | 75% | 75% | 75% | 75% | 75% | 75% | 75%
| Maximum injection time for MS2(ms) | 22ms | 30ms | 30ms | 45ms | 54ms | 22ms | 45ms | 22ms | 22ms | 22ms | 22ms | 22ms | 22ms | 22ms | 22ms
DIA method | Acquisition mode | DIA | DIA | DIA | DIA | DIA | DIA | DIA | DIA | DIA | DIA | DIA | DIA | DIA | DIA | DIA
| Advanced Peak Determination | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE
| FAIMS Mode | Not installed | Standard Resolution(-45, -65) | Standard Resolution(-45, -65) | Not installed | Not installed | Not installed | Not installed | Not installed | Not installed | Not installed | Not installed | Not installed | Not installed | Not installed | Not installed
| Detection type for full MS | Orbitrap | Orbitrap | Orbitrap | Orbitrap | Orbitrap | Orbitrap | Orbitrap | Orbitrap | Orbitrap | Orbitrap | Orbitrap | Orbitrap | Orbitrap | Orbitrap | Orbitrap
| Resolution for full MS | 120000 | 60000 | 120000 | 60000 | 120000 | 50000 | 120000 | 120000 | 120000 | 120000 | 120000 | 120000 | 120000 | 120000 | 120000
| Scan range for full MS | 348-1202 | 400-1300 | 350-1200 | 400-1200 | 350-1200 | 350-1500 | 375-1500 | 400-1200 | 350-1200 | 350-1200 | 350-1200 | 350-1200 | 350-1200 | 350-1200 | 350-1200
| Maximum injection time for full MS | 40ms | 20ms | 50ms | 50ms | 50ms | 50ms | 50ms | 50ms | 50ms | 50ms | 50ms | 50ms | 50ms | 50ms | 50ms
| AGC target for full MS | 3000000 | custom | custom | 400000 | 1200000 | 500000 | 400000 | custom | 1000000 | 1200000 | 1000000 | custom | custom | custom | custom
| Normalized AGC target for full MS(%) |  | 300% | 300% | 100% | 300% | 125% | 100% | 300% |  | 300% | 300% | 300% | 300% | 300% | 300%
| cycle time(s) |  | 3s | 3s | 3s | 3s | 3s | 3s | 3s | 3s | 3s | 3s | 3s | 3s | 3s | 3s
| Activation type | HCD | HCD | HCD | HCD | HCD | HCD | HCD | HCD | HCD | HCD | HCD | HCD | HCD | HCD | HCD
| Isolation window(m/z) | custom | custom | custom | custom | custom | custom | 6.7 | 20 | custom | custom | custom | custom | custom | custom | custom
| Normalized collision energy (%) | 28 | 28, 30, 32 | 30 | 30 | 30 | 32 | 32 | 30 | 30 | 32 | 32 | 30 | 30 | 30 | 30
| Detection type for MS2 | Orbitrap | Orbitrap | Orbitrap | Orbitrap | Orbitrap | Orbitrap | Orbitrap | Orbitrap | Orbitrap | Orbitrap | Orbitrap | Orbitrap | Orbitrap | Orbitrap | Orbitrap
| Resolution for MS2 | 30000 | 15000 | 30000 | 30000 | 30000 | 30000 | 30000 | 30000 | 30000 | 30000 | 30000 | 30000 | 30000 | 30000 | 30000
| Fixed first mass for MS2 | 200 | 110 | 200 | 200-2000 | Auto | 200-2000 | 200-2000 | 200-2000 | 120 | 120 | 120 | 120 | 120 | 120 | 120
| AGC target for MS2 | 3000000 | custom | custom | 500000 | 50000 | 500000 | 500000 | custom | custom | custom | custom | custom | custom | custom | custom
| Normalized AGC target for MS2(%) |  | 500% | 200% | 1000% | 100% | 100% | 1000% | 1000% | 200% | 200% | 200% | 200% | 200% | 200% | 200%
| Maximum injection time for MS2(ms) | Auto | 50ms | 50ms | 54ms | 54ms | 54ms | 54ms | Auto | 50ms | 50ms | 50ms | 50ms | 50ms | 50ms | 50ms
| Scan windows number | 60(20+20+20) | 72（-45 36window, -65 36window） | 160（-45 80window, -65 80window） | 24 | 31 | 36 | 75(1+25+1+25+1+25) | 40 | 80(26+26+28) | 80(26+26+28) | 80(26+26+28) | 80(26+26+28) | 80(26+26+28) | 80(26+26+28) | 80(26+26+28)
| DIA window | m/z¦Isolation Window¦ | m/z¦Isolation Window¦CV | m/z¦Isolation Window¦CV | m/z¦Isolation Window¦ | m/z¦Isolation Window¦ | m/z¦Isolation Window¦ | m/z¦Isolation Window¦ | m/z¦Isolation Window¦ | m/z¦Isolation Window¦ | m/z¦Isolation Window¦ | m/z¦Isolation Window¦ | m/z¦Isolation Window¦ | m/z¦Isolation Window¦ | m/z¦Isolation Window¦ | m/z¦Isolation Window¦
|  | Full MS [348-1202] | Full MS [400-1300] -45¦ | Full MS [350-1200] -45¦ | 411¦ 23¦ | 362¦ 25¦ | 370¦ 40¦ | Full MS [375-1500] | 410¦ 20¦ | Full MS [350-1200] | Full MS [350-1200] | Full MS [350-1200] | Full MS [350-1200] | Full MS [350-1200] | Full MS [350-1200] | Full MS [350-1200]
|  | 360¦ 21¦ | 412.5¦ 26¦ -45¦ | 361.5¦ 23¦ -45¦ | 431¦ 19¦ | 384¦ 17¦ | 398¦ 16¦ | 403.35¦ 6.7¦ | 430¦ 20¦ | 358.5¦ 17¦ | 358.5¦ 17¦ | 358.5¦ 17¦ | 358.5¦ 17¦ | 358.5¦ 17¦ | 358.5¦ 17¦ | 358.5¦ 17¦
|  | 380¦ 21¦ | 437.5¦ 26¦ -45¦ | 383.5¦ 23¦ -45¦ | 449¦ 18¦ | 401¦ 16¦ | 413¦ 14¦ | 410.05¦ 6.7¦ | 450¦ 20¦ | 373.5¦ 13¦ | 373.5¦ 13¦ | 373.5¦ 13¦ | 373.5¦ 13¦ | 373.5¦ 13¦ | 373.5¦ 13¦ | 373.5¦ 13¦
|  | 400¦ 21¦ | 462.5¦ 26¦ -45¦ | 398¦ 8¦ -45¦ | 466¦ 19¦ | 417¦ 13¦ | 426.5¦ 13¦ | 416.75¦ 6.7¦ | 470¦ 20¦ | 384.5¦ 9¦ | 384.5¦ 9¦ | 384.5¦ 9¦ | 384.5¦ 9¦ | 384.5¦ 9¦ | 384.5¦ 9¦ | 384.5¦ 9¦
|  | 415¦ 11¦ | 487.5¦ 26¦ -45¦ | 405¦ 8¦ -45¦ | 482¦ 15¦ | 431¦ 13¦ | 439¦ 12¦ | 423.45¦ 6.7¦ | 490¦ 20¦ | 395.0¦ 13¦ | 395.0¦ 13¦ | 395.0¦ 13¦ | 395.0¦ 13¦ | 395.0¦ 13¦ | 395.0¦ 13¦ | 395.0¦ 13¦
|  | 425¦ 11¦ | 512.5¦ 26¦ -45¦ | 412¦ 8¦ -45¦ | 496¦ 16¦ | 444¦ 13¦ | 450.5¦ 11¦ | 430.15¦ 6.7¦ | 510¦ 20¦ | 404.5¦ 7¦ | 404.5¦ 7¦ | 404.5¦ 7¦ | 404.5¦ 7¦ | 404.5¦ 7¦ | 404.5¦ 7¦ | 404.5¦ 7¦
|  | 435¦ 11¦ | 537.5¦ 26¦ -45¦ | 419¦ 8¦ -45¦ | 511¦ 15¦ | 457¦ 12¦ | 461.5¦ 11¦ | 436.85¦ 6.7¦ | 530¦ 20¦ | 411.5¦ 7¦ | 411.5¦ 7¦ | 411.5¦ 7¦ | 411.5¦ 7¦ | 411.5¦ 7¦ | 411.5¦ 7¦ | 411.5¦ 7¦
|  | 445¦ 11¦ | 562.5¦ 26¦ -45¦ | 426¦ 8¦ -45¦ | 526¦ 16¦ | 469¦ 12¦ | 472¦ 10¦ | 443.55¦ 6.7¦ | 550¦ 20¦ | 419.0¦ 8¦ | 419.0¦ 8¦ | 419.0¦ 8¦ | 419.0¦ 8¦ | 419.0¦ 8¦ | 419.0¦ 8¦ | 419.0¦ 8¦
|  | 455¦ 11¦ | 587.5¦ 26¦ -45¦ | 433¦ 8¦ -45¦ | 542¦ 18¦ | 481¦ 12¦ | 482¦ 10¦ | 450.25¦ 6.7¦ | 570¦ 20¦ | 426.0¦ 6¦ | 426.0¦ 6¦ | 426.0¦ 6¦ | 426.0¦ 6¦ | 426.0¦ 6¦ | 426.0¦ 6¦ | 426.0¦ 6¦
|  | 465¦ 11¦ | 612.5¦ 26¦ -45¦ | 440¦ 8¦ -45¦ | 558¦ 17¦ | 492¦ 10¦ | 502¦ 10¦ | 456.95¦ 6.7¦ | 590¦ 20¦ | 432.0¦ 6¦ | 432.0¦ 6¦ | 432.0¦ 6¦ | 432.0¦ 6¦ | 432.0¦ 6¦ | 432.0¦ 6¦ | 432.0¦ 6¦
|  | 475¦ 11¦ | 637.5¦ 26¦ -45¦ | 447¦ 8¦ -45¦ | 574¦ 16¦ | 503¦ 12¦ | 512¦ 10¦ | 463.65¦ 6.7¦ | 610¦ 20¦ | 437.5¦ 5¦ | 437.5¦ 5¦ | 437.5¦ 5¦ | 437.5¦ 5¦ | 437.5¦ 5¦ | 437.5¦ 5¦ | 437.5¦ 5¦
|  | 485¦ 11¦ | 662.5¦ 26¦ -45¦ | 454¦ 8¦ -45¦ | 590¦ 19¦ | 514¦ 12¦ | 522¦ 10¦ | 470.35¦ 6.7¦ | 630¦ 20¦ | 442.5¦ 5¦ | 442.5¦ 5¦ | 442.5¦ 5¦ | 442.5¦ 5¦ | 442.5¦ 5¦ | 442.5¦ 5¦ | 442.5¦ 5¦
|  | 495¦ 11¦ | 687.5¦ 26¦ -45¦ | 461¦ 8¦ -45¦ | 608¦ 18¦ | 526¦ 12¦ | 532¦ 10¦ | 477.05¦ 6.7¦ | 650¦ 20¦ | 447.5¦ 5¦ | 447.5¦ 5¦ | 447.5¦ 5¦ | 447.5¦ 5¦ | 447.5¦ 5¦ | 447.5¦ 5¦ | 447.5¦ 5¦
|  | 505¦ 11¦ | 712.5¦ 26¦ -45¦ | 468¦ 8¦ -45¦ | 626¦ 21¦ | 538¦ 12¦ | 542¦ 10¦ | 483.75¦ 6.7¦ | 670¦ 20¦ | 452.5¦ 5¦ | 452.5¦ 5¦ | 452.5¦ 5¦ | 452.5¦ 5¦ | 452.5¦ 5¦ | 452.5¦ 5¦ | 452.5¦ 5¦
|  | 515¦ 11¦ | 737.5¦ 26¦ -45¦ | 475¦ 8¦ -45¦ | 648¦ 24¦ | 550¦ 12¦ | 550¦ 12¦ | 490.45¦ 6.7¦ | 690¦ 20¦ | 458.0¦ 6¦ | 458.0¦ 6¦ | 458.0¦ 6¦ | 458.0¦ 6¦ | 458.0¦ 6¦ | 458.0¦ 6¦ | 458.0¦ 6¦
|  | 525¦ 11¦ | 762.5¦ 26¦ -45¦ | 482¦ 8¦ -45¦ | 669¦ 21¦ | 563¦ 15¦ | 552¦ 10¦ | 497.15¦ 6.7¦ | 710¦ 20¦ | 463.5¦ 5¦ | 463.5¦ 5¦ | 463.5¦ 5¦ | 463.5¦ 5¦ | 463.5¦ 5¦ | 463.5¦ 5¦ | 463.5¦ 5¦
|  | 535¦ 11¦ | 787.5¦ 26¦ -45¦ | 489¦ 8¦ -45¦ | 693¦ 29¦ | 578¦ 14¦ | 562.5¦ 11¦ | 503.85¦ 6.7¦ | 730¦ 20¦ | 469.0¦ 6¦ | 469.0¦ 6¦ | 469.0¦ 6¦ | 469.0¦ 6¦ | 469.0¦ 6¦ | 469.0¦ 6¦ | 469.0¦ 6¦
|  | 545¦ 11¦ | 812.5¦ 26¦ -45¦ | 496¦ 8¦ -45¦ | 721¦ 28¦ | 592¦ 15¦ | 573.5¦ 11¦ | 510.55¦ 6.7¦ | 750¦ 20¦ | 474.0¦ 4¦ | 474.0¦ 4¦ | 474.0¦ 4¦ | 474.0¦ 4¦ | 474.0¦ 4¦ | 474.0¦ 4¦ | 474.0¦ 4¦
|  | 555¦ 11¦ | 837.5¦ 26¦ -45¦ | 503¦ 8¦ -45¦ | 752¦ 37¦ | 606¦ 14¦ | 585¦ 12¦ | 517.25¦ 6.7¦ | 770¦ 20¦ | 479.0¦ 6¦ | 479.0¦ 6¦ | 479.0¦ 6¦ | 479.0¦ 6¦ | 479.0¦ 6¦ | 479.0¦ 6¦ | 479.0¦ 6¦
|  | 565¦ 11¦ | 862.5¦ 26¦ -45¦ | 510¦ 8¦ -45¦ | 790¦ 41¦ | 622¦ 17¦ | 597.5¦ 13¦ | 523.95¦ 6.7¦ | 790¦ 20¦ | 484.0¦ 4¦ | 484.0¦ 4¦ | 484.0¦ 4¦ | 484.0¦ 4¦ | 484.0¦ 4¦ | 484.0¦ 4¦ | 484.0¦ 4¦
|  | 575¦ 11¦ | 887.5¦ 26¦ -45¦ | 517¦ 8¦ -45¦ | 833¦ 48¦ | 639¦ 16¦ | 611¦ 14¦ | 530.65¦ 6.7¦ | 810¦ 20¦ | 488.5¦ 5¦ | 488.5¦ 5¦ | 488.5¦ 5¦ | 488.5¦ 5¦ | 488.5¦ 5¦ | 488.5¦ 5¦ | 488.5¦ 5¦
|  | Full MS [348-1202] | 912.5¦ 26¦ -45¦ | Full MS [350-1200] -65¦ | 891¦ 69¦ | 656¦ 18¦ | 625.5¦ 15¦ | 537.35¦ 6.7¦ | 830¦ 20¦ | 493.0¦ 4¦ | 493.0¦ 4¦ | 493.0¦ 4¦ | 493.0¦ 4¦ | 493.0¦ 4¦ | 493.0¦ 4¦ | 493.0¦ 4¦
|  | 585¦ 11¦ | 937.5¦ 26¦ -45¦ | 361.5¦ 23¦ -65¦ | 969¦ 90¦ | 675¦ 21¦ | 641¦ 16¦ | 544.05¦ 6.7¦ | 850¦ 20¦ | 498.0¦ 6¦ | 498.0¦ 6¦ | 498.0¦ 6¦ | 498.0¦ 6¦ | 498.0¦ 6¦ | 498.0¦ 6¦ | 498.0¦ 6¦
|  | 595¦ 11¦ | 962.5¦ 26¦ -45¦ | 383.5¦ 23¦ -65¦ | 1107¦ 188¦ | 698¦ 23¦ | 657.5¦ 17¦ | 550.75¦ 6.7¦ | 870¦ 20¦ | 503.0¦ 4¦ | 503.0¦ 4¦ | 503.0¦ 4¦ | 503.0¦ 4¦ | 503.0¦ 4¦ | 503.0¦ 4¦ | 503.0¦ 4¦
|  | 605¦ 11¦ | 987.5¦ 26¦ -45¦ | 398¦ 8¦ -65¦ |  | 723¦ 28¦ | 693.5¦ 19¦ | 557.45¦ 6.7¦ | 890¦ 20¦ | 507.0¦ 4¦ | 507.0¦ 4¦ | 507.0¦ 4¦ | 507.0¦ 4¦ | 507.0¦ 4¦ | 507.0¦ 4¦ | 507.0¦ 4¦
|  | 615¦ 11¦ | 1012.5¦ 26¦ -45¦ | 405¦ 8¦ -65¦ |  | 756¦ 37¦ | 713¦ 20¦ | 564.15¦ 6.7¦ | 910¦ 20¦ | 511.5¦ 5¦ | 511.5¦ 5¦ | 511.5¦ 5¦ | 511.5¦ 5¦ | 511.5¦ 5¦ | 511.5¦ 5¦ | 511.5¦ 5¦
|  | 625¦ 11¦ | 1037.5¦ 26¦ -45¦ | 412¦ 8¦ -65¦ |  | 793¦ 36¦ | 734¦ 22¦ | Full MS [375-1500] | 930¦ 20¦ | 516.0¦ 4¦ | 516.0¦ 4¦ | 516.0¦ 4¦ | 516.0¦ 4¦ | 516.0¦ 4¦ | 516.0¦ 4¦ | 516.0¦ 4¦
|  | 635¦ 11¦ | 1062.5¦ 26¦ -45¦ | 419¦ 8¦ -65¦ |  | 838¦ 53¦ | 757¦ 24¦ | 570.85¦ 6.7¦ | 950¦ 20¦ | Full MS [350-1200] | Full MS [350-1200] | Full MS [350-1200] | Full MS [350-1200] | Full MS [350-1200] | Full MS [350-1200] | Full MS [350-1200]
|  | 645¦ 11¦ | 1087.5¦ 26¦ -45¦ | 426¦ 8¦ -65¦ |  | 908¦ 88¦ | 782.5¦ 27¦ | 577.55¦ 6.7¦ | 970¦ 20¦ | 521.0¦ 6¦ | 521.0¦ 6¦ | 521.0¦ 6¦ | 521.0¦ 6¦ | 521.0¦ 6¦ | 521.0¦ 6¦ | 521.0¦ 6¦
|  | 655¦ 11¦ | 1112.5¦ 26¦ -45¦ | 433¦ 8¦ -65¦ |  | 1034¦ 164¦ | 811¦ 30¦ | 584.25¦ 6.7¦ | 990¦ 20¦ | 526.0¦ 4¦ | 526.0¦ 4¦ | 526.0¦ 4¦ | 526.0¦ 4¦ | 526.0¦ 4¦ | 526.0¦ 4¦ | 526.0¦ 4¦
|  | 665¦ 11¦ | 1137.5¦ 26¦ -45¦ | 440¦ 8¦ -65¦ |  | 1155¦ 75¦ | 843¦ 34¦ | 590.95¦ 6.7¦ | 1010¦ 20¦ | 530.5¦ 5¦ | 530.5¦ 5¦ | 530.5¦ 5¦ | 530.5¦ 5¦ | 530.5¦ 5¦ | 530.5¦ 5¦ | 530.5¦ 5¦
|  | 675¦ 11¦ | 1162.5¦ 26¦ -45¦ | 447¦ 8¦ -65¦ |  |  | 880¦ 40¦ | 597.65¦ 6.7¦ | 1030¦ 20¦ | 536.0¦ 6¦ | 536.0¦ 6¦ | 536.0¦ 6¦ | 536.0¦ 6¦ | 536.0¦ 6¦ | 536.0¦ 6¦ | 536.0¦ 6¦
|  | 685¦ 11¦ | 1187.5¦ 26¦ -45¦ | 454¦ 8¦ -65¦ |  |  | 925¦ 50¦ | 604.35¦ 6.7¦ | 1050¦ 20¦ | 541.0¦ 4¦ | 541.0¦ 4¦ | 541.0¦ 4¦ | 541.0¦ 4¦ | 541.0¦ 4¦ | 541.0¦ 4¦ | 541.0¦ 4¦
|  | 695¦ 11¦ | 1212.5¦ 26¦ -45¦ | 461¦ 8¦ -65¦ |  |  | 988¦ 76¦ | 611.05¦ 6.7¦ | 1070¦ 20¦ | 545.0¦ 4¦ | 545.0¦ 4¦ | 545.0¦ 4¦ | 545.0¦ 4¦ | 545.0¦ 4¦ | 545.0¦ 4¦ | 545.0¦ 4¦
|  | 705¦ 11¦ | 1237.5¦ 26¦ -45¦ | 468¦ 8¦ -65¦ |  |  | 1113¦ 174¦ | 617.75¦ 6.7¦ | 1090¦ 20¦ | 549.5¦ 5¦ | 549.5¦ 5¦ | 549.5¦ 5¦ | 549.5¦ 5¦ | 549.5¦ 5¦ | 549.5¦ 5¦ | 549.5¦ 5¦
|  | 715¦ 11¦ | 1262.5¦ 26¦ -45¦ | 475¦ 8¦ -65¦ |  |  |  | 624.45¦ 6.7¦ | 1110¦ 20¦ | 554.0¦ 4¦ | 554.0¦ 4¦ | 554.0¦ 4¦ | 554.0¦ 4¦ | 554.0¦ 4¦ | 554.0¦ 4¦ | 554.0¦ 4¦
|  | 725¦ 11¦ | 1287.5¦ 26¦ -45¦ | 482¦ 8¦ -65¦ |  |  |  | 631.15¦ 6.7¦ | 1130¦ 20¦ | 558.5¦ 5¦ | 558.5¦ 5¦ | 558.5¦ 5¦ | 558.5¦ 5¦ | 558.5¦ 5¦ | 558.5¦ 5¦ | 558.5¦ 5¦
|  | 735¦ 11¦ | Full MS [400-1300] -65¦ | 489¦ 8¦ -65¦ |  |  |  | 637.85¦ 6.7¦ | 1150¦ 20¦ | 564.0¦ 6¦ | 564.0¦ 6¦ | 564.0¦ 6¦ | 564.0¦ 6¦ | 564.0¦ 6¦ | 564.0¦ 6¦ | 564.0¦ 6¦
|  | 745¦ 11¦ | 412.5¦ 26¦ -65¦ | 496¦ 8¦ -65¦ |  |  |  | 644.55¦ 6.7¦ | 1170¦ 20¦ | 570.0¦ 6¦ | 570.0¦ 6¦ | 570.0¦ 6¦ | 570.0¦ 6¦ | 570.0¦ 6¦ | 570.0¦ 6¦ | 570.0¦ 6¦
|  | 755¦ 11¦ | 437.5¦ 26¦ -65¦ | 503¦ 8¦ -65¦ |  |  |  | 651.25¦ 6.7¦ | 1190¦ 20¦ | 576.0¦ 6¦ | 576.0¦ 6¦ | 576.0¦ 6¦ | 576.0¦ 6¦ | 576.0¦ 6¦ | 576.0¦ 6¦ | 576.0¦ 6¦
|  | 765¦ 11¦ | 462.5¦ 26¦ -65¦ | 510¦ 8¦ -65¦ |  |  |  | 657.95¦ 6.7¦ |  | 581.5¦ 5¦ | 581.5¦ 5¦ | 581.5¦ 5¦ | 581.5¦ 5¦ | 581.5¦ 5¦ | 581.5¦ 5¦ | 581.5¦ 5¦
|  | 775¦ 11¦ | 487.5¦ 26¦ -65¦ | 517¦ 8¦ -65¦ |  |  |  | 664.65¦ 6.7¦ |  | 587.0¦ 6¦ | 587.0¦ 6¦ | 587.0¦ 6¦ | 587.0¦ 6¦ | 587.0¦ 6¦ | 587.0¦ 6¦ | 587.0¦ 6¦
|  | Full MS [348-1202] | 512.5¦ 26¦ -65¦ | Full MS [350-1200] -45¦ |  |  |  | 671.35¦ 6.7¦ |  | 593.0¦ 6¦ | 593.0¦ 6¦ | 593.0¦ 6¦ | 593.0¦ 6¦ | 593.0¦ 6¦ | 593.0¦ 6¦ | 593.0¦ 6¦
|  | 785¦ 11¦ | 537.5¦ 26¦ -65¦ | 524¦ 8¦ -45¦ |  |  |  | 678.05¦ 6.7¦ |  | 598.5¦ 5¦ | 598.5¦ 5¦ | 598.5¦ 5¦ | 598.5¦ 5¦ | 598.5¦ 5¦ | 598.5¦ 5¦ | 598.5¦ 5¦
|  | 795¦ 11¦ | 562.5¦ 26¦ -65¦ | 531¦ 8¦ -45¦ |  |  |  | 684.75¦ 6.7¦ |  | 603.5¦ 5¦ | 603.5¦ 5¦ | 603.5¦ 5¦ | 603.5¦ 5¦ | 603.5¦ 5¦ | 603.5¦ 5¦ | 603.5¦ 5¦
|  | 805¦ 11¦ | 587.5¦ 26¦ -65¦ | 538¦ 8¦ -45¦ |  |  |  | 691.45¦ 6.7¦ |  | 609.0¦ 6¦ | 609.0¦ 6¦ | 609.0¦ 6¦ | 609.0¦ 6¦ | 609.0¦ 6¦ | 609.0¦ 6¦ | 609.0¦ 6¦
|  | 815¦ 11¦ | 612.5¦ 26¦ -65¦ | 545¦ 8¦ -45¦ |  |  |  | 698.15¦ 6.7¦ |  | 614.5¦ 5¦ | 614.5¦ 5¦ | 614.5¦ 5¦ | 614.5¦ 5¦ | 614.5¦ 5¦ | 614.5¦ 5¦ | 614.5¦ 5¦
|  | 825¦ 11¦ | 637.5¦ 26¦ -65¦ | 552¦ 8¦ -45¦ |  |  |  | 704.85¦ 6.7¦ |  | 619.5¦ 5¦ | 619.5¦ 5¦ | 619.5¦ 5¦ | 619.5¦ 5¦ | 619.5¦ 5¦ | 619.5¦ 5¦ | 619.5¦ 5¦
|  | 840¦ 21¦ | 662.5¦ 26¦ -65¦ | 559¦ 8¦ -45¦ |  |  |  | 711.55¦ 6.7¦ |  | 625.5¦ 7¦ | 625.5¦ 7¦ | 625.5¦ 7¦ | 625.5¦ 7¦ | 625.5¦ 7¦ | 625.5¦ 7¦ | 625.5¦ 7¦
|  | 860¦ 21¦ | 687.5¦ 26¦ -65¦ | 566¦ 8¦ -45¦ |  |  |  | 718.25¦ 6.7¦ |  | 631.5¦ 5¦ | 631.5¦ 5¦ | 631.5¦ 5¦ | 631.5¦ 5¦ | 631.5¦ 5¦ | 631.5¦ 5¦ | 631.5¦ 5¦
|  | 880¦ 21¦ | 712.5¦ 26¦ -65¦ | 573¦ 8¦ -45¦ |  |  |  | 724.95¦ 6.7¦ |  | 636.5¦ 5¦ | 636.5¦ 5¦ | 636.5¦ 5¦ | 636.5¦ 5¦ | 636.5¦ 5¦ | 636.5¦ 5¦ | 636.5¦ 5¦
|  | 900¦ 21¦ | 737.5¦ 26¦ -65¦ | 580¦ 8¦ -45¦ |  |  |  | 731.65¦ 6.7¦ |  | 641.5¦ 5¦ | 641.5¦ 5¦ | 641.5¦ 5¦ | 641.5¦ 5¦ | 641.5¦ 5¦ | 641.5¦ 5¦ | 641.5¦ 5¦
|  | 920¦ 21¦ | 762.5¦ 26¦ -65¦ | 587¦ 8¦ -45¦ |  |  |  | Full MS [375-1500] |  | 647.0¦ 6¦ | 647.0¦ 6¦ | 647.0¦ 6¦ | 647.0¦ 6¦ | 647.0¦ 6¦ | 647.0¦ 6¦ | 647.0¦ 6¦
|  | 940¦ 21¦ | 787.5¦ 26¦ -65¦ | 594¦ 8¦ -45¦ |  |  |  | 738.35¦ 6.7¦ |  | 653.5¦ 7¦ | 653.5¦ 7¦ | 653.5¦ 7¦ | 653.5¦ 7¦ | 653.5¦ 7¦ | 653.5¦ 7¦ | 653.5¦ 7¦
|  | 960¦ 21¦ | 812.5¦ 26¦ -65¦ | 601¦ 8¦ -45¦ |  |  |  | 745.05¦ 6.7¦ |  | Full MS [350-1200] | Full MS [350-1200] | Full MS [350-1200] | Full MS [350-1200] | Full MS [350-1200] | Full MS [350-1200] | Full MS [350-1200]
|  | 980¦ 21¦ | 837.5¦ 26¦ -65¦ | 608¦ 8¦ -45¦ |  |  |  | 751.75¦ 6.7¦ |  | 659.5¦ 5¦ | 659.5¦ 5¦ | 659.5¦ 5¦ | 659.5¦ 5¦ | 659.5¦ 5¦ | 659.5¦ 5¦ | 659.5¦ 5¦
|  | 1005¦ 31¦ | 862.5¦ 26¦ -65¦ | 615¦ 8¦ -45¦ |  |  |  | 758.45¦ 6.7¦ |  | 665.5¦ 7¦ | 665.5¦ 7¦ | 665.5¦ 7¦ | 665.5¦ 7¦ | 665.5¦ 7¦ | 665.5¦ 7¦ | 665.5¦ 7¦
|  | 1035¦ 31¦ | 887.5¦ 26¦ -65¦ | 622¦ 8¦ -45¦ |  |  |  | 765.15¦ 6.7¦ |  | 672.0¦ 6¦ | 672.0¦ 6¦ | 672.0¦ 6¦ | 672.0¦ 6¦ | 672.0¦ 6¦ | 672.0¦ 6¦ | 672.0¦ 6¦
|  | 1065¦ 31¦ | 912.5¦ 26¦ -65¦ | 629¦ 8¦ -45¦ |  |  |  | 771.85¦ 6.7¦ |  | 678.5¦ 7¦ | 678.5¦ 7¦ | 678.5¦ 7¦ | 678.5¦ 7¦ | 678.5¦ 7¦ | 678.5¦ 7¦ | 678.5¦ 7¦
|  | 1095¦ 31¦ | 937.5¦ 26¦ -65¦ | 636¦ 8¦ -45¦ |  |  |  | 778.55¦ 6.7¦ |  | 686.5¦ 9¦ | 686.5¦ 9¦ | 686.5¦ 9¦ | 686.5¦ 9¦ | 686.5¦ 9¦ | 686.5¦ 9¦ | 686.5¦ 9¦
|  | 1125¦ 31¦ | 962.5¦ 26¦ -65¦ | 643¦ 8¦ -45¦ |  |  |  | 785.25¦ 6.7¦ |  | 694.5¦ 7¦ | 694.5¦ 7¦ | 694.5¦ 7¦ | 694.5¦ 7¦ | 694.5¦ 7¦ | 694.5¦ 7¦ | 694.5¦ 7¦
|  | 1155¦ 31¦ | 987.5¦ 26¦ -65¦ | 650¦ 8¦ -45¦ |  |  |  | 791.95¦ 6.7¦ |  | 702.5¦ 9¦ | 702.5¦ 9¦ | 702.5¦ 9¦ | 702.5¦ 9¦ | 702.5¦ 9¦ | 702.5¦ 9¦ | 702.5¦ 9¦
|  | 1185¦ 31¦ | 1012.5¦ 26¦ -65¦ | 657¦ 8¦ -45¦ |  |  |  | 798.65¦ 6.7¦ |  | 711.0¦ 8¦ | 711.0¦ 8¦ | 711.0¦ 8¦ | 711.0¦ 8¦ | 711.0¦ 8¦ | 711.0¦ 8¦ | 711.0¦ 8¦
|  |  | 1037.5¦ 26¦ -65¦ | Full MS [350-1200] -65¦ |  |  |  | 805.35¦ 6.7¦ |  | 719.0¦ 8¦ | 719.0¦ 8¦ | 719.0¦ 8¦ | 719.0¦ 8¦ | 719.0¦ 8¦ | 719.0¦ 8¦ | 719.0¦ 8¦
|  |  | 1062.5¦ 26¦ -65¦ | 524¦ 8¦ -65¦ |  |  |  | 812.05¦ 6.7¦ |  | 728.0¦ 10¦ | 728.0¦ 10¦ | 728.0¦ 10¦ | 728.0¦ 10¦ | 728.0¦ 10¦ | 728.0¦ 10¦ | 728.0¦ 10¦
|  |  | 1087.5¦ 26¦ -65¦ | 531¦ 8¦ -65¦ |  |  |  | 818.75¦ 6.7¦ |  | 738.0¦ 10¦ | 738.0¦ 10¦ | 738.0¦ 10¦ | 738.0¦ 10¦ | 738.0¦ 10¦ | 738.0¦ 10¦ | 738.0¦ 10¦
|  |  | 1112.5¦ 26¦ -65¦ | 538¦ 8¦ -65¦ |  |  |  | 825.45¦ 6.7¦ |  | 748.0¦ 10¦ | 748.0¦ 10¦ | 748.0¦ 10¦ | 748.0¦ 10¦ | 748.0¦ 10¦ | 748.0¦ 10¦ | 748.0¦ 10¦
|  |  | 1137.5¦ 26¦ -65¦ | 545¦ 8¦ -65¦ |  |  |  | 832.15¦ 6.7¦ |  | 758.5¦ 11¦ | 758.5¦ 11¦ | 758.5¦ 11¦ | 758.5¦ 11¦ | 758.5¦ 11¦ | 758.5¦ 11¦ | 758.5¦ 11¦
|  |  | 1162.5¦ 26¦ -65¦ | 552¦ 8¦ -65¦ |  |  |  | 838.85¦ 6.7¦ |  | 768.5¦ 9¦ | 768.5¦ 9¦ | 768.5¦ 9¦ | 768.5¦ 9¦ | 768.5¦ 9¦ | 768.5¦ 9¦ | 768.5¦ 9¦
|  |  | 1187.5¦ 26¦ -65¦ | 559¦ 8¦ -65¦ |  |  |  | 845.55¦ 6.7¦ |  | 779.0¦ 12¦ | 779.0¦ 12¦ | 779.0¦ 12¦ | 779.0¦ 12¦ | 779.0¦ 12¦ | 779.0¦ 12¦ | 779.0¦ 12¦
|  |  | 1212.5¦ 26¦ -65¦ | 566¦ 8¦ -65¦ |  |  |  | 852.25¦ 6.7¦ |  | 791.5¦ 13¦ | 791.5¦ 13¦ | 791.5¦ 13¦ | 791.5¦ 13¦ | 791.5¦ 13¦ | 791.5¦ 13¦ | 791.5¦ 13¦
|  |  | 1237.5¦ 26¦ -65¦ | 573¦ 8¦ -65¦ |  |  |  | 858.95¦ 6.7¦ |  | 803.5¦ 11¦ | 803.5¦ 11¦ | 803.5¦ 11¦ | 803.5¦ 11¦ | 803.5¦ 11¦ | 803.5¦ 11¦ | 803.5¦ 11¦
|  |  | 1262.5¦ 26¦ -65¦ | 580¦ 8¦ -65¦ |  |  |  | 865.65¦ 6.7¦ |  | 815.0¦ 12¦ | 815.0¦ 12¦ | 815.0¦ 12¦ | 815.0¦ 12¦ | 815.0¦ 12¦ | 815.0¦ 12¦ | 815.0¦ 12¦
|  |  | 1287.5¦ 26¦ -65¦ | 587¦ 8¦ -65¦ |  |  |  | 872.35¦ 6.7¦ |  | 828.0¦ 14¦ | 828.0¦ 14¦ | 828.0¦ 14¦ | 828.0¦ 14¦ | 828.0¦ 14¦ | 828.0¦ 14¦ | 828.0¦ 14¦
|  |  |  | 594¦ 8¦ -65¦ |  |  |  | 879.05¦ 6.7¦ |  | 843.0¦ 16¦ | 843.0¦ 16¦ | 843.0¦ 16¦ | 843.0¦ 16¦ | 843.0¦ 16¦ | 843.0¦ 16¦ | 843.0¦ 16¦
|  |  |  | 601¦ 8¦ -65¦ |  |  |  | 885.75¦ 6.7¦ |  | 859.5¦ 17¦ | 859.5¦ 17¦ | 859.5¦ 17¦ | 859.5¦ 17¦ | 859.5¦ 17¦ | 859.5¦ 17¦ | 859.5¦ 17¦
|  |  |  | 608¦ 8¦ -65¦ |  |  |  | 892.45¦ 6.7¦ |  | 876.5¦ 17¦ | 876.5¦ 17¦ | 876.5¦ 17¦ | 876.5¦ 17¦ | 876.5¦ 17¦ | 876.5¦ 17¦ | 876.5¦ 17¦
|  |  |  | 615¦ 8¦ -65¦ |  |  |  | 899.15¦ 6.7¦ |  | 898.5¦ 27¦ | 898.5¦ 27¦ | 898.5¦ 27¦ | 898.5¦ 27¦ | 898.5¦ 27¦ | 898.5¦ 27¦ | 898.5¦ 27¦
|  |  |  | 622¦ 8¦ -65¦ | [5 empty] | 923.5¦ 23¦ | 923.5¦ 23¦ | 923.5¦ 23¦ | 923.5¦ 23¦ | 923.5¦ 23¦ | 923.5¦ 23¦ | 923.5¦ 23¦
|  |  |  | 629¦ 8¦ -65¦ | [5 empty] | 943.5¦ 17¦ | 943.5¦ 17¦ | 943.5¦ 17¦ | 943.5¦ 17¦ | 943.5¦ 17¦ | 943.5¦ 17¦ | 943.5¦ 17¦
|  |  |  | 636¦ 8¦ -65¦ | [5 empty] | 962.5¦ 21¦ | 962.5¦ 21¦ | 962.5¦ 21¦ | 962.5¦ 21¦ | 962.5¦ 21¦ | 962.5¦ 21¦ | 962.5¦ 21¦
|  |  |  | 643¦ 8¦ -65¦ | [5 empty] | 997.5¦ 49¦ | 997.5¦ 49¦ | 997.5¦ 49¦ | 997.5¦ 49¦ | 997.5¦ 49¦ | 997.5¦ 49¦ | 997.5¦ 49¦
|  |  |  | 650¦ 8¦ -65¦ | [5 empty] | 1061.0¦ 78¦ | 1061.0¦ 78¦ | 1061.0¦ 78¦ | 1061.0¦ 78¦ | 1061.0¦ 78¦ | 1061.0¦ 78¦ | 1061.0¦ 78¦
657¦ 8¦ -65¦
Full MS [350-1200] -45¦
664¦ 8¦ -45¦
671¦ 8¦ -45¦
678¦ 8¦ -45¦
685¦ 8¦ -45¦
692¦ 8¦ -45¦
699¦ 8¦ -45¦
706¦ 8¦ -45¦
713¦ 8¦ -45¦
720¦ 8¦ -45¦
727¦ 8¦ -45¦
734¦ 8¦ -45¦
741¦ 8¦ -45¦
748¦ 8¦ -45¦
755¦ 8¦ -45¦
762¦ 8¦ -45¦
769¦ 8¦ -45¦
776¦ 8¦ -45¦
783¦ 8¦ -45¦
790¦ 8¦ -45¦
797¦ 8¦ -45¦
Full MS [350-1200] -65¦
664¦ 8¦ -65¦
671¦ 8¦ -65¦
678¦ 8¦ -65¦
685¦ 8¦ -65¦
692¦ 8¦ -65¦
699¦ 8¦ -65¦
706¦ 8¦ -65¦
713¦ 8¦ -65¦
720¦ 8¦ -65¦
727¦ 8¦ -65¦
734¦ 8¦ -65¦
741¦ 8¦ -65¦
748¦ 8¦ -65¦
755¦ 8¦ -65¦
762¦ 8¦ -65¦
769¦ 8¦ -65¦
776¦ 8¦ -65¦
783¦ 8¦ -65¦
790¦ 8¦ -65¦
797¦ 8¦ -65¦
Full MS [350-1200] -45¦
807.5¦ 15¦ -45¦
821.5¦ 15¦ -45¦
835.5¦ 15¦ -45¦
849.5¦ 15¦ -45¦
863.5¦ 15¦ -45¦
877.5¦ 15¦ -45¦
891.5¦ 15¦ -45¦
905.5¦ 15¦ -45¦
919.5¦ 15¦ -45¦
933.5¦ 15¦ -45¦
947.5¦ 15¦ -45¦
961.5¦ 15¦ -45¦
975.5¦ 15¦ -45¦
989.5¦ 15¦ -45¦
1003.5¦ 15¦ -45¦
1029.5¦ 39¦ -45¦
1067.5¦ 39¦ -45¦
1105.5¦ 39¦ -45¦
1143.5¦ 39¦ -45¦
1181.5¦ 39¦ -45¦
Full MS [350-1200] -65¦
807.5¦ 15¦ -65¦
821.5¦ 15¦ -65¦
835.5¦ 15¦ -65¦
849.5¦ 15¦ -65¦
863.5¦ 15¦ -65¦
877.5¦ 15¦ -65¦
891.5¦ 15¦ -65¦
905.5¦ 15¦ -65¦
919.5¦ 15¦ -65¦
933.5¦ 15¦ -65¦
947.5¦ 15¦ -65¦
961.5¦ 15¦ -65¦
975.5¦ 15¦ -65¦
989.5¦ 15¦ -65¦
1003.5¦ 15¦ -65¦
1029.5¦ 39¦ -65¦
1067.5¦ 39¦ -65¦
1105.5¦ 39¦ -65¦
1143.5¦ 39¦ -65¦
1181.5¦ 39¦ -65¦

## Supplementary Data 2C
Supplementary Data 2C. LC-MS parameters for Bruker timsTOF data.
Lab information | Lab name | M02-tims Pro | U08-tims Pro2 | U09-tims Pro2
| LC type | nanoElute | UltiMate 3000 | nanoElute
| MS type | timsTOF Pro | timsTOF Pro 2 | timsTOF Pro 2
LC method | Loading amount | 500 ng | 500 ng | 500 ng
| Precolumn | no | no | no
| Analytical column | C18, 25 cm × 75 μm, 1.6 μm ionopticks | Monolithic Column, 50 cm × 50 μm | Monolithic Column, 50 cm × 50 μm
| Mobile phase A | 100% H2O/0.1% FA | 100% H2O/0.1% FA | 100% H2O/0.1% FA
| Mobile phase B | 80% ACN/20% H2O/0.1% FA | 100% ACN/0.1% FA | 100% ACN/0.1% FA
| Flow rate（nL/min） | 300 | 500 | 500
| Gradient | 30min | 30min | 30min
| Mobile phase | 0min, 2%B; 20min, 22%B 24min, 37%B 26min, 80%B 30min, 80%B | 0min, 4%B; 22min, 16%B 25.3min, 24%B 25.4min, 72%B 26.4min, 72%B 26.5min, 1%B 30min, 1%B | 0min, 4%B; 22min, 16%B 25.3min, 24%B 25.4min, 72%B 26.4min, 72%B 26.5min, 1%B 30min, 1%B
DDA method | Acquisition time | 30min | 30min | 30min
| Acquisition mode | DDA | DDA | DDA
| Detection type for MS | TOF | TOF | TOF
MS settings | Scan Mode | PASEF | PASEF | PASEF
| Scan range for MS | 100-1700 m/z | 100-1700 m/z | 100-1700 m/z
TIMS settings | 1/K0 start (V.s/cm2) | 0.75 | 0.75 | 0.75
| 1/K0 End (V.s/cm2) | 1.3 | 1.3 | 1.3
| Ramp time | 100ms | 100ms | 100ms
| Accu. Time | 100ms | 100ms | 100ms
| Ducy cycle | 100% | 100% | 100%
| Ramp rate | 9.43Hz | 9.43Hz | 9.43Hz
MS/MS | Number of PASEF ramps | 9 | 9 | 9
| Total Cycle time | 1.06 s | 1.06 s | 1.06 s
| Charge minimum | 2 | 2 | 2
| Charge maximum | 5 | 5 | 5
| Target intensity | 5000 | 5000 | 5000
| Intensity Threshold | 1000 | 1000 | 1000
Active exclusion | Active exclusion | TRUE | TRUE | TRUE
| Release after | 0.25min | 0.25min | 0.25min
Collision energy setting | 1/k0 (V.s/cm2) | 0.6; 1.6 | 0.6; 1.6 | 0.6; 1.6
| Collision energy | 20; 59 | 20; 59 | 20; 59
Isolation width setting | mass (m/z) | 700; 800 | 700; 800 | 700; 800
| width (m/z) | 2; 3 | 2; 3 | 2; 3
DIA method | Scan Mode | dia-PASEF | dia-PASEF | dia-PASEF
| Scan range for MS | 100-1700 m/z | 100-1700 m/z | 100-1700 m/z
TIMS settings | 1/K0 start (V.s/cm2) | 0.75 | 0.75 | 0.75
| 1/K0 End (V.s/cm2) | 1.3 | 1.3 | 1.3
| Ramp time | 100 ms | 100 ms | 100 ms
| Accu. Time | 100 ms | 100 ms | 100 ms
| Ducy cycle | 100 | 100 | 100
| Ramp rate | 9.43 Hz | 9.43 Hz | 9.43 Hz
MS/MS windows setting | MS width | Variable window | 16 Da | 16 Da
| MS overlap | 0 Da | 0 Da | 0 Da
| Mass steps per cycle | 32 | 50 | 50
| calculate from polygon | Mass steps | Mass steps | Mass steps
| mobility overlap | 0 1/k0 | 0 1/k0 | 0 1/k0
| 1/k0 (V.s/cm2) | 0.6; 1.6 | 0.6; 1.6 | 0.6; 1.6
| Collision energy | 20; 59 | 20; 59 | 20; 59
| Mobility range (1/k0) | 0.75-1.3 | 0.75-1.29 | 0.75-1.29
| Mass range | 398-1230 | 369-1169 | 369-1169
| Mean cycle time estimate | 0.95 s | 1.48 s | 1.48 s
| detailed Scan window | Cycle Id¦ Start IM [1/K0]¦ End IM [1/K0]¦ Start Mass [m/z]¦ End Mass [m/z]¦ | Cycle Id¦ Start IM [1/K0]¦ End IM [1/K0]¦ Start Mass [m/z]¦ End Mass [m/z]¦ | Cycle Id¦ Start IM [1/K0]¦ End IM [1/K0]¦ Start Mass [m/z]¦ End Mass [m/z]¦
|  | 1¦ 0.97¦ 1.15¦ 713.33¦ 730.32¦ | 1¦ 1.2289¦ 1.2849¦ 1105¦ 1121¦ | 1¦ 1.2289¦ 1.2849¦ 1105¦ 1121¦
|  | 1¦ 0.9¦ 0.97¦ 600.83¦ 613.84¦ | 1¦ 1.0851¦ 1.2199¦ 897¦ 913¦ | 1¦ 1.0851¦ 1.2199¦ 897¦ 913¦
|  | 1¦ 0.83¦ 0.9¦ 511.76¦ 522.28¦ | 1¦ 0.9523¦ 1.0766¦ 705¦ 721¦ | 1¦ 0.9523¦ 1.0766¦ 705¦ 721¦
|  | 1¦ 0.75¦ 0.83¦ 398.21¦ 422.22¦ | 1¦ 0.8305¦ 0.9452¦ 529¦ 545¦ | 1¦ 0.8305¦ 0.9452¦ 529¦ 545¦
|  | 2¦ 0.98¦ 1.15¦ 730.32¦ 750.84¦ | 1¦ 0.7501¦ 0.8258¦ 369¦ 385¦ | 1¦ 0.7501¦ 0.8258¦ 369¦ 385¦
|  | 2¦ 0.9¦ 0.98¦ 613.84¦ 627.31¦ | 2¦ 1.2400¦ 1.2847¦ 1121¦ 1137¦ | 2¦ 1.2400¦ 1.2847¦ 1121¦ 1137¦
|  | 2¦ 0.84¦ 0.9¦ 522.28¦ 532.3¦ | 2¦ 1.0961¦ 1.2318¦ 913¦ 929¦ | 2¦ 1.0961¦ 1.2318¦ 913¦ 929¦
|  | 2¦ 0.75¦ 0.84¦ 422.22¦ 437.24¦ | 2¦ 0.9633¦ 1.0885¦ 721¦ 737¦ | 2¦ 0.9633¦ 1.0885¦ 721¦ 737¦
|  | 3¦ 0.99¦ 1.15¦ 750.84¦ 776.84¦ | 2¦ 0.8416¦ 0.9571¦ 545¦ 561¦ | 2¦ 0.8416¦ 0.9571¦ 545¦ 561¦
|  | 3¦ 0.91¦ 0.99¦ 627.31¦ 638.82¦ | 2¦ 0.7500¦ 0.8377¦ 385¦ 401¦ | 2¦ 0.7500¦ 0.8377¦ 385¦ 401¦
|  | 3¦ 0.84¦ 0.91¦ 532.3¦ 544.28¦ | 3¦ 1.2511¦ 1.2846¦ 1137¦ 1153¦ | 3¦ 1.2511¦ 1.2846¦ 1137¦ 1153¦
|  | 3¦ 0.75¦ 0.84¦ 437.24¦ 451.25¦ | 3¦ 1.1072¦ 1.2437¦ 929¦ 945¦ | 3¦ 1.1072¦ 1.2437¦ 929¦ 945¦
|  | 4¦ 1.01¦ 1.2¦ 776.84¦ 801.95¦ | 3¦ 0.9744¦ 1.1004¦ 737¦ 753¦ | 3¦ 0.9744¦ 1.1004¦ 737¦ 753¦
|  | 4¦ 0.92¦ 1.01¦ 638.82¦ 652.34¦ | 3¦ 0.8527¦ 0.9691¦ 561¦ 577¦ | 3¦ 0.8527¦ 0.9691¦ 561¦ 577¦
|  | 4¦ 0.85¦ 0.92¦ 544.28¦ 555.23¦ | 3¦ 0.7500¦ 0.8497¦ 401¦ 417¦ | 3¦ 0.7500¦ 0.8497¦ 401¦ 417¦
|  | 4¦ 0.75¦ 0.85¦ 451.25¦ 464.25¦ | 4¦ 1.2621¦ 1.2844¦ 1153¦ 1169¦ | 4¦ 1.2621¦ 1.2844¦ 1153¦ 1169¦
|  | 5¦ 1.01¦ 1.2¦ 801.95¦ 833.81¦ | 4¦ 1.1183¦ 1.2557¦ 945¦ 961¦ | 4¦ 1.1183¦ 1.2557¦ 945¦ 961¦
|  | 5¦ 0.93¦ 1.01¦ 652.34¦ 665.33¦ | 4¦ 0.9855¦ 1.1124¦ 753¦ 769¦ | 4¦ 0.9855¦ 1.1124¦ 753¦ 769¦
|  | 5¦ 0.86¦ 0.93¦ 555.23¦ 565.33¦ | 4¦ 0.8637¦ 0.981¦ 577¦ 593¦ | 4¦ 0.8637¦ 0.981¦ 577¦ 593¦
|  | 5¦ 0.75¦ 0.86¦ 464.25¦ 477.2¦ | 4¦ 0.7531¦ 0.8616¦ 417¦ 433¦ | 4¦ 0.7531¦ 0.8616¦ 417¦ 433¦
|  | 6¦ 1.02¦ 1.2¦ 833.81¦ 880.94¦ | 5¦ 1.1293¦ 1.2676¦ 961¦ 977¦ | 5¦ 1.1293¦ 1.2676¦ 961¦ 977¦
|  | 6¦ 0.94¦ 1.02¦ 665.33¦ 678.33¦ | 5¦ 0.9965¦ 1.1243¦ 769¦ 785¦ | 5¦ 0.9965¦ 1.1243¦ 769¦ 785¦
|  | 6¦ 0.86¦ 0.94¦ 565.33¦ 577.29¦ | 5¦ 0.8748¦ 0.993¦ 593¦ 609¦ | 5¦ 0.8748¦ 0.993¦ 593¦ 609¦
|  | 6¦ 0.75¦ 0.86¦ 477.2¦ 488.75¦ | 5¦ 0.7641¦ 0.8736¦ 433¦ 449¦ | 5¦ 0.7641¦ 0.8736¦ 433¦ 449¦
|  | 7¦ 1.04¦ 1.3¦ 880.94¦ 960.96¦ | 6¦ 1.1404¦ 1.2796¦ 977¦ 993¦ | 6¦ 1.1404¦ 1.2796¦ 977¦ 993¦
|  | 7¦ 0.95¦ 1.04¦ 678.33¦ 695.38¦ | 6¦ 1.0076¦ 1.1363¦ 785¦ 801¦ | 6¦ 1.0076¦ 1.1363¦ 785¦ 801¦
|  | 7¦ 0.87¦ 0.95¦ 577.29¦ 589.76¦ | 6¦ 0.8859¦ 1.0049¦ 609¦ 625¦ | 6¦ 0.8859¦ 1.0049¦ 609¦ 625¦
|  | 7¦ 0.75¦ 0.87¦ 488.75¦ 500.75¦ | 6¦ 0.7752¦ 0.8855¦ 449¦ 465¦ | 6¦ 0.7752¦ 0.8855¦ 449¦ 465¦
|  | 8¦ 1.06¦ 1.3¦ 960.96¦ 1230.67¦ | 7¦ 1.1625¦ 1.2859¦ 1009¦ 1025¦ | 7¦ 1.1625¦ 1.2859¦ 1009¦ 1025¦
|  | 8¦ 0.97¦ 1.06¦ 695.38¦ 713.33¦ | 7¦ 1.0297¦ 1.1601¦ 817¦ 833¦ | 7¦ 1.0297¦ 1.1601¦ 817¦ 833¦
|  | 8¦ 0.88¦ 0.97¦ 589.76¦ 600.83¦ | 7¦ 0.908¦ 1.0288¦ 641¦ 657¦ | 7¦ 0.908¦ 1.0288¦ 641¦ 657¦
|  | 8¦ 0.75¦ 0.88¦ 500.75¦ 511.76¦ | 7¦ 0.7863¦ 0.8974¦ 465¦ 481¦ | 7¦ 0.7863¦ 0.8974¦ 465¦ 481¦
|  |  | 8¦ 1.1736¦ 1.2857¦ 1025¦ 1041¦ | 8¦ 1.1736¦ 1.2857¦ 1025¦ 1041¦
|  |  | 8¦ 1.0408¦ 1.1721¦ 833¦ 849¦ | 8¦ 1.0408¦ 1.1721¦ 833¦ 849¦
|  |  | 8¦ 0.9191¦ 1.0407¦ 657¦ 673¦ | 8¦ 0.9191¦ 1.0407¦ 657¦ 673¦
|  |  | 8¦ 0.7973¦ 0.9094¦ 481¦ 497¦ | 8¦ 0.7973¦ 0.9094¦ 481¦ 497¦
|  |  | 9¦ 1.2068¦ 1.2852¦ 1073¦ 1089¦ | 9¦ 1.2068¦ 1.2852¦ 1073¦ 1089¦
|  |  | 9¦ 1.0629¦ 1.196¦ 865¦ 881¦ | 9¦ 1.0629¦ 1.196¦ 865¦ 881¦
|  |  | 9¦ 0.9301¦ 1.0527¦ 673¦ 689¦ | 9¦ 0.9301¦ 1.0527¦ 673¦ 689¦
|  |  | 9¦ 0.8084¦ 0.9213¦ 497¦ 513¦ | 9¦ 0.8084¦ 0.9213¦ 497¦ 513¦
|  |  | 10¦ 1.2179¦ 1.2851¦ 1089¦ 1105¦ | 10¦ 1.2179¦ 1.2851¦ 1089¦ 1105¦
|  |  | 10¦ 1.074¦ 1.2079¦ 881¦ 897¦ | 10¦ 1.074¦ 1.2079¦ 881¦ 897¦
|  |  | 10¦ 0.9412¦ 1.0646¦ 689¦ 705¦ | 10¦ 0.9412¦ 1.0646¦ 689¦ 705¦
|  |  | 10¦ 0.8195¦ 0.9333¦ 513¦ 529¦ | 10¦ 0.8195¦ 0.9333¦ 513¦ 529¦
|  |  | 11¦ 1.1515¦ 1.286¦ 993¦ 1009¦ | 11¦ 1.1515¦ 1.286¦ 993¦ 1009¦
|  |  | 11¦ 1.0187¦ 1.1482¦ 801¦ 817¦ | 11¦ 1.0187¦ 1.1482¦ 801¦ 817¦
|  |  | 11¦ 0.8969¦ 1.0169¦ 625¦ 641¦ | 11¦ 0.8969¦ 1.0169¦ 625¦ 641¦
|  |  | 12¦ 1.1847¦ 1.2856¦ 1041¦ 1057¦ | 12¦ 1.1847¦ 1.2856¦ 1041¦ 1057¦
|  |  | 12¦ 1.0519¦ 1.184¦ 849¦ 865¦ | 12¦ 1.0519¦ 1.184¦ 849¦ 865¦
|  |  | 13¦ 1.1957¦ 1.2854¦ 1057¦ 1073¦ | 13¦ 1.1957¦ 1.2854¦ 1057¦ 1073¦

## Supplementary Data 2D
Supplementary Data 2D. LC-MS parameters for SCIEX zenoTOF data.
Lab information | Lab name | M10-Zeno7600 | U10-Zeno7600
| LC type | Waters M Class | Waters M Class
| MS type | ZenoTOF 7600 | ZenoTOF 7600
| Loading amount | 2.5 μL/2.5μg | 2 μL/2 μg
LC method | Precolumn | no | no
| Analytical column | C18, 15 cm × 0.3 mm, 2.1 μm, Phenominex | Monolithic Column, 50 cm × 50 μm
| Mobile phase A | 100% H2O/0.1% FA | 100% H2O/0.1% FA
| Mobile phase B | 100% ACN/0.1% FA | 100% ACN/0.1% FA
| Flow rate（nL/min） | 5000 | 500
| Gradient | 30 min | 30 min
| Mobile phase | 0min, 5%B; 22min, 20%B 25.3min, 30%B 25.4min, 80%B 26.4min, 80%B 26.5min, 2%B 30min, 2%B | 0min, 4%B; 22min, 16%B 25.5min, 24%B 25.6min, 72%B 26.6min, 72%B 26.7min, 2%B 30min, 2%B
DDA method | Acquisition time | 30 min | 30 min
| Acquisition mode | DDA | DDA
| Detection type for MS | TOF | TOF
| Total scan time (seconds) | 1.369 | 1.368
| Estimated cycles | 1271 | 1316
MS settings | Scan type | TOFMS | TOFMS
| Polarity | Positive | Positive
| Spray voltage (V) | 5500 | 3000
| Scan range for MS | 350-1500 | 350-1200
| Accumulation time (s) | 0.1 | 0.1
| Declustering potential (V) | 80 | 80
| Collision energy (V) | 10 | 10
| Time bins to sum | 8 | 8
IDA Criteria | Maximum candidates ion | 50 | 50
| Intensity threshold exceeds (counts/s) | 100 | 100
| Exclude former candidate ions | TRUE | TRUE
| For (sec) | 5 | 12
| After (occurrences) | 1 | 1
| Dynamic CE for MS/MS | TRUE | TRUE
| Charge state | 2-5 | 2-5
| Isotope to select | Monoisotopic | Monoisotopic
IDA Dependent | Scan type | TOFMSMS | TOFMSMS
| Polarity | Positive | Positive
| Fragmentation mode | CID | CID
| Scan range for MS/MS | 140-1800 | 110-1500
| Accumulation time (s) | 0.02 | 0.02
| Zeno threshold (cps) | 100000 | 100000
| Declustering potential (V) | 80 | 80
| Tims bins to sum | 8 | 8
DIA method | Acquisition time | 30 min | 30 min
| Acquisition mode | SWATH | SWATH
| Total scan time (seconds) | 1.545 | 2.095
| Estimated cycles | 1165 | 859
MS settings | Scan type | TOFMS | TOFMS
| Polarity | positive | positive
| Spray voltage (V) | 5500 | 3000
| Scan range for MS | 350-1200 | 350-1200
| Accumulation time (s) | 0.1 | 0.25
| Declustering potential (V) | 80 | 80
| Collision energy (V) | 10 | 10
| Time bins to sum | 8 | 8
TOFMSMS | Fragmentation mode | CID | CID
| Scan range for MS/MS | 120-1500 | 120-1500
| Accumulation time (s) | 0.013 | 0.018
| Dynamic collision energy | TRUE | TRUE
| Time bins to sum | 8 | 8
| Scan windows number | 80 | 80
| DIA window | m/z¦Isolation Window¦DP¦CE | m/z¦Isolation Window¦DP¦CE
|  | 350-367¦ 17¦ 80¦ 19¦ | 350-367¦ 17¦ 80¦ 16¦
|  | 367-380¦ 13¦ 80¦ 19¦ | 367-380¦ 13¦ 80¦ 17¦
|  | 380-389¦ 9¦ 80¦ 19¦ | 380-389¦ 9¦ 80¦ 18¦
|  | 389-401¦ 2¦ 80¦ 20¦ | 389-401¦ 2¦ 80¦ 18¦
|  | 401-408¦ 7¦ 80¦ 20¦ | 401-408¦ 7¦ 80¦ 19¦
|  | 408-415¦ 7¦ 80¦ 20¦ | 408-415¦ 7¦ 80¦ 19¦
|  | 415-423¦ 7¦ 80¦ 20¦ | 415-423¦ 7¦ 80¦ 19¦
|  | 423-429¦ 6¦ 80¦ 21¦ | 423-429¦ 6¦ 80¦ 20¦
|  | 429-435¦ 6¦ 80¦ 21¦ | 429-435¦ 6¦ 80¦ 20¦
|  | 435-440¦ 5¦ 80¦ 21¦ | 435-440¦ 5¦ 80¦ 20¦
|  | 440-445¦ 5¦ 80¦ 22¦ | 440-445¦ 5¦ 80¦ 21¦
|  | 445-450¦ 5¦ 80¦ 22¦ | 445-450¦ 5¦ 80¦ 21¦
|  | 450-455¦ 5¦ 80¦ 22¦ | 450-455¦ 5¦ 80¦ 21¦
|  | 455-461¦ 6¦ 80¦ 22¦ | 455-461¦ 6¦ 80¦ 21¦
|  | 461-466¦ 5¦ 80¦ 23¦ | 461-466¦ 5¦ 80¦ 22¦
|  | 466-472¦ 6¦ 80¦ 23¦ | 466-472¦ 6¦ 80¦ 22¦
|  | 472-476¦ 4¦ 80¦ 23¦ | 472-476¦ 4¦ 80¦ 22¦
|  | 476-482¦ 6¦ 80¦ 23¦ | 476-482¦ 6¦ 80¦ 22¦
|  | 482-486¦ 4¦ 80¦ 23¦ | 482-486¦ 4¦ 80¦ 23¦
|  | 486-491¦ 5¦ 80¦ 24¦ | 486-491¦ 5¦ 80¦ 23¦
|  | 491-495¦ 4¦ 80¦ 24¦ | 491-495¦ 4¦ 80¦ 23¦
|  | 495-501¦ 6¦ 80¦ 24¦ | 495-501¦ 6¦ 80¦ 23¦
|  | 501-505¦ 4¦ 80¦ 24¦ | 501-505¦ 4¦ 80¦ 24¦
|  | 505-509¦ 4¦ 80¦ 25¦ | 505-509¦ 4¦ 80¦ 24¦
|  | 509-514¦ 5¦ 80¦ 25¦ | 509-514¦ 5¦ 80¦ 24¦
|  | 514-518¦ 4¦ 80¦ 25¦ | 514-518¦ 4¦ 80¦ 24¦
|  | 518-524¦ 6¦ 80¦ 25¦ | 518-524¦ 6¦ 80¦ 24¦
|  | 524-528¦ 4¦ 80¦ 26¦ | 524-528¦ 4¦ 80¦ 25¦
|  | 528-533¦ 5¦ 80¦ 26¦ | 528-533¦ 5¦ 80¦ 25¦
|  | 533-539¦ 6¦ 80¦ 26¦ | 533-539¦ 6¦ 80¦ 25¦
|  | 539-543¦ 4¦ 80¦ 26¦ | 539-543¦ 4¦ 80¦ 25¦
|  | 543-547¦ 4¦ 80¦ 27¦ | 543-547¦ 4¦ 80¦ 26¦
|  | 547-552¦ 5¦ 80¦ 27¦ | 547-552¦ 5¦ 80¦ 26¦
|  | 552-556¦ 4¦ 80¦ 27¦ | 552-556¦ 4¦ 80¦ 26¦
|  | 556-561¦ 5¦ 80¦ 27¦ | 556-561¦ 5¦ 80¦ 26¦
|  | 561-567¦ 5¦ 80¦ 28¦ | 561-567¦ 5¦ 80¦ 27¦
|  | 567-573¦ 6¦ 80¦ 28¦ | 567-573¦ 6¦ 80¦ 27¦
|  | 573-579¦ 6¦ 80¦ 28¦ | 573-579¦ 6¦ 80¦ 27¦
|  | 579-584¦ 5¦ 80¦ 28¦ | 579-584¦ 5¦ 80¦ 27¦
|  | 584-590¦ 6¦ 80¦ 29¦ | 584-590¦ 6¦ 80¦ 28¦
|  | 590-596¦ 6¦ 80¦ 29¦ | 590-596¦ 6¦ 80¦ 28¦
|  | 596-601¦ 5¦ 80¦ 29¦ | 596-601¦ 5¦ 80¦ 28¦
|  | 601-606¦ 5¦ 80¦ 29¦ | 601-606¦ 5¦ 80¦ 28¦
|  | 606-612¦ 6¦ 80¦ 30¦ | 606-612¦ 6¦ 80¦ 29¦
|  | 612-617¦ 5¦ 80¦ 30¦ | 612-617¦ 5¦ 80¦ 29¦
|  | 617-622¦ 5¦ 80¦ 30¦ | 617-622¦ 5¦ 80¦ 29¦
|  | 622-629¦ 7¦ 80¦ 30¦ | 622-629¦ 7¦ 80¦ 30¦
|  | 629-634¦ 5¦ 80¦ 31¦ | 629-634¦ 5¦ 80¦ 30¦
|  | 634-639¦ 5¦ 80¦ 31¦ | 634-639¦ 5¦ 80¦ 30¦
|  | 639-644¦ 5¦ 80¦ 31¦ | 639-644¦ 5¦ 80¦ 30¦
|  | 644-650¦ 6¦ 80¦ 31¦ | 644-650¦ 6¦ 80¦ 31¦
|  | 650-657¦ 7¦ 80¦ 32¦ | 650-657¦ 7¦ 80¦ 31¦
|  | 657-662¦ 5¦ 80¦ 32¦ | 657-662¦ 5¦ 80¦ 31¦
|  | 662-669¦ 7¦ 80¦ 32¦ | 662-669¦ 7¦ 80¦ 32¦
|  | 669-675¦ 6¦ 80¦ 33¦ | 669-675¦ 6¦ 80¦ 32¦
|  | 675-682¦ 7¦ 80¦ 33¦ | 675-682¦ 7¦ 80¦ 32¦
|  | 682-691¦ 9¦ 80¦ 33¦ | 682-691¦ 9¦ 80¦ 32¦
|  | 691-698¦ 7¦ 80¦ 33¦ | 691-698¦ 7¦ 80¦ 33¦
|  | 698-707¦ 9¦ 80¦ 34¦ | 698-707¦ 9¦ 80¦ 33¦
|  | 707-715¦ 8¦ 80¦ 34¦ | 707-715¦ 8¦ 80¦ 34¦
|  | 715-723¦ 8¦ 80¦ 34¦ | 715-723¦ 8¦ 80¦ 34¦
|  | 723-733¦ 10¦ 80¦ 35¦ | 723-733¦ 10¦ 80¦ 35¦
|  | 733-743¦ 10¦ 80¦ 35¦ | 733-743¦ 10¦ 80¦ 35¦
|  | 743-753¦ 10¦ 80¦ 35¦ | 743-753¦ 10¦ 80¦ 35¦
|  | 753-764¦ 11¦ 80¦ 35¦ | 753-764¦ 11¦ 80¦ 36¦
|  | 764-773¦ 9¦ 80¦ 36¦ | 764-773¦ 9¦ 80¦ 37¦
|  | 773-785¦ 12¦ 80¦ 36¦ | 773-785¦ 12¦ 80¦ 37¦
|  | 785-798¦ 13¦ 80¦ 36¦ | 785-798¦ 13¦ 80¦ 38¦
|  | 798-809¦ 11¦ 80¦ 37¦ | 798-809¦ 11¦ 80¦ 38¦
|  | 809-821¦ 12¦ 80¦ 37¦ | 809-821¦ 12¦ 80¦ 39¦
|  | 821-835¦ 14¦ 80¦ 37¦ | 821-835¦ 14¦ 80¦ 39¦
|  | 835-851¦ 16¦ 80¦ 38¦ | 835-851¦ 16¦ 80¦ 40¦
|  | 851-868¦ 17¦ 80¦ 38¦ | 851-868¦ 17¦ 80¦ 41¦
|  | 868-885¦ 17¦ 80¦ 38¦ | 868-885¦ 17¦ 80¦ 42¦
|  | 885-912¦ 17¦ 80¦ 39¦ | 885-912¦ 17¦ 80¦ 43¦
|  | 912-935¦ 13¦ 80¦ 39¦ | 912-935¦ 13¦ 80¦ 44¦
|  | 935-952¦ 17¦ 80¦ 40¦ | 935-952¦ 17¦ 80¦ 45¦
|  | 952-973¦ 21¦ 80¦ 40¦ | 952-973¦ 21¦ 80¦ 46¦
|  | 973-1022¦ 49¦ 80¦ 40¦ | 973-1022¦ 49¦ 80¦ 47¦
|  | 1022-1100¦ 78¦ 80¦ 41¦ | 1022-1100¦ 78¦ 80¦ 50¦