## Donor 1
patient information | 94 years old woman, 45 kg, 140 cm tall, body donor, right sylvian and right cerebellar stroke, cognitive disorders of vascular origin, depressive syndrome atrial fibrillation and hypertensive heart disease micro-crystalline arthritis (gout), right lung pneumopathy (3 years before death) cataract of the left eye, squamous cell carcinoma of the skin (left temporal region)
sample | Donor 1 heart |  |  | Donor 1 left lung | [5 empty] | Donor 1 kidney / spleen
voxel size (µm) | 25.08 µm | 6.05 µm | 2.22 µm | 25.08 µm | 6.05 µm | 2.45 | 25.25 | 6.5 | 2.51 | 25.08 µm | 6.05 | 1.29
samnple jar diameter (mm) | 140 mm |  |  | 140 mm | [5 empty] | 80 mm
sample part | whole organ | 4 different locations | 3 different locations | whole organ | 3 different locations | 3 different locations | 1 ROI scanned twice for FSC | 1 ROI scanned twice for FSC | 1 ROI scanned twice for FSC | whole organs | central columns | central columns
optic | zoom demag Hasselblad | zoom demag Hasselblad | Hasselblad 100/300 | zoom demag Hasselblad | zoom demag Hasselblad | zoom mag canon | zoom demag Hasselblad | zoom demag Hasselblad | zoom mag canon | zoom demag Hasselblad | zoom demag Hasselblad | zoom mag canon
optic configuration | x 0.24 | x 1 | x 3 | x 0.24 | x 1 | x 2.61 | x 0.24 | x 1 | x 2.61 | x 0.24 | x 1 | x 5
date | 2020-07-05 | 2020-07-07 | 2020-07-08 | 2020-07-11 | 2020-07-26 |  | 2021-06-04 |  |  | 2020-07-12 | 2020-07-13 | 2020-07-20
average energy of the incoming beam (keV) | ~85 keV | ~87 keV | ~76 keV | ~85 keV | ~81 keV | ~70 keV | ~80 keV | ~80 keV | ~70 keV | ~85 keV | ~85 keV | ~69 keV
Average energy after sample (keV) | ~93 keV | ~95 keV | ~84 keV | ~93 keV | ~89 keV | ~79 keV | ~88 keV | ~88 keV | ~77 keV | ~89 keV | ~89 keV | ~74 keV
estimated water equivalent surface dose rate on the sample (Gy/s) | 10.5 | 4.5 | 122 | 10.5 | 12.7 | 83.3 | 48.3 | 48.3 | 128 | 10.5 | 10.5 | 161
estimated water equivalent surface dose rate on the ROI (Gy/s) | 10.5 | 1.8 | 35.9 | 10.5 | 4.9 | 26.5 | 48.3 | 18.1 | 34.7 | 10.5 | 6.6 | 83.1
integrated surface dose per scan (kGy) | 2.1137 | 0.43902 | 12.924 | 2.205 | 1.029 | 12.9585 | 7.9695 | 2.1991 | 4.2161 | 1.575 | 0.99 | 29.916
Filters (mm) | Mo 0.1 mm | Al 2mm | Mo 0.25 mm | Mo 0.1 mm | Al 3mm | Al 3mm | Mo 0.23 mm | Mo 0.23 mm | Mo 0.1 mm | Mo 0.1mm | Mo 0.1mm | Mo 0.1mm
| Al 2 mm | SiO2 bars 16*5mm | SiO2 bars 3*4mm | Al 2 mm | SiO2 bars 12*5mm | SiO2 bars 6*5mm | Al 2.3 mm | Al 2.3 mm | Al 2.3 | Al 2mm | Al 2mm | Al 2mm
| SiO2 bars 12*5mm |  |  | SiO2 bars 12*5mm |  |  | SiO2 bars 6*4mm | SiO2 bars 6*4mm | SiO2 bars 4*4mm | SiO2 bars 12*5mm | SiO2 bars 12*5mm | SiO2 bars 3*5mm
propagation distance (mm) | 3475 mm |  | 970 mm | 3475 mm |  | 970 mm | 3500 mm |  | 1440 mm | 3475 mm |  | 500 mm
sensor | sCMOS PCO edge 4.2 |  |  | sCMOS PCO edge 4.2 |  |  | sCMOS PCO edge 4.2 |  |  | sCMOS PCO edge 4.2
Scintillator | LuAG: Ce 2000 µm |  | LuAG: Ce 100 µm + lead glass meniscus | LuAG: Ce 2000 µm |  | LuAG: Ce 250 µm | LuAG: Ce 1000 µm |  | LuAG: Ce 250 µm | LuAG: Ce 2000 µm |  | LuAG: Ce 250 µm
insertion device | BM5-EBS 2PWA out of axis | BM5-EBS 2PWA central beam |  | BM5-EBS 2PWA out of axis | BM5-EBS 2PWA central beam |  | BM5-EBS 2PWA out of axis | BM5-EBS 2PWA central beam |  | BM5-EBS 2PWA central beam
filling mode | 200 mA |  |  | 200 mA |  |  | 200 mA |  |  | 200 mA
projection number | 2* 9990 | 6000 |  | 2*9990 | 6000 |  | 2*9990 | 6000 |  | 6000
scanning protocol | quarter-acquisition : one scan in half-acquisition with 800 pixels offset + one annular scan to cover the complete sample diameter | half-acquisition, offset 900 pixels |  | quarter-acquisition : one scan in half-acquisition with 800 pixels offset + one annular scan to cover the complete sample diameter | half-acquisition, offset 900 pixels |  | quarter-acquisition : one scan in half-acquisition with 800 pixels offset + one annular scan to cover the complete sample diameter | half-acquisition, offset 900 pixels |  | half-acquisition, offset 900 pixels
remarks | several problems during scanning that led to repetitive scans => some bubbling | scanned after new degasing and remounting, still few bubbles, but manageable |  | N.A. | beam crashes, but no bubbling | N.A. | twice the same series of scans for Fourier Shell Coerrelation, vertically shifted by 0.1mm |  |  | N.A. | N.A. | too long original propagation distance (1500 mm), scan done a second time with 500 mm, darkened optic, electron orbit perturbations
scanning geometry | on axis | off axis |  | on axis | off axis |  | on axis | off axis |  | on axis
references strategy | two flatfields by averaging all projections from two complete reference scans in ethanol-filled jar (tetra-acquisition protocol) | multi-references generated every 100 projections from reference scan in ethanol-filled jar |  | two flatfields by averaging all projections from two complete reference scans in ethanol-filled jar (tetra-acquisition protocol) | multi-references generated every 100 projections from reference scan in ethanol-filled jar or from averaging throughout the subscans |  | two flatfields by averaging all projections from two complete reference scans in ethanol-filled jar (tetra-acquisition protocol) | multi-references generated every 100 projections from reference scan in ethanol-filled jar or from averaging throughout the subscans |  | One flatfield by averaging all projections from a reference scans in ethanol-filled jar
subframe time (ms) | 7 | 30 | 120 | 6 | 15 | 50 | 5 | 6 | 6 | 4 | 10 | 120
exposure time (ms) | 35 | 90 | 120 | 36 | 60 | 150 | 25 | 30 | 30 | 40 | 40 | 120
accumulation level | 5 | 3 | 1 | 6 | 4 | 3 | 5 | 5 | 5 | 10 | 4 | 1
time per scan (min) | 6.71 | 8.13 | 12 | 7 | 7 | 16.3 | 5.5 | 4.05 | 4.05 | 5 | 5 | 12
number of scans | 156 | 36 | 17 | 196 | 28 | 9 | 18 | 13 | 9 | 50 | 13 | 12
subscan step (mm) | 2.3 | 4 | 4 | 2.2 | 3.5 | 1.5 | 2.2 | 2.2 | 2.2 | 2 | 3.5 | 2.3
total time (h) | 18.446 | 5.878 | 4.4 | 23.8667 | 4.2667 | 3.445 | 2.65 | 1.8775 | 1.6075 | 5.1667 | 2.0833 | 3.4
pre-processing | beam level normalisation, borders correction, lateral concatenation of central and annular scans, low frequencies adjustement from common parts, calculation of ring artefacts correction mask from high-frequencies of all projections averaging. | beam level normalisation, low frequencies adjustement based on common parts for half-acquisition, calculation of ring artefacts correction mask from high-frequencies of all projections averaging |  | beam level normalisation, borders correction, lateral concatenation of central and annular scans, low frequencies adjustement from common parts, calculation of ring artefacts correction mask from high-frequencies of all projections averaging, vertical concatenation of radiographs, correction of horizontal linear artefacts due to beam vertical power profile | beam level normalisation, low frequencies adjustement based on common parts for half-acquisition, calculation of ring artefacts correction mask from high-frequencies of all projections averaging |  | beam level normalisation, borders correction, lateral concatenation of central and annular scans, low frequencies adjustement from common parts, calculation of ring artefacts correction mask from high-frequencies of all projections averaging. | beam level normalisation, low frequencies adjustement based on common parts for half-acquisition, calculation of ring artefacts correction mask from high-frequencies of all projections averaging |  | beam level normalisation, low frequencies adjustement based on common parts for half-acquisition, calculation of ring artefacts correction mask from high-frequencies of all projections averaging
reconstruction | single distance phase retrieval with unsharp mask, filtered back-projection tomographic reconstruction, vertical concatenation of sub-scans, 16 bits conversion, ring artefacts correction, binning |  |  | single distance phase retrieval with unsharp mask, filtered back-projection tomographic reconstruction, 16 bits conversion, ring artefacts correction, binning | single distance phase retrieval with unsharp mask, filtered back-projection tomographic reconstruction, vertical concatenation of sub-scans, 16 bits conversion, ring artefacts correction, binning |  | single distance phase retrieval with unsharp mask, filtered back-projection tomographic reconstruction, vertical concatenation of sub-scans, 16 bits conversion, ring artefacts correction. Loading of the two stacks in VGStudioMax 2.3.4, precise realignment along z axis and exportation to new stacks of same size and location for the FSC tests |  |  | single distance phase retrieval with unsharp mask, filtered back-projection tomographic reconstruction, vertical concatenation of sub-scans, 16 bits conversion, ring artefacts correction, binning

## Donor 2
patient information | 69 years old woman, 40 kg, 145 cm, body donor type 2 diabetes, pelvic radiation to treat cancer of the uterus, right colectomy (benign lesion on histopathology) bilateral nephrostomy for acute obstructive renal failure, cystectomy omentectomy and peritoneal carcinoma with occlusive syndrome
sample | Donor 2 brain |  |  | Donor 2 kidney
voxel size (µm) | 25.08 µm | 6.05 µm | 2.45 µm | 25 µm | 2.5 µm
samnple jar diameter (mm) | 140 mm |  |  | 85 mm
sample part | whole organ | 4 different locations | 3 different locations | whole organs | lateral transect
optic | zoom demag Hasselblad | zoom demag Hasselblad | zoom mag Canon | zoom demag Hasselblad | zoom mag canon
optic configuration | x 0.24 | x 1 | x 2.58 | x 0.24 | x 2.5
date | 2020-07-24 | 2020-07-25 | 2020-07-25 | 2021-04-03 | 2021-04-15
average energy of the incoming beam (keV) | ~85 keV | ~81 keV | ~74 keV | ~84 keV | ~80 keV
Average energy after sample (keV) | ~93 keV | ~89 keV | ~83 keV | ~88 keV | ~84 keV
estimated water equivalent surface dose rate on the sample (Gy/s) | 10.5 | 12.7 | 42.4 | 25 | 60
estimated water equivalent surface dose rate on the ROI (Gy/s) | 10.5 | 4.9 | 14.1 | 25 | 32
integrated surface dose per scan (kGy) | 2.4822 | 1.029 | 6.6834 | 4.335 | 5.5488
Filters (mm) | Mo 0.1 mm | Al 3mm | Al 3mm | Mo 0.24mm | Mo 0.24mm
| Al 2 mm | SiO2 bars 12*5mm | SiO2 bars 8*5mm | SiO2 bars 10*5mm | Al 0.51mm
| SiO2 bars 12*5mm |  |  |  | SiO2 bars 6*4mm
propagation distance (mm) | 3475 mm |  | 1500 mm | 3475 mm | 1400 mm
sensor | sCMOS PCO edge 4.2 |  |  | sCMOS PCO edge 4.2
Scintillator | LuAG: Ce 2000 µm |  | LuAG: Ce 250 µm | LuAG: Ce 1000 µm | LuAG: Ce 250 µm
insertion device | BM5-EBS 2PWA out of axis | BM5-EBS 2PWA central beam |  | BM5-EBS 2PWA central beam
filling mode | 200 mA |  |  | 200 mA
projection number | 2* 9990 | 6000 |  | 6000
scanning protocol | quarter-acquisition : one scan in half-acquisition with 800 pixels offset + one annular scan to cover the complete sample diameter, double scan | half-acquisition, offset 900 pixels |  | half-acquisition, offset 900 pixels
remarks | N.A. | some bubbles in the hippocampus | no additional bubbling | the optic is partially darkened (~40% efficacity loss) | N.A.
scanning geometry | on axis | off axis |  | on axis | increasing off-axis from center to border of the kidney
references strategy | two flatfields by averaging all projections from two complete reference scans in ethanol-filled jar (tetra-acquisition protocol) | multi-references generated every 100 projections from reference scan in ethanol-filled jar |  | One flatfield by averaging all projections from a reference scans in a jar filled with ethanol and crushed agar gel at 70% of ethanol | One reference scan in equivalent geometry for each of the lateral position in a jar filled with ethanol and crushed agar gel at 70% of ethanol. 4 scans with sample + 4 reference scans
subframe time (ms) | 6 | 15 | 50 | 7 | 50
exposure time (ms) | 36 | 60 | 150 | 49 | 50
accumulation level | 6 | 4 | 3 | 7 | 1
time per scan (min) | 7.88 | 7 | 15.8 | 5.78 | 5.78
number of scans | 158 | 32 | 14 | 49 | 4
subscan step (mm) | 2.2 | 3.5 | 3.5 | 2.2 | 7 (horizontal shift)
total time (h) | 21.7507 | 4.7333 | 4.6867 | 5.7203 | 1.3853
pre-processing | beam level normalisation, borders correction, lateral concatenation of central and annular scans, low frequencies adjustement from common parts, calculation of ring artefacts correction mask from high-frequencies of all projections averaging. | beam level normalisation, low frequencies adjustement based on common parts for half-acquisition, calculation of ring artefacts correction mask from high-frequencies of all projections averaging |  | beam level normalisation, low frequencies adjustement based on common parts for half-acquisition, calculation of ring artefacts correction mask from high-frequencies of all projections averaging, vertical concatenation of the radiographs to generate a single scan of the complete organ, correction of the horizontal linear artefacts due to the beam vertical power profile | beam level normalisation, low frequencies adjustement based on common parts for half-acquisition, calculation of ring artefacts correction mask from high-frequencies of all projections averaging
reconstruction | single distance phase retrieval with unsharp mask, filtered back-projection tomographic reconstruction, vertical concatenation of sub-scans, 16 bits conversion, ring artefacts correction, binning |  |  | single distance phase retrieval with unsharp mask, filtered back-projection tomographic reconstruction, 16 bits conversion, ring artefacts correction, binning | single distance phase retrieval with unsharp mask, filtered back-projection tomographic reconstruction, 16 bits conversion, ring artefacts correction, lateral concatenation of reconstructed volumes, binning

## Donor 3
patient information | - 54 years old male who died of covid-19 after 21 days of hospitalization. - Mechanically ventilated. - Pulmonary failure, renal failure, bacterial pneumonia with Klebsiella aerogenes, general brain edema, subarachnoidal and intracranial bleeding. - No prior known diseases or prior medication
sample | Donor 3 upper lobe of left lung |  |  |  | Donor 3 core biopsies
voxel size (µm) | 26.38 µm | 6.24 µm | 6.24 µm | 2.22 µm | 2.25 µm
samnple jar diameter (mm) | 100 mm |  |  |  | 10 mm
sample part | whole lung lobe | 3 different locations | 5 different locations | 9 different locations | 5 core biopsies
optic | zoom demag Hasselblad | zoom demag Hasselblad | zoom demag Hasselblad | Hasselblad 100/300 | Hasselblad 100/300
optic configuration | x 0.24 | x 1 | x 1 | x 3 | x 3
date | 2020-06-02 | 2020-06-03 | 06-07/06/2020 | 09-10/06/2020 | 2020-05-29
average energy of the incoming beam (keV) | ~66 keV |  | ~64 keV | ~76 keV | ~43 keV
Average energy after sample (keV) | ~74 keV |  | ~73 keV | ~81 keV | ~44 keV
estimated water equivalent surface dose rate on the sample (Gy/s) | 5.5 |  | 8.7 | 122 | 25.6
estimated water equivalent surface dose rate on the ROI (Gy/s) | 5.5 | 2.7 | 4.2 | 56.7 | 25.1
integrated surface dose per scan (kGy) | 1.485 | 1.215 | 1.0723 | 11.4817 | 6.9276
Filters (mm) | Al 3.5 mm |  | Al 3.5 mm | Mo 0.25 mm | Al 1mm
| glassy C rods 27*5 mm |  | glassy C rods 27*5 mm | SiO2 bars 3*4mm | SiO2 rods 3*4mm
| glassy C rods 12*4 mm |  | glassy C rods 8*4 mm
propagation distance (mm) | 3675 mm |  |  | 970 mm | 605 mm
sensor | sCMOS PCO edge 4.2 |  |  |  | sCMOS PCO edge 4.2
Scintillator | LuAG: Ce 2000 µm |  |  | LuAG: Ce 100 µm + lead glass meniscus | LuAG: Ce 100 µm + lead glass meniscus
insertion device | BM5-EBS 2PWA out of axis | BM5-EBS 2PWA central beam |  |  | BM5-EBS lateral beam from DQ2C 0.39T magnet
filling mode | 200 mA |  |  |  | 200 mA
projection number | 6000 |  |  |  | 6000
scanning protocol | half-acquisition, offset 900 pixels |  |  |  | half-acquisition, offset 900 pixels
remarks | some bubbling induced by the scans | strong bubbling induced by scans | new scans performed after long degassing | no bubbling | first scan faster but unstable and leading to some movements in the samples
scanning geometry | on axis | off axis |  |  | on axis
references strategy | single flatfield generated from the averaging of all the frames of the last scan of the series (without sample in the mounting media) | multi-references generated every 100 projections from averaging of frames in each series of scans |  |  | references calculated for each vial using averaging of projections from all the scans with exclusion of scans having highly contrasted structures
subframe time (ms) | 15 | 50 | 25 | 60 | 25
exposure time (ms) | 75 | 150 | 75 | 60 | 75
accumulation level | 5 | 3 | 3 | 1 | 3
time per scan (min) | 9 | 15 | 8.51 | 6.75 | 9.2
number of scans | 68 | 93 | 107 | 52 | 88
subscan step (mm) | 2.25 | 3.5 | 3.5 | 4 | 2.2
total time (h) | 11.2 | 24.25 | 16.1762 | 6.85 | 14.4933
pre-processing | beam level normalisation, low frequencies adjustement based on common parts for half-acquisition, calculation of ring artefacts correction mask from high-frequencies of all projections averaging
reconstruction | single distance phase retrieval with unsharp mask, filtered back-projection tomographic reconstruction, vertical concatenation of sub-scans, 16 bits conversion, ring artefacts correction, binning

## Donor 4
patient information | male 77 years right lower lobe segment 6, resection of the segment due to small pulmonary adenocarcinoma (1.4 cm) coronary heart disease, arterial hypertension, chronic rheumatic disease (polymyalgia rheumatica)
sample | Donor 4
voxel size (µm) | 2.25 µm
samnple jar diameter (mm) | 10 mm
sample part | 1 core biopsy
optic | zoom mag Canon
optic configuration | x 2.58
date | 2020-07-25
average energy of the incoming beam (keV) | ~81 keV
Average energy after sample (keV) | ~82 keV
estimated water equivalent surface dose rate on the sample (Gy/s) | 12.7
estimated water equivalent surface dose rate on the ROI (Gy/s) | 12.6
integrated surface dose per scan (kGy) | 3.7044
Filters (mm) | Al 3mm
SiO2 rods 12*5mm
propagation distance (mm) | 1500 mm
sensor | sCMOS PCO edge 4.2
Scintillator | LuAG: Ce 250 µm
insertion device | BM5-EBS 2PWA central beam
filling mode | 200 mA
projection number | 6000
scanning protocol | half-acquisition, offset 900 pixels
remarks | N.A.
scanning geometry | on axis
references strategy | real references + references calculated for the complete vial using averaging of projections from all the scans with exclusion of scans having highly contrasted structures
subframe time (ms) | 20
exposure time (ms) | 80
accumulation level | 4
time per scan (min) | 9.8
number of scans | 4
subscan step (mm) | 2.2
total time (h) | 1.6533
pre-processing | beam level normalisation, low frequencies adjustement based on common parts for half-acquisition, calculation of ring artefacts correction mask from high-frequencies of all projections averaging
reconstruction | single distance phase retrieval with unsharp mask, filtered back-projection tomographic reconstruction, vertical concatenation of sub-scans, 16 bits conversion, ring artefacts correction, binning