Supplementary Material 1

Protocol for high-throughput quantification of distinct LPC species by ESI-MS/MS (for more information see https://www.lipidomicnet.org/index.php/Lysophosphatidylcholine_-_ESI-MS/MS_-_Liebisch_et_al.)1

| Sample preparation | - Lipid extraction according to Bligh and Dyer2 - Internal standards are added prior lipid extraction - To avoid loss of lipids, extraction was performed in glassware - Samples are dissolved in methanol containing 10 mM ammonium acetate/chloroform (3/1 = v/v) Material; Material used; Internal Standard(s); Internal Standard(s) added. Cultured cells; 100µg protein; LPC 13:0, LPC 19:0; 50ng each. Human plasma; 20µl; LPC 13:0, LPC 19:0; 1500ng each. |
| --- | --- |
| Instrumentation and method | Pump - Type: binary high pressure gradient (Agilent 1100) - Mode: isocratic flow gradient - Solvent(s): Methanol containing 10 mM ammonium acetate/chloroform (3/1 = v/v) - Flow gradient: Time [min]; Flow [ml/min]; % Solvent A; % Solvent B. 0; 0.05; 100; 0. 0.1; 0.03; 100; 0. 1.1; 0.2; 100; 0. 1.3; 0.05; 100; 0. Autosampler - Type: CTC Pal - Injection volume: 20µl - Wash solvent: methanol/chloroform = 1/1 Mass spectrometer - Type: Triple quadrupole (Micromass, Quattro Ultima) - Source temperature: 300°C - Ionization mode: ESI positive - Ionization voltage: 3500 V - Collision gas: Argon - Collision gas pressure: 1.0 10-3 Torr - Collision energy: 24 V - MS/MS-mode: precursor ion scan of m/z 184.1 |
| Data analysis and quantification | Data handling - Combine spectra above half peak height - Smooth combined spectrum (if necessary), centroid combined spectrum, pick peak intensities Isotope correction - Using Excel Macros correcting the peak intensities in a sequential algorithm starting from low mass species - Five isotope peaks including the monoisotopic were used - The detailed algorithm is described in the appendix of Liebisch G, et al’s study3 Calibration and quantification - Calibration type: matrix calibration - addition of naturally occurring species - Species used for calibration: Species; Cultured cells; Human plasma. LPC 16:0; 0 - 100 pmol; 0 - 1000 pmol. LPC 18:0; 0 - 100 pmol; 0 - 1000 pmol. LPC 18:1; 0 - 100 pmol; 0 - 1000 pmol. |
| Method validation | Precision - CV within-run: 3 % (major), 12 % (minor), CV total: 12 % (major), 25 % (minor) Detection limit - 0.6 - 0.8µM in plasma |

Supplementary Material 2

Protocol for high-throughput quantification of distinct PLPE species by ESI-MS/MS (for more information see https://www.lipidomicnet.org/index.php/Phosphatidylethanolamine-based_plasmalogens_-_ESI-MS/MS_-_Liebisch_et_al.)3,4

| Sample preparation | - Lipid extraction according to Bligh and Dyer2 - Internal standards are added prior lipid extraction - To avoid loss of lipids, extraction was performed in glassware - Samples are dissolved in methanol containing 10 mM ammonium acetate/chloroform (3/1 = v/v) Material; Material used; Internal Standard(s); Internal Standard(s) added. Cultured cells; 100µg protein; PE 28:0, PE 40:0 (diphytanoyl); 500ng each. Human plasma; 20µl; PE 28:0, PE 40:0 (diphytanoyl); 400ng each. |
| --- | --- |
| Instrumentation and method | Pump - Type: binary high pressure gradient (Agilent 1100) - Mode: isocratic flow gradient - Solvent(s): Methanol containing 10 mM ammonium acetate/chloroform (3/1 = v/v) - Flow gradient: Time [min]; Flow [ml/min]; % Solvent A; % Solvent B. 0; 0.05; 100; 0. 0.1; 0.03; 100; 0. 1.1; 0.2; 100; 0. 1.3; 0.05; 100; 0. Autosampler - Type: CTC Pal - Injection volume: 20µl - Wash solvent: methanol/chloroform = 1/1 Mass spectrometer - Type: Triple quadrupole (Micromass, Quattro Ultima) - Source temperature: 300°C - Ionization mode: ESI positive - Ionization voltage: 3500 V - Collision gas: Argon - Collision gas pressure: 1.0 10-3 Torr - Collision energy: 18 V - MS/MS-mode: - PE p16:0/x: neutral loss scan of m/z 364.3 - PE p18:1/x: neutral loss scan of m/z 390.3 - PE p18:0/x: neutral loss scan of m/z 392.3 |
| Data analysis and quantification | Data handling - Combine spectra above half peak height - Smooth combined spectrum (if necessary), centroid combined spectrum, pick peak intensities Isotope correction - Using Excel Macros correcting the peak intensities in a sequential algorithm starting from low mass species - Five isotope peaks including the monoisotopic were used - The detailed algorithm is described in the appendix of Liebisch G, et al’s study3 Calibration and quantification - Calibration type: matrix calibration - addition of naturally occurring species - Species used for calibration: Species; Cultured cells; Human plasma. PE p16:0/20:4; 0 - 500 pmol; 0 - 250 pmol. |

Supplementary Material 3

Protocol for high-throughput quantification of distinct PC and SPM species by ESI-MS/MS (for more information see https://www.lipidomicnet.org/index.php/Phosphatidylcholine_/_Sphingomyelin_-_ESI-MS/MS_-_Liebisch_et_al.)3

| Sample preparation | - Lipid extraction according to Bligh and Dyer2 - Internal standards are added prior lipid extraction - To avoid loss of lipids, extraction was performed in glassware - Samples are dissolved in methanol containing 10 mM ammonium acetate/chloroform (3/1 = v/v) Material; Material used; Internal Standard(s); Internal Standard(s) added. Cultured cells; 100µg protein; PC 28:0, PC 44:0; 1250ng each. Human plasma; 20µl; PC 28:0, PC 44:0; 3500ng PC 28:0, 5000ng PC 44:0. |
| --- | --- |
| Instrumentation and method | Pump - Type: binary high pressure gradient (Agilent 1100) - Mode: isocratic flow gradient - Solvent(s): Methanol containing 10 mM ammonium acetate/chloroform (3/1 = v/v) - Flow gradient: Time [min]; Flow [ml/min]; % Solvent A; % Solvent B. 0; 0.05; 100; 0. 0.1; 0.03; 100; 0. 1.1; 0.2; 100; 0. 1.3; 0.05; 100; 0. Autosampler - Type: CTC Pal - Injection volume: 20µl - Wash solvent: methanol/chloroform = 1/1 Mass spectrometer - Type: Triple quadrupole (Micromass, Quattro Ultima) - Source temperature: 300°C - Ionization mode: ESI positive - Ionization voltage: 3500 V - Collision gas: Argon - Collision gas pressure: 1.0 10-3 Torr - Collision energy: 30 V - MS/MS-mode: precursor ion scan of m/z 184.1 |
| Data analysis and quantification | Data handling - Combine spectra above half peak height - Smooth combined spectrum (if necessary), centroid combined spectrum, pick peak intensities Isotope correction - Using Excel Macros correcting the peak intensities in a sequential algorithm starting from low mass species - Five isotope peaks including the monoisotopic were used - The detailed algorithm is described in the appendix of Liebisch G, et al’s study3 Calibration and quantification - Calibration type: matrix calibration - addition of naturally occurring species - Species used for calibration: Species; Cultured cells; Human plasma. PC 34:1; 0 - 1000 pmol; 0 - 3000 pmol. PC 36:2; 0 - 1000 pmol; 0 - 3000 pmol. PC 38:4; 0 - 1000 pmol; 0 - 3000 pmol. PC 40:0; 0 - 1000 pmol; 0 - 3000 pmol. PC O 16:0/20:4; 0 - 1000 pmol; 0 - 3000 pmol. SM 16:0; 0 - 1000 pmol; 0 - 3000 pmol. SM 18:1; 0 - 1000 pmol; 0 - 3000 pmol. SM 18:0; 0 - 1000 pmol; 0 - 3000 pmol. |
| Method validation | Precision - CV within-run: 4 % (major), 5-10 % (minor); CV total: 10 % (major), 15 % (minor) Detection limit - 0.6 µM in plasma |

Supplementary Material 4

Protocol for high-throughput quantification of distinct PE species by ESI-MS/MS (for more information see https://www.lipidomicnet.org/index.php/Phosphatidylethanolamine_-_ESI-MS/MS_-_Liebisch_et_al.)3,5

| Sample preparation | - Lipid extraction according to Bligh and Dyer2 - Internal standards are added prior lipid extraction - To avoid loss of lipids, extraction was performed in glassware - Samples are dissolved in methanol containing 10 mM ammonium acetate/chloroform (3/1 = v/v) Material; Material used; Internal Standard(s); Internal Standard(s) added. Cultured cells; 100µg protein; PE 28:0, PE 40:0 (diphytanoyl); 500ng each. Human plasma; 20µl; PE 28:0, PE 40:0 (diphytanoyl); 400ng each. |
| --- | --- |
| Instrumentation and method | Pump - Type: binary high pressure gradient (Agilent 1100) - Mode: isocratic flow gradient - Solvent(s): Methanol containing 10 mM ammonium acetate/chloroform (3/1 = v/v) - Flow gradient: Time [min]; Flow [ml/min]; % Solvent A; % Solvent B. 0; 0.05; 100; 0. 0.1; 0.03; 100; 0. 1.1; 0.2; 100; 0. 1.3; 0.05; 100; 0. Autosampler - Type: CTC Pal - Injection volume: 20µl - Wash solvent: methanol/chloroform = 1/1 Mass spectrometer - Type: Triple quadrupole (Micromass, Quattro Ultima) - Source temperature: 300°C - Ionization mode: ESI positive - Ionization voltage: 3500 V - Collision gas: Argon - Collision gas pressure: 1.0 10-3 Torr - Collision energy: 20 V - MS/MS-mode: precursor ion scan of m/z 141.0 |
| Data analysis and quantification | Data handling - Combine spectra above half peak height - Smooth combined spectrum (if necessary), centroid combined spectrum, pick peak intensities Isotope correction - Using Excel Macros correcting the peak intensities in a sequential algorithm starting from low mass species - Five isotope peaks including the monoisotopic were used - The detailed algorithm is described in the appendix of Liebisch G, et al’s study3 Calibration and quantification - Calibration type: matrix calibration - addition of naturally occurring species - Species used for calibration: Species; Cultured cells; Human plasma. PE 34:1; 0 - 500 pmol; 0 – 250 pmol. PE 36:2; 0 - 500 pmol; 0 – 250 pmol. PE 38:4; 0 - 500 pmol; 0 – 250 pmol. PE 40:6; 0 - 500 pmol; 0 – 250 pmol. |

Supplementary Material 5

Protocol for high-throughput quantification of distinct ceramide (CER) species by ESI-MS/MS (for more information see https://www.lipidomicnet.org/index.php/Ceramide_-_ESI-MS/MS_-_Liebisch_et_al.)6

| Sample preparation | - Lipid extraction according to Bligh and Dyer2 - Internal standards are added prior lipid extraction - To avoid loss of lipids, extraction was performed in glassware - Samples are dissolved in methanol containing 5 mM ammonium acetate/chloroform (3/1 = v/v) Material; Material used; Internal Standard(s); Internal Standard(s) added. Cultured cells; 100µg protein; Cer 14:0, Cer 17:0; 50ng each. Human plasma; 20µl; Cer 14:0, Cer 17:0; 50ng each. |
| --- | --- |
| Instrumentation and method | Pump - Type: binary high pressure gradient (Agilent 1100) - Mode: isocratic flow gradient - Solvent(s): Methanol containing 10 mM ammonium acetate/chloroform (3/1 = v/v) - Flow gradient: Time [min]; Flow [ml/min]; % Solvent A; % Solvent B. 0; 0.05; 100; 0. 0.1; 0.03; 100; 0. 1.1; 0.2; 100; 0. 1.3; 0.05; 100; 0. Autosampler - Type: CTC Pal - Injection volume: 20µl - Wash solvent: methanol/chloroform = 1/1 Mass spectrometer - Type: Triple quadrupole (Micromass, Quattro Ultima) - Source temperature: 250°C - Ionization mode: ESI positive - Ionization voltage: 3500 V - Collision gas: Argon - Collision gas pressure: 1.0 10-3 Torr - Collision energy: 30 V - recursor ion scan of m/z 264.2 - multiple reaction monitoring table of species observed frequently Analyte; Precursor [m/z]; Precursor [m/z]; Collision energy [eV]. Cer 14:0 - H2O (IS); 492.5; 264.2; 25 V. Cer 14:0 (IS); 510.5; 264.2; 25 V. Cer 16:0 - H2O; 520.5; 264.2; 25 V. Cer 16:0; 538.5; 264.2; 25 V. Cer 17:0 - H2O (IS); 534.5; 264.2; 25 V. Cer 17:0 (IS); 552.5; 264.2; 25 V. Cer 18:0 - H2O; 548.5; 264.2; 25 V. Cer 18:0; 566.6; 264.2; 25 V. Cer 20:0 - H2O; 576.6; 264.2; 25 V. Cer 20:0; 594.6; 264.2; 25 V. Cer 22:1 - H2O; 602.6; 264.2; 25 V. Cer 22:1; 620.6; 264.2; 25 V. Cer 22:0 - H2O; 604.6; 264.2; 25 V. Cer 22:0; 622.6; 264.2; 25 V. Cer 23:0 - H2O; 618.6; 264.2; 25 V. Cer 23:0; 636.6; 264.2; 25 V. Cer 24:1 - H2O; 630.6; 264.2; 25 V. Cer 24:1; 648.6; 264.2; 25 V. Cer 24:0 - H2O; 632.6; 264.2; 25 V. Cer 24:0; 650.6; 264.2; 25 V. |
| Data analysis and quantification | Data handling - Combine spectra above half peak height - Smooth combined spectrum (if necessary), centroid combined spectrum, pick peak intensities Isotope correction - Using Excel Macros correcting the peak intensities in a sequential algorithm starting from low mass species - Five isotope peaks including the monoisotopic were used - The detailed algorithm is described in the appendix of Liebisch G, et al’s study3 Calibration and quantification - Calibration type: matrix calibration - addition of naturally occurring species - Both [M+H]+ and [M+H-H2O]+ are use - Species used for calibration: Species; Cultured cells; Human plasma. Cer 16:0; 0 - 35 pmol; 0 - 50 pmol. Cer 18:0; 0 - 35 pmol; 0 - 50 pmol. Cer 20:0; 0 - 35 pmol; 0 - 50 pmol. Cer 24:1; 0 - 35 pmol; 0 - 50 pmol. Cer 24:0; 0 - 35 pmol; 0 - 50 pmol. |
| Method validation | Precision - CV over all: below 10% in plasma samples Detection limit - 0.3 pmol injected amount (in fibroblast lipid extract) Recovery - Recovery of Cer 6:0 above 90% |

References

1.	Liebisch, G., Drobnik, W., Lieser, B. & Schmitz, G. High-throughput quantification of lysophosphatidylcholine by electrospray ionization tandem mass spectrometry. Clin. Chem. 48, 2217–2224 (2002).

2.	Bligh, E. G. & Dyer, W. J. A rapid method of total lipid extraction and purification. Can. J. Biochem. Physiol. 37, 911–917 (1959).

3.	Liebisch, G., Lieser, B., Rathenberg, J., Drobnik, W. & Schmitz, G. High-throughput quantification of phosphatidylcholine and sphingomyelin by electrospray ionization tandem mass spectrometry coupled with isotope correction algorithm. Biochim. Biophys. Acta 1686, 108–117 (2004).

4.	Zemski Berry, K. A. & Murphy, R. C. Electrospray ionization tandem mass spectrometry of glycerophosphoethanolamine plasmalogen phospholipids. J. Am. Soc. Mass Spectrom. 15, 1499–1508 (2004).

5.	Brügger, B., Erben, G., Sandhoff, R., Wieland, F. T. & Lehmann, W. D. Quantitative analysis of biological membrane lipids at the low picomole level by nano-electrospray ionization tandem mass spectrometry. Proc. Natl. Acad. Sci. U. S. A. 94, 2339–2344 (1997).

6.	Liebisch, G. et al. Quantitative measurement of different ceramide species from crude cellular extracts by electrospray ionization tandem mass spectrometry (ESI-MS/MS). J. Lipid Res. 40, 1539–1546 (1999).